AL05-013									June 8, 2005

Item:  	Guidant Model 1861 Ventak Prizm 2 DR, Implantable Cardioverter-Defibrillators (ICD) manufactured before November 2002

### General Information:	Guidant Corporation submitted voluntary advice to physicians, by letter on May 23, 2005, about an unpredictable breach of an insulator in the device. This insulator defect could result in at least partial diversion of current in the high-voltage output circuitry, thereby preventing the device from delivering high-voltage shock therapy when ventricular tachycardia or fibrillation is detected. This random component failure can occur in any Model 1861 ICD manufactured before November 13, 2002.

Specific Incident:	Guidant reported that there have been 26 cases exhibiting the above failure mode from 37,000 devices implanted worldwide.  In one of those cases, the patient may have died when the device failed to deliver high-voltage shock therapy.

Actions:	1. Within two weeks, electrophysiology/cardiology staff or other appropriate parties must identify all affected patients by implementing each of the following steps a through d. It is important that ALL INFORMATION sources be reviewed to insure that patients will not be missed, as they may be found on one list and not on another.

a)  Review the manufacturer’s letter (attached).

b)  Review the patient list posted on the VA National ICD Surveillance Center intranet website (https://icd.sanfrancisco.med.va.gov).  This list, provided by Guidant, consists of all the VA patients in the company’s database that have an implanted model 1861 ICD which had been manufactured prior to November 13, 2002 and VA patients having this implant that are being followed at a VA facility.

- Review the VA National registry for ICD implants Washington VAMC, point of contact is Ronald.Jones1@va.gov.

- Review your patient records for all patients with implanted Guidant model 1861 Ventak Prizm 2 DR, ICD devices.

2. Within the next 45 calendar days, interrogate all Guidant Model 1861 ICDs for “no telemetry” or “warning screen” conditions. These conditions indicate that the ICD may be inoperative. If one of these conditions is present, replace with a suitable new device.

a) If the interrogation of the Guidant model 1861 ICD does not reveal a problem, the patient should be followed at the manufacturer’s recommended intervals of every 3 months.  However patients should be instructed to return immediately for device interrogation following any shock delivery, and ICD replacement should be considered at that time.

3. Follow the actions contained in Attachment 2. This guidance was prepared by Dr. Edmund Keung of the VA National ICD Surveillance Center as the best course of action for your patients.

Additional Information:	See Guidant website www.guidant.com

Source:   		Guidant Corporation.

Contact: 	Dr. Edmund Keung at VA National ICD Surveillance Center at

(415) 221-4810 Extension 3182 Or

Mr. Bryanne Patail at National Center for Patient Safety (NCPS) at (734) 930-5890

ATTACHMENT 1

ATTACHMENT 2

|  | DEPARTMENT OF VETERANS AFFAIRS Medical Center 4150 Clement Street San Francisco CA 94121 VA National ICD Surveillance Center |  |
| --- | --- | --- |

June 1, 2005

Dear colleagues:

This document is to provide you with more technical information on the issue involving the Guidant Model 1861 Ventak Prizm 2 DR ICD performance.

- The Problem:

  - The ICD can fail to deliver shock of any magnitude when it detects ventricular tachy arrhythmias (VT or VF).  The occurrence is random.

In order for the problem to occur, the following 4 events have to occur:

- The wire (which acts as the negative pole) has to position close enough to the tube in the header that is filled with the inert gas (which is the positive pole)

- There has to be a breach of insulation on the wire

- There has to be a variation in the adhesive that is between the tube and the wire

- The shock has to jump from the negative pole to the positive pole to short out the can

- The problem does not occur during cap reform.  Normal cap reform does not have predictive value.

  - The history of the ICD performance has no predictive value.  A device can provide successful shock yesterday and fail to deliver shock today when VT/VF are detected.  Similarly, doing a successful defibrillation test has no predictive value.

- The bottom line: There is no way to predict which device will have this problem.  Routine follow-up will not disclose the problem.  The only way to find out is when the patient needs a therapeutic shock (and it fails)!

(B) Try to identify as many affected patients as possible.  There are several sources:

(1)	The list of your affected patients that has been provided by your local Guidant representative.

(2)	Bryanne Patail, the biomedical engineer for the VA National Center for Patient Safety obtained a list of 225 patients from Guidant.  These patients had either a model 1861 ICD manufactured prior to November 13, 2002 implanted or are being followed at a VA facility.  You can go to the VA National ICD Surveillance Center website (<https://icd.sanfrancisco.med.va.gov>) to search for your own patients.  Please remember, this list is not complete. There are likely to be patients who are being followed in your clinics and not on this list.

(3)	Washington DC VAMC maintains a national registry for ICD implants.  Mr. Ron Jones can help you with your search. Ronald.Jones1@va.gov

(4)	Your own local database.  This is the best, if you have one.

(C)	What to do?

Guidant does not recommend replacement of the devices.  We have to weigh the incidence and consequence of battery failure against potential complications with generator replacement.  The incidence of circuitry failure is low, at 1 in 2000 (0.05% or 0.002% per service month).  However, if it fails, the result can be catastrophic for the patient.  In patients with a history of frequent shock therapy or who are totally pacemaker-dependant, device replacement may be considered.  We have to evaluate individual patient’s clinical conditions, discuss the situation with the patients and arrive at the best course of action.

Ed Keung, MD

Ph: 415-221-4810, extension 3182

Edmund.Keung@med.va.gov

# Text boxes

Alert

6

5

4

3

Instructions on how to access the VA National ICD Surveillance center database.

Perform the following steps to access your patient list supplied by Guidant on the VA National ICD Surveillance Center:

- VA intranet URL: https://icd.sanfrancisco.med.va.gov

- You have to register as a user first. Len Roberts, our administrator (Leonard.Roberts@med.va.gov)   	will review the information you provided and grant you access within 24 hours or less

- After you log in, click on <Queries & Reports> and <Patient DeviceSearch>.

- Select ICD Generator from the Filter by Device Type dropdown list.

- Select Guidant from the Filter by Manufacturer drop-down list.

- Enter “1861” in the Search Model textbox.

- Click Go

You will see a list of all your patients with the affected devices.  The device alert status is listed in the far right corner under the column heading Alert (Y=Yes).  Do not forget that there may be more than one page for the list, depending on how many patients you have.  You can export the table to an Excel spreadsheet or just print it.

You will notice that almost all your patients have SSN (social security number) and phone numbers of 888-88-8888 and (415) 221-4810, respectively.  Guidant did not include the patients’ SSN and phone numbers in their list. The medical centers listed under the column VAMC are the hospitals where they had their device implanted or the follow-up clinics, according to Guidant’s records.

Status: R=registration only