Cycle Sequencing using BigDye v3.1

Performed on Rolling Circle Amplified DNA Template

1/16th Reaction (0.5BD)

Version Number:		2.1

Production Start Date:	10/01/03

Version 2.1 Date:		05/06/04

Authors:			Arshi Khan, Nancy Hammon, Catherine Gordon

Reviewed/Revised by:	Amber Nivens

# Summary

The BigDye v3.1 sequencing kit is used to cycle sequence plasmid DNA, which has been amplified by rolling circle amplification, in preparation for sequencing on the ABI 3730xl instruments.

# Materials & Reagents

| Materials/Reagents/Equipment Disposables 384-well PCR Microplate, Clear 50-ml Centrifuge Tube PlateLoc Peelable Heat Seal LTS Pipette Tips (2, 10, 100, 200, 1000μl) 175 Graduated Conical tube with cap Reagents BigDye Terminator v3.1 Cycle Sequencing Kit (with 5X BigDye Terminator Seq Buffer) FW Sequencing Primer RV Sequencing Primer Coulter Clenz 100% Ethanol (200 Proof) Milli-Q Water Equipment 384-well Hydra Twister Hydra Wash Reservoir Cavro Dispensing Robot Pipettes (2, 10, 100, 200, 1000μl) GeneAmp PCR System 9700 PlateLoc Thermal Plate Sealer 250 ml, 500 ml bottles 600 ml beakers PTFE Release Agent Dry Lubricant 125 ml double pharmaceutical graduate | Vendor Axygen/ISC BioExpress Corning/VWR Velocity 11 Rainin Corning Applied Biosystems IDT IDT Beckman Coulter AAPER Alcohol (LBNL) Millipore Milli-Q System Robbins Scientific Robbins Scientific JGI Rainin Perkin Elmer (Applied Biosystems) Velocity 11 Pyrex Pyrex Miller-Stephenson Falcon/VWR | Stock Number T-3059-1CS - 430290 06643-001 352076 4336923 - PN 8546930 6810-13289 - - - - - - - - - - - |
| --- | --- | --- |

# EH&S

JGI employee performing this procedure must wear a lab coat and gloves.  In situation were there might be a chance of an accidental splash to the eyes, safety glasses must be worn.

# Calibration Check

Hydra Twisters

NOTE: This should be performed on one FW and one RV plate within the first 10 plates per batch per Hydra Twister at step 2.12.

- Without disturbing the Hydra Twister, retrieve an empty FW destination plate from Stacker 4.

- Place the empty FW destination plate on the weighing balance and press the TARE button.

- Return the tared empty FW destination plate back to Stacker 4.

- After the tared FW destination plate has been aliquoted, remove it from Stacker 5 and replace it with an empty Axygen plate.

- Reweigh the aliquoted FW destination plate on the weighing balance.

NOTE 1: If the weight ranges fall outside of 0.3366g – 0.3743g, WARNING.

NOTE 2: If the weight ranges fall outside of 0.3272g – 0.3837g, DO NOT USE.

- Enter the weight into the Plate Weight Excel Spreadsheet.

NOTE 1: To open the Excel spreadsheet, click on Shortcut to Plate Weight on desktop.

NOTE 2: If the weight entered in the excel spreadsheet is too close to the high or low range, the number entered in the weight column will be highlighted in a red box and the text WARNING will appear in orange under the Spec. column.   If this occurs, continue to use the Hydra but monitor it carefully so that it does not fall out of range.  If it is continuously high or low and in the WARNING zone, contact your supervisor to determine how to proceed.

NOTE 3:  If the weight entered in the excel spreadsheet falls out of the range, the number entered in the weight column will be highlighted in a red box and the text DO NOT USE will appear in a red box under the Spec. column.  If this occurs, discontinue use of the hydra and submit a work request for calibration.

- Remove the empty Axygen plate from Stacker 5 on the Hydra Twister and replace it with the aliquoted weighed FW destination plate.

- Repeat steps 1 – 7 with a RV destination plate.

- Click Save on the Excel spreadsheet.

Cavro 3

NOTE: This should be performed on the first plate of each batch dispensed during step 3.6c.

- Place the first FW sample plate on the weighing balance and press the TARE button.

- Place the first RV sample plate on the weighing balance and note the difference with the FW plate.

- Place both plates onto the appropriate Cavro stage.

- Click on START.

  - Watch for droplets on the tips and missed wells on the plate.

  - Visually check the bottom of the plate to make sure that the volumes in all of the wells are uniform.

NOTE:  If tips are clogged, wells are missing, or volumes are not dispensed correctly, submit a work request & call Instrumentation to get the problem fixed.

- Upon completion of dispensing, reweigh both plates to check the volume of sequencing chemistry cocktail that is being dispensed and record the data in the Cavro Excel spreadsheet.

- If weights are all within the range continue to dispense all the plates in the batch. Acceptable dispense range for 4ul of BD v3.1 (0.5rxn) chemistry is 1.501-1.571g.

NOTE 1: If the weight entered in the excel spreadsheet is too close to the high or low range, the number entered in the weight column will be highlighted in a red box and the text WARNING will appear in orange under the Spec. column.   If this occurs, continue to use the cavro but monitor it carefully so that it does not fall out of range.  If it is continuously high or low and in the WARNING zone, contact your supervisor to determine how to proceed.

NOTE 2:  If the weight entered in the excel spreadsheet falls out of the range, the number entered in the weight column will be highlighted in a red box and the text DO NOT USE will appear in a red box under the Spec. column.  If this occurs, submit a work request to instrumentation for calibration.

# Procedure

NOTE 1:  Make sure to wear gloves throughout the entire Sequencing Chemistry process.

NOTE 2: Heat Inactivation (Step 8 of the Rolling Circle Amplification on Plasmid Glycerol Stock protocol) must be completed prior to aliquoting the sequencing template into the FW and RV plates.

NOTE 3: If you are resequencing plates, go to the Venonat database ( click on RCA DB ( click on Scan in RCA plates for Re-Sequencing only! ( enter in: operator and direction ( click on SUBMIT ( print the barcodes and affix to 384-well PCR plates.

- Sample Preparation

  - Remove the BigDye v3.1 Terminator Kit, respective FW/RV primers, & any leftover cocktail corresponding with respective primers from the -20(C freezer #202 located in Rm 140.

  - Let the reagents thaw on ice.

NOTE 1:  Only remove the one primer type at a time.  Return it back to the freezer as soon as the mix is made.  Label the primer tube with the date and the batch number prior to replacing in the freezer.  The ice bucket may only contain one primer type at any time.

NOTE 2:  DO NOT CONSOLIDATE PRIMER TUBES.

NOTE 3: The BigDye Kit usually takes 30 minutes to thaw, but may take up to two hours.

  - Remove the 5X buffer from the 4oC deli located in Rm140 and place it in the same ice bucket as BigDye reagents.

  - Place a 175 ml Corning Conical tube of fresh Milli-Q water and place it on ice.

NOTE: Make sure water is chilled prior to making the sequencing chemistry cocktail

- Aliquoting (Hydra Twister Robot)

NOTE:  Make sure to randomize aliquoting batches with at least 2 or more RCA batches that have the same primer association.

  - At the completion of Heat Inactivation (Step 8 of the Rolling Circle Amplification on Plasmid Glycerol Stock protocol) FW and RV destination barcodes were printed. Affix these FW and RV barcodes to new Axygen plates (on the numbered and indented side of the plate).

NOTE: Make sure that the barcodes on the destination plates are marked with the correct primer association color against the database and against the tubes you pulled (key located on the countertop).

2.2	Obtain the RCA source plates from final heat off of the thermocyclers or from the 4oC deli in the Sequencing Preparation room (Rm140) and remove the foil seals.

Checkpoint:  Visually check the RCA source plate after removing the seal to ensure that the proper volume is in each plate.

2.3	Match the source plates with their corresponding destination plates (the stacks of plates should be in the same order).

2.4	In the Venonat database click on RCA DB ( STEP 4A - Scan RCA Plates to Update Aliquot Information ( enter in:

- Select the correct information from the drop-down menus: Hydra, Aliquot operator.

- Use the barcode scanner to scan in the RCA source plates.

- Click on SUBMIT.

- Scan in the corresponding FW and RV plates and click on SUBMIT.

Checkpoint: If there are any mismatched plates correct the mismatches before loading these plates onto the Hydra Twister stackers.

2.5	Place the RCA templates (source plates) in Stacker 1 and the destination plates (FW and RV pairs) in Stackers 3 and 4 (load stacker 4 up with a maximum of 32 plates before loading stacker 3). The top source plate should match the top destination plate in stacker 4. To distinguish Twister 2 plates from Twister 1 plates, add black dots to the FW and RV barcodes of Twister 2 plates.

NOTE: Press down the emergency stop button located next to the Hydra Twister to manually move the Twister arm.

2.6	At the beginning of each week, move the supply and drain Masterflex tubing to a new slot and write the date on the tubing.

NOTE: There are 4 slots on the tubing. When all slots have been filled once, submit a work request to have the tubing changed (approx. every month).

2.7	At the beginning of each day, turn on the Hydra Twister computers, log on, then push down and twist up each emergency stop button located next to the Hydra Twister.

2.8	Open the Aliquot AUTOSTART1 program on the computer monitor.

  - Run the WASH_100μlx3 program.

2.10	After completion of the WASH_100μlx3 program, exit the program.

NOTE: Each time a task is completed in the Aliquot Autostart1 program, you must close the program and re-open it.

2.11	Re-open Aliquot AUTOSTART1 and click on Aliquot Axygen 384’s.

Checkpoint: The program will ask if there is enough water. Check the supply container and if it is less than half full, press the START button located above the supply container. If the water container is greater than half full, click OK on the program window. This must be done for every batch.

NOTE 1: This program will aliquot 1μl per well from the RCA plates into each well of the corresponding FW and RV plates in preparation for sequencing.

NOTE 2: Periodically monitor the water trough to ensure water is filling and draining correctly. If water trough is not filling or draining properly or if the program was stopped abruptly, then manually fill or drain the trough by pushing the corresponding button towards manual on the Twister Hydra Pump Controller located next to the Hydra Twister. Make sure to switch the button back to automatic when finished. If the water trough is not filling or draining properly, also submit a work request.

2.12	Weigh one FW and one RV plate from each Hydra Twister (Refer to Hydra Twisters under the Calibration Check section).

2.13	After the program is complete and the pump has stopped running, close the window and push down and twist up the emergency stop button.

NOTE: This will move the robotic arm so that the RCA source plates may be removed.

2.14	Remove source and destination plates from the Hydra-Twister stackers.

2.15	Re-open Aliquot AUTOSTART1 and run the WASH_100μlx3 program.

NOTE: This program must be run in between each batch/run.

2.16	For the source RCA plates:

- Place clear adhesive seals on all source RCA plates.

- From each RCA batch:

- Isolate 6 random plates.

  - Wrap the plates with labeling tape.

  - Label: RCA batch # and RCA date.

  - Put the plates in the 4oC deli located in Rm 140 in the QC Gels bin.

- Wrap the remaining plates in each batch with labeling tape.

- Label: ABI sequencing batch # and date.

- Put the plates in the 4oC deli located in Rm 140 on the correct shelf for that day.

NOTE: Steps 2.16 a-e can be done while running the Cavro during step 3.6. After the plates have been stored in the 4oC deli for 1 week, throw them away in the autoclave waste container.

2.17	For the FW and RV plates:

a. Stack FW and RV plates into FW and RV sets.

b. Visually check destination plates to make sure that all 384 wells contain RCA                                                                                                                                                                                               template.

Checkpoint: If 4 or more wells per plate are empty or overfilled, re-aliquot the plate manually with a Hydra. If there is something wrong with a tip (e.g. clogged tip), submit a work request. If the source plate is the cause of the missing wells, see the supervisor to make a determination on how to proceed.

c.   Place a clear adhesive seal on the top plate of each stack.

d. Store the stacks in the 4oC deli located in Rm 140 until it is time to add the cocktail. If the plates are aliquoted one hour or more before the cocktail will be added, add a clear adhesive seal to each plate in the stack before storing the plates in the deli.

2.18	At the end of the day, place a trough with Coulter Clenz cleaning reagent (located under the sink in room 140) in the destination area and run the Trough Wash program.

- Enter needle fill amount: 80%. Click OK.

- Enter the number of wash cycles: 5. Click OK.

- Enter the soak time: 60 sec. Click OK.

- Enter the number of bath wash cycles: 5. Click OK.

2.19      To shut down the Hydra Twister Robot for overnight storage:

- Click on End of the Day Fill to fill the needles with water.

- Turn off the robot by pushing down on the red emergency stop button located next to the Hydra Twister.

- Log off of the computers.

- Sequencing Chemistry

  - Label two 175-ml Corning conical tubes FW or RV, and place them on ice.

  - Prepare the BigDye cocktail (Refer to the Reagent/Stock Preparation section) using the two 175-ml tubes (one for FW and one for RV).

NOTE 1: When dispensing on Cavro 3 – make enough cocktail for 6 extra plates of FW and RV so the syringes don’t run dry. When dispensing by hand – make 20% extra cocktail for each.

  - Cap each FW and RV BigDye cocktail tube, and mix by swirling gently and inverting each tube 3 to 4 times.

  - Place two 175ml Corning conical Tubes into the respective two 600-ml beakers of ice labeled FW and RV.

  - Using Cavro 3, dispense 4μl of the FW and RV BigDye cocktail into each well of corresponding FW and RV 384-well plates.

  - Log onto the computer for Cavro 3.

  - Click on the Cavro 3 Barcodes icon located on the desktop.

  - Check the Cavro 3 calibration for each batch (Refer to the Calibration Check section).

  - Remove the FW Cavro tubing from the FW water bottle.

  - Replace the FW water bottle located on the FW Cavro stand with the FW beaker containing FW cocktail.

  - Place the FW tubing into the FW 175 ml corning conical tube.

  - Make sure the FW stopper is at the bottom of the tube.

  - Repeat steps 3.6.d – 3.6.g for the RV Cavro and RV cocktail.

  - Under Select Stations, select BOTH.

  - Under Select Process, select PRIME.

NOTE 1: Before priming, make sure that the troughs do not have water in them.

NOTE 2: Ensure that the feed tubes are submersed in BigDye cocktail and contain no air bubbles

  - Dump the FW and RV trough contents back into corresponding FW and RV conical tubes of BigDye cocktail.

  - Remove the FW and RV plate stacks from the 4oC deli located in Rm140.

  - Place a FW plate onto its corresponding FW plate stand (bottom Cavro stage) and place a RV plate onto its corresponding RV plate stand (top Cavro stage).

  - Click on START.

  - Upon completion of the dispensing run, take the plates off of the stands and visually check to make sure that cocktail has been dispensed into each well.

Checkpoint: If 4 wells or less per plate are empty, manually add 4μl of cocktail to the empty wells with a pipette. If more than 4 wells are empty or overfilled, re-aliquot the plate with the hydra and run it through the Cavro again. If there is something wrong with a tip (e.g. clogged tip), submit a work request and contact Instrumentation to resolve the problem.

NOTE: While Cavro is dispensing, the source plates can be sealed.

  - Place the plates on ice.

  - Repeat steps 3.6m – 3.6p until cocktail has been dispensed into all of the plates.

NOTE: Instead of clicking on START in step 3.6n, you may now press ENTER on the keyboard to start a dispensing run.

  - Save the leftover BigDye cocktail after each batch:

  - Under Select Process, select PRIME.

  - Dump the remaining cocktail from the FW and RV 175 ml Corning conical tubes and FW and RV troughs into corresponding FW and RV 50-ml tubes labeled with: the date, either FW or RV, batch number, primer type, primer lot number, BigDye lot #, & 5X Buffer Lot #.

  - Place the tubes in the -20oC freezer located in Rm 140 or keep on ice if they will be used again.

  - DO NOT USE leftover cocktail from a previous batch during the day.  This should be refrozen after completing the batch.

  - Place plates in a centrifuge for a quick spin (1 minute at 1,000rpm).

  - Use the PlateLoc Thermal Plate Sealer to seal each plate with a white foil seal.

  - In the Venonat database, click on RCA DB ( STEP 4B - Scan Plates for Cycle Sequencing:

    - Fill in the boxes: Dispense Operator, Dispense Instrument, FW Primer Lot, FW Kit Lot, RV Primer Lot, RV Kit Lot, Mix Operator, Days.

    - Scan in the plates in FW and RV pairs to thermocyclers according to which side of the thermocycler a plate will be loaded.

    - Click on SUBMIT.

    - Print this page and add it to the sequencing chemistry binder.

  - Place the plates on the correctly matched thermocyclers.

  - Run the BigDye program on the thermocyclers (refer to Instrument Maintenance & Settings for cycle program).

NOTE: Use the computer to start thermocyclers if all thermocyclers are filled on the corresponding rack. Otherwise, start thermocyclers individually.  After starting thermocyclers, verify that each thermocycler has been started on the correct program.

  - Wash Cavro 3 after each batch.

  - Place the 175ml Corning conical tubes in the autoclave bin.

  - Place the 250-ml water bottles back on their respective FW and RV stands.

  - Place the Cavro tubing back into the bottles.

  - Under Select Process, select PRIME and repeat 3 times.

NOTE: Empty the FW and RV troughs into waste containers after each wash.

  - Rinse out the FW and RV troughs by running them under Milli-Q water.

  - Add fresh Milli-Q water to each 250-ml bottle when the water levels get low.

  - Place the FW and RV troughs back into their slots.

  - After the thermocycler program ends (approx. 2.5 hours later), remove the paired plates from the thermocyclers and place them in a bin labeled: ABI batch # and date.

NOTE: Stop the thermocyclers either manually or by computer.  If the plates do not come out easily, use another plate or a plate release tool to pry the plates up off the thermocyclers. When the plates start getting difficult to remove, spray the thermocyclers with PTFE Release Agent Dry Lubricant.

  - Place the plates in the 4oC fridge located in Rm141 for the BET cleanup protocol.

  - At the end of the day:

a.  On Mondays, clean the Cavro by priming 70% Ethanol through the FW & RV tubes 3 times. Follow the ethanol wash with a water wash by priming Milli-Q water through the tubes 5 times.

- Replace the two 250-ml water bottles on the FW and RV Cavro stands with clean bottles filled with fresh Milli-Q water.

- Under Select Process, select PRIME.

- Leave the Cavro tubing filled with water.

- Empty & rinse out the FW and RV troughs by running them under Milli-Q water.

- Close the Cavro 3 Barcodes program.

- Log off of the computer

# Reagent/Stock Preparation

BigDye v3.1 Cocktail, 0.5BD reaction for Plasmid Template

Per one 384-well Plate:

1.05 ml Milli-Q Water

190 μl BigDye v3.1

290 μl 5X Sequencing Buffer

6.1 μl Primer (250μM)

- When preparing fresh cocktail mix, prepare enough cocktail for the number of source plates in your sample plus 6 additional plates according to the cocktail volume per number of plates chart located in Sequencing Chemistry bay and also at the end of this protocol in Appendix B. To calculate how much to prepare when using recycled BigDye: Take the total volume you need and subtract the volume of leftover BigDye cocktail. Prepare cocktail based on the final volume.

Example: 20-source plate batch.

  - You need enough cocktail volume for 26 plates/FW primer and 26 plates/RV primer (total volume per primer = 39.896ml).

  - You have 5ml leftover BigDye cocktail.

  - You need to make 34.896ml fresh BigDye cocktail mix.

  - Prepare enough for a 23-source plate batch (total volume per primer = 35.293ml) and add to the leftover BigDye cocktail.

- Place two 175 ml Corning conical tubes in the bucket.  Mark each tube with FW or RV for the two different primer directions.

- Add buffer, water, BigDye, primer, and any leftover cocktail to each 250-ml bottle and mix by swirling and inverting the bottle.

NOTE 1: Keep all reagents on ice while preparing cocktail and dispensing into the 384-well plates.

NOTE 2: If a tube of leftover cocktail has not been used within the last month, throw it out and make completely fresh mix.

70% Ethanol

- Pour 350 ml of 100% ethanol into a 500-ml bottle.

- Add 150 ml of Milli-Q water.

- Mix well by inverting bottle.

- Label bottle: 70% ethanol, date prepared, and your initials.

# Instrument Maintenance & Settings

Perkin Elmer 9700/Applied Biosystems 9700

Program Big Dye

95oC for 1 minute

95oC for 30 seconds

50oC for 20 seconds	   25 cycles

60oC for 4 minutes

Hold at 4oC

Volume = 5μl

Cavro 3

Setup:

  - Log onto the computers located next to the Cavros.

  - Double-click on the Cavro 3 Barcodes icon.

Calibration:

Refer to the Calibration Check section on page 2.

Shut-down:

  - Close the Cavro 3 Barcodes program.

NOTE:  Be sure that the tubes are full of water and that the water bottles are replaced daily with clean glassware and Milli-Q water.

  - Log off the computer.

# Troubleshooting

Procedure for Thermal Cycler Recovery After a Power Outage

Summary

This procedure details how to ensure thermal cycler recovery after a power outage and how to handle production plates running on the cyclers at the time.

NOTE: Action is Based on Building Power Outage Duration

<15secs

No action other than an inspection needs to be taken for an outage less than 15 seconds long. The operator should check all the cyclers affected to ensure the runs are continuing as required.

>15secs

If the power is off for 15 seconds or longer during a run, then the method currently running will still restart. If the power fails while executing a hold, or approaching a hold, then that hold temperature will restart from the beginning.

>5 minutes

Turn off each cycler at the front power switch. Remove plates from the cyclers and store them in the deli if the power is not restored in 5 minutes.

<2 hours

If building power is restored in 2 hours or less then put the plates back onto the same cyclers and resume the run by turning each cycler back on again at the front power switch. The operator should check all the cyclers affected to ensure the runs are continuing as required.

>2 hours

If building power is not restored within 2 hours the plates should be held for further testing.

NOTE: Any instrument that was idle prior to the power failure or did not automatically recover after an outage should be powered off and then back on after a 15 second pause.

PCR Network (Rm141) Recovery when Effected Runs are Completed

- Power off and then back on after a 15 second pause all the instruments in the network one at a time, wait for them all to reach the initial menu window.

- Re-boot the PCR net computer and restart the PCR Network application.

- Confirm that the PCR net software recognizes the units as they are returned to the system.

Work Request

If required submit an Instrumentation Work Request for any instruments continuing to have issues after the restoration of building power and shutting off & on again.

# SOP Approval

| DEPARTMENT | APPROVED BY | DATE |
| --- | --- | --- |
| Lab Supervisor |  |  |
| Research & Development |  |  |
| Instrumentation |  |  |
| QC |  |  |
| Purchasing |  |  |
| EH & S |  |  |
| Informatics |  |  |
| Seq Assessment & Analysis |  |  |
| Dept Head of Prod Seq |  |  |

# Appendix A

AUDIT TRACKING

5/27/04 – Duane Kubischta and Bill Feil performed an audit between 6:00AM – 12:00PM. The operator was David Robinson.

# Appendix B

|  |  |  | ABI Sequencing Chemistry- Reduced Buffer (0.5 BD rxn) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| When dispensing with Cavaro 3, prepare mix for 6 more plates than the number of plates in the batch. |
| Plate # | Water (ml) | BigDye (ml) | 5X Buffer (ml) | Primer (ul) | Total in ea. Mix (ml) |
| 1 | 1.05 | 0.19 | 0.29 | 6.1 | 1.534 |  |
| 2 | 2.10 | 0.38 | 0.58 | 12.3 | 3.069 |  |
| 3 | 3.14 | 0.58 | 0.86 | 18.4 | 4.603 |  |
| 4 | 4.19 | 0.77 | 1.15 | 24.6 | 6.138 |  |
| 5 | 5.24 | 0.96 | 1.44 | 30.7 | 7.672 |  |
| 6 | 6.29 | 1.15 | 1.73 | 36.9 | 9.207 |  |
| 7 | 7.34 | 1.34 | 2.02 | 43.0 | 10.741 |  |
| 8 | 8.39 | 1.54 | 2.30 | 49.2 | 12.276 |  |
| 9 | 9.43 | 1.73 | 2.59 | 55.3 | 13.810 |  |
| 10 | 10.48 | 1.92 | 2.88 | 61.4 | 15.345 |  |
| 11 | 11.53 | 2.11 | 3.17 | 67.6 | 16.879 |  |
| 12 | 12.58 | 2.30 | 3.46 | 73.7 | 18.414 |  |
| 13 | 13.63 | 2.50 | 3.74 | 79.9 | 19.948 |  |
| 14 | 14.68 | 2.69 | 4.03 | 86.0 | 21.482 |  |
| 15 | 15.72 | 2.88 | 4.32 | 92.2 | 23.017 |  |
| 16 | 16.77 | 3.07 | 4.61 | 98.3 | 24.551 |  |
| 17 | 17.82 | 3.26 | 4.90 | 104.4 | 26.086 |  |
| 18 | 18.87 | 3.46 | 5.18 | 110.6 | 27.620 |  |
| 19 | 19.92 | 3.65 | 5.47 | 116.7 | 29.155 |  |
| 20 | 20.97 | 3.84 | 5.76 | 122.9 | 30.689 |  |
| 21 | 22.01 | 4.03 | 6.05 | 129.0 | 32.224 |  |
| 22 | 23.06 | 4.22 | 6.34 | 135.2 | 33.758 |  |
| 23 | 24.11 | 4.42 | 6.62 | 141.3 | 35.293 |  |
| 24 | 25.16 | 4.61 | 6.91 | 147.5 | 36.827 |  |
| 25 | 26.21 | 4.80 | 7.20 | 153.6 | 38.362 |  |
| 26 | 27.26 | 4.99 | 7.49 | 159.7 | 39.896 |  |
| 27 | 28.30 | 5.18 | 7.78 | 165.9 | 41.431 |  |
| 28 | 29.35 | 5.38 | 8.06 | 172.0 | 42.965 |  |
| 29 | 30.40 | 5.57 | 8.35 | 178.2 | 44.499 |  |
| 30 | 31.45 | 5.76 | 8.64 | 184.3 | 46.034 |  |
| 31 | 32.50 | 5.95 | 8.93 | 190.5 | 47.568 |  |
| 32 | 33.55 | 6.14 | 9.22 | 196.6 | 49.103 |  |
| 33 | 34.59 | 6.34 | 9.50 | 202.8 | 50.637 |  |
| 34 | 35.64 | 6.53 | 9.79 | 208.9 | 52.172 |  |
| 35 | 36.69 | 6.72 | 10.08 | 215.0 | 53.706 |  |
| 36 | 37.74 | 6.91 | 10.37 | 221.2 | 55.241 |  |
| 37 | 38.79 | 7.10 | 10.66 | 227.3 | 56.775 |  |
| 38 | 39.84 | 7.30 | 10.94 | 233.5 | 58.310 |  |
| 39 | 40.88 | 7.49 | 11.23 | 239.6 | 59.844 |  |
| 40 | 41.93 | 7.68 | 11.52 | 245.8 | 61.379 |  |
| 41 | 42.98 | 7.87 | 11.81 | 251.9 | 62.913 |  |
| 42 | 44.03 | 8.06 | 12.10 | 258.0 | 64.447 |  |
| 43 | 45.08 | 8.26 | 12.38 | 264.2 | 65.982 |  |
| 44 | 46.13 | 8.45 | 12.67 | 270.3 | 67.516 |  |
| 45 | 47.17 | 8.64 | 12.96 | 276.5 | 69.051 |  |
| 46 | 48.22 | 8.83 | 13.25 | 282.6 | 70.585 |  |
| 47 | 49.27 | 9.02 | 13.54 | 288.8 | 72.120 |  |
| 48 | 50.32 | 9.22 | 13.82 | 294.9 | 73.654 |  |
| 49 | 51.37 | 9.41 | 14.11 | 301.1 | 75.189 |  |
| 50 | 52.42 | 9.60 | 14.40 | 307.2 | 76.723 |  |
| 51 | 53.46 | 9.79 | 14.69 | 313.3 | 78.258 |  |
| 52 | 54.51 | 9.98 | 14.98 | 319.5 | 79.792 |  |
| 53 | 55.56 | 10.18 | 15.26 | 325.6 | 81.327 |  |
| 54 | 56.61 | 10.37 | 15.55 | 331.8 | 82.861 |  |
| 55 | 57.66 | 10.56 | 15.84 | 337.9 | 84.396 |  |
| 56 | 58.71 | 10.75 | 16.13 | 344.1 | 85.930 |  |
| 57 | 59.75 | 10.94 | 16.42 | 350.2 | 87.464 |  |
| 58 | 60.80 | 11.14 | 16.70 | 356.4 | 88.999 |  |
| 59 | 61.85 | 11.33 | 16.99 | 362.5 | 90.533 |  |
| 60 | 62.90 | 11.52 | 17.28 | 368.6 | 92.068 |  |
| 61 | 63.95 | 11.71 | 17.57 | 374.8 | 93.602 |  |
| 62 | 65.00 | 11.90 | 17.86 | 380.9 | 95.137 |  |
| 63 | 66.04 | 12.10 | 18.14 | 387.1 | 96.671 |  |
| 64 | 67.09 | 12.29 | 18.43 | 393.2 | 98.206 |  |
| 65 | 68.14 | 12.48 | 18.72 | 399.4 | 99.740 |  |
| 66 | 69.19 | 12.67 | 19.01 | 405.5 | 101.275 |  |
| 67 | 70.24 | 12.86 | 19.30 | 411.6 | 102.809 |  |
| 68 | 71.29 | 13.06 | 19.58 | 417.8 | 104.344 |  |
| 69 | 72.33 | 13.25 | 19.87 | 423.9 | 105.878 |  |
| 70 | 73.38 | 13.44 | 20.16 | 430.1 | 107.412 |