Cycle Sequencing using BigDye v3.1

Performed on Fosmid DNA Template

Version Number:		1.3

Production Start Date:		02/01/04

Version 1.3 Date:		06/11/04

Authors:			Arshi Khan, Jamie Jett

Reviewed/Revised by:	Catherine Gordon, Megha Ambati

# Summary

The BigDye v3.1 sequencing kit is used to cycle sequence fosmid DNA, prepared via Agencourt SprintPrep or CosMCPrep, in preparation for sequencing on the ABI 3730xl instruments.

# Materials & Reagents

| Materials/Reagents/Equipment Disposables 384-well PCR Microplate, Clear 50-ml Centrifuge Tube 25-ml Serological Pipet 10-ml Serological Pipet PlateLoc Peelable Heat Seal LTS Pipette Tips (2, 10, 100, 200, 1000μl) 100-ml Reagent Reservoir Reagents BigDye Terminator v3.1 Cycle Sequencing Kit 5X BigDye Terminator v1.1/v3.1 Sequencing Buffer T7 Fosmid FW Sequencing Primer (250μm) 5’ TAA TAC GAC TCA CTA TAG GG 3’ T7 Fosmid RV Sequencing Primer (250μm) 5’ ATG ACC ATG ATT ACG CCA AG 3’ 100% Ethanol (200 Proof) Milli-Q Water Coulter Clenz Equipment 96/384-well Multidrop Micro Multidrop Micro Cassette Head 384-well Hydra Twister Hydra Wash Reservoir GeneAmp PCR System 9700 PlateLoc Thermal Plate Sealer Centrifuge 5810 Easypet | Vendor Axygen/ISC BioExpress Corning/VWR Falcon/VWR Falcon/VWR Velocity 11 Rainin Matrix Applied Biosystems Applied Biosystems IDT IDT AAPER Alcohol (LBNL) Millipore Milli-Q System - Thermo electron Thermo electron Robbins Scientific Robbins Scientific Perkin Elmer (Applied Biosystems) Velocity 11 Eppendorf Eppendorf | Stock Number T-3059-1CS 2533-50 357525 357551 06643-001 - 8086 4337458 4336701 - - 6810-13289 - - - - - - - - - - |
| --- | --- | --- |

# Calibration Check

Hydra Twisters

This should be performed on one FW and one RV plate within the first 10 plates per batch per Hydra Twister at step 2.12.

- Without disturbing the Hydra Twister, retrieve an empty FW destination plate from Stacker 4.

- Place the empty FW destination plate on the weighing balance and press the TARE button.

- Return the tared empty FW destination plate back to Stacker 4.

- After the tared FW destination plate has been aliquoted, remove it from Stacker 5 and replace it with an empty Axygen plate.

- Reweigh the aliquoted FW destination plate on the weighing balance.

NOTE 1: If the weight ranges fall outside of 1.463g – 1.55g, WARNING.

NOTE 2: If the weight ranges fall outside of 1.441g – 1.572g DO NOT USE.

- Enter the weight into the Plate Weight Excel Spreadsheet.

NOTE: To open the Excel spreadsheet, click on Shortcut to Plate Weight on desktop. If shortcut is unavailable, go to Network Places on desktop and click on Entire Network( click on Entire Contents( click on Microsoft Windows Network( click on jgi_psf( click on Octopus( click on ProdSeq( click on RCA(click on Calibrations (  click on Hydra Twister Calibration ( click on Plate WeightData( select Fosmid worksheet.

NOTE 2: If the weight entered in the excel spreadsheet is too close to the high or low range, the number entered in the weight column will be highlighted in a red box and the text WARNING will appear in orange under the Spec. column.   If this occurs, continue to use the Hydra but monitor it carefully so that it does not fall out of range.  If it is continuously high or low and in the WARNING zone, contact your supervisor to determine how to proceed.

NOTE 3:  If the weight entered in the excel spreadsheet falls out of the range, the number entered in the weight column will be highlighted in a red box and the text DO NOT USE will appear in a red box under the Spec. column.  If this occurs, discontinue use of the hydra and submit a work request for calibration.

- Remove the empty Axygen plate from Stacker 5 on the Hydra Twister and replace it with the aliquoted weighed FW destination plate.

- Repeat steps 1 – 7 with a RV destination plate.

- Click Save on the Excel spreadsheet.

Multidrop Micro

NOTE 1: This should be performed on the 1st plate per Multidrop Micro instrument per each Fosmid sequencing chemistry batch.

NOTE 2:  Make sure the Multidrop Micro is set for a 384 plate, 2µl volume, and 24 columns (col).

NOTE 3:  Refer to the Multidrop Micro in the Instrument Maintenance & Settings at the end of this protocol for proper Priming, Dispensing, and Purging techniques.

- Prior to dispensing any liquid, go to the balance and tare an empty 384-well Axygen plate.

- Dispense 2μl of Milli-Q water into each well.

- Visually check to make sure that Milli-Q water was dispensed into all 384 wells and that the well volumes are uniform.

- Return to the balance and obtain the new weight.

- If weights are all within the range continue to dispense all plates in batch. If not, set the cassette aside for the Instrumentation group and change to a new cassette; when changing to a new cassette, perform a calibration check on new cassette head by repeating steps 1 – 4.

NOTE:  Acceptable weight range after 2μl Milli-Q water addition: 0 .741g –0 .875 g.

- Record these values in the MD Calibrations Excel Spreadsheet, along with the serial number of the Multidrop Micro cassette head.

NOTE 1: To open the Excel spreadsheet, click on Shortcut to MD Calibrations on desktop. If shortcut is unavailable, go to Network Places on desktop and click on Entire Network( click on Entire Contents( click on Microsoft Windows Network( click on jgi_psf( click on Octopus( click on ProdSeq( click on RCA(click on MD Calibration(click on MD Calibrations( select Fosmid FW or Fosmid RV worksheet.

NOTE 2: If the weight entered in the excel spreadsheet is too close to the high or low range, the number entered in the weight column will be highlighted in a red box and the text WARNING will appear in orange under the Spec. column.   If this occurs, continue to use the cassette but monitor it carefully so that it does not fall out of range.  If it is continuously high or low and in the WARNING zone, contact your supervisor to determine how to proceed.

NOTE 3:  If the weight entered in the excel spreadsheet falls out of the range, the number entered in the weight column will be highlighted in a red box and the text DO NOT USE will appear in a red box under the Spec. column.  If this occurs, set the bad cassette aside for the Instrumentation group and change to a new cassette (after performing a calibration check on the new cassette).

# Procedure

NOTE 1:  Make sure to wear gloves throughout the entire Sequencing Chemistry process.

NOTE 2: All reagents/stock solutions should be prepared prior to the start of the procedure.

NOTE 3:  If you are resequencing plates, go to the Venonat database ( click on RCA DB ( click on Scan in RCA plates for Re-Sequencing only! ( enter in: operator and direction ( click on SUBMIT ( print the barcodes and affix to 384-well PCR plates.

- Sample Preparation

  - Remove the BigDye v3.1 Terminator Kit, respective FW/RV primers, & any leftover cocktail corresponding with respective primers from the -20(C freezer #202 located in Rm 140.

  - Let the reagents thaw on ice.

NOTE 1: The BigDye Kit usually takes 30 minutes to thaw, but may take up to two hours.

  - Remove the 5X buffer from the 4oC deli located in Rm140 and place it in the same ice bucket as BigDye reagents.

- Aliquoting the Fosmid Template

  - Take the FW and RV fosmid barcodes located on the shelf on the edge of the Sequencing Chemistry bay C (placed there by operators in the Libraries group who processed the fosmid DNA) and mark a regular green stripe next to the barcodes.

  - Affix the barcodes to new 384-well Axygen plates (on the indented side of the plate), stacking the RV direction of each set on top of the FW direction.

  - Remove the fosmid template plates from the bin labeled “Fosmids for Sequencing” from the 4oC deli in the Sequencing Preparation Rm140.

  - Unseal all the fosmid source plates.

  - Match the fosmid source plates to the correct destination plates (the bottom plate in the source stack should match the bottom FW/RV plates of the destination stack).

  - In the Venonat database, click on RCA DB ( STEP 4A - Scan RCA Plates to Update Aliquot Information ( enter in:

- Select the correct information from the drop-down menus: Hydra, Aliquot operator.

- Use the barcode scanner to scan in the fosmid source plates.

- Click on SUBMIT.

- Scan in the corresponding FW and RV plates and click on SUBMIT.

Checkpoint: If there are any mismatched plates correct the mismatches before loading these plates onto the Hydra Twister stackers.

2.5	Place the fosmid templates (source plates) in Stacker 1 and the destination plates (FW and RV pairs) in Stackers 3 and 4 (load stacker 4 up with a maximum of 32 plates before loading stacker 3). The top source plate should match the top destination plate in stacker 4. To distinguish Twister 2 plates from Twister 1 plates, add black dots to the FW and RV barcodes of Twister 2 plates.

NOTE: Press down the emergency stop button located next to the Hydra Twister to manually move the Twister arm.

2.6	At the beginning of each week, move the supply and drain Masterflex tubing to a new slot and write the date on the tubing.

NOTE: There are 4 slots on the tubing. When all slots have been filled once, submit a work request to have the tubing changed (approx. every month).

2.7	At the beginning of each day, log onto the Hydra Twister computers, then push down and twist up each emergency stop button located next to each Hydra Twister.

2.8	Open the Fosmid Autostart program on the monitor.

2.9	Run the WASH_100μlx3 program.

2.10	After completion of the WASH_100μlx3 program, exit the program.

NOTE: Each time a task is completed in the Aliquot FOSMID 1 program, close the program and re-open it.

2.11	Re-open Fosmid Autostart and click on Aliquot Axygen 384’s.

Checkpoint: The program will ask if there is enough water. Check the supply container and if it is less than half full, press the START button located above the supply container. If the water container is greater than half full, click OK on the program window. This must be done for every batch.

NOTE 1: This program will aliquot 4μl per well from the Fosmid source plates into each well of the corresponding FW and RV plates in preparation for sequencing.

NOTE 2: Periodically monitor the water trough to ensure water is filling and draining correctly. If water trough is not filling or draining properly or if the program was stopped abruptly, then manually fill or drain the trough by pushing the corresponding button towards manual on the Twister Hydra Pump Controller located next to the Hydra Twister. Make sure to switch the button back to automatic when finished. If the water trough is not filling or draining properly, also submit a work request.

2.12	Weigh one FW and one RV plate from each Hydra Twister (Refer to Hydra Twisters under the Calibration Check section).

2.13	After the program is complete and the pump has stopped running, close the window and push down and twist up the emergency stop button.

NOTE: This will move the robotic arm so that the fosmid source plates may be removed.

2.14	Remove source and destination plates from the Hydra-Twister stackers.

  - Re-open Fosmid Autostart and run the WASH_100μlx3 program.

NOTE: This program must be run in between each batch/run.

  - Visually check plates to make sure that all 384 wells contain fosmid template.

Checkpoint: If 4 or more wells are missing or overfilled per plate, re-aliquot the plate with the hydra. If there is something wrong with the tip, submit a work request. If the source plate is the source of the missing wells, see a supervisor to make a determination on how to proceed.

2.16	For the source fosmid plates:

- Place clear adhesive seals on all source fosmid plates.

- From each fosmid batch:

- Isolate 1 plate per operator per library per day.

  - Wrap the plates with labeling tape.

  - Label: Fosmids and date.

  - Put the plates in the 4oC deli located in Rm 140 in the QC Gels bin.

- Wrap the remaining plates in each batch with labeling tape.

- Label: Fosmids and date.

- Put the plates in the Fosmids bin located on the bottom shelf of the 4oC deli in Rm 140.

2.17	For the FW and RV plates:

    - Stack FW and RV plates into FW and RV sets.

b. Visually check destination plates to make sure that all 384 wells contain fosmid                                                                                                                                                                                               template.

Checkpoint: If 4 or more wells per plate are empty or overfilled, re-aliquot the plate manually with a Hydra. If there is something wrong with a single tip (e.g. clogged tip), submit a work request and have it repaired immediately. If the source plate is the cause of the missing wells, see the supervisor to make a determination on how to proceed.

c.   Place a clear adhesive seal on the top plate of each stack.

d.  Store the stacks in the 4oC deli located in Rm 140 until it is time to add the cocktail. NOTE: If the plates are aliquoted one hour or more before the cocktail will be added, add a clear adhesive seal to each plate in the stack before storing the plates in the deli.

2.18	At the end of the day, place a trough with Coulter Clenz cleaning reagent (located under the sink in room 140) in the destination area and run the Trough Wash program.

- Enter needle fill amount: 80%. Click OK.

- Enter the number of wash cycles: 5. Click OK.

- Enter the soak time: 60 sec. Click OK.

- Enter the number of bath wash cycles: 5. Click OK.

2.19      To shut down the Hydra Twister Robot for overnight storage:

- Click on End of the Day Fill to fill the needles with water.

- Turn off the robot by pushing down on the red emergency stop button located next to the Hydra Twister.

- Log off of the computer.

- Sequencing Chemistry

  - Label two 50-ml Corning tubes: FW and RV, respectively, and place the tubes on ice to chill. Affix a 2” piece of tape to the side of each tube, with the date, direction, primer type, primer lot#, BigDye lot#, and buffer lot#.

  - Prepare the Fosmid BigDye cocktail (Refer to the Reagent/Stock Preparation section) using the two 50-ml tubes (one tube for FW and one for RV).

NOTE 1: When dispensing with the Multidrop Micro, prepare 30% extra cocktail mix for batches of 19 source plates or less and 6 extra plates for batches of 20 or more source plates.

NOTE 2:  Make fresh Fosmid BigDye cocktail at the beginning of every week and continue to recycle the leftovers throughout the week.

  - Cap each FW and RV Fosmid BigDye cocktail tube, and mix by swirling gently and inverting each tube 3 times.

  - Install the Multidrop Micro cassette head labeled “Fosmid FW” into either MD Micro 1 or MD Micro 4 and install the Multidrop Micro cassette head labeled “Fosmid RV” into the remaining MD Micro (Refer to the Installing the cassette instructions under the Multidrop Micro section in Instrument Maintenance & Settings).

NOTE:  Fosmid BigDye Cocktail can be dispensed into both FW and RV plates simultaneously by using the 2 different Multidrop Micro instruments; one for the FW direction and one for the RV direction.

  - Wash the Multidrop Micro tubing by priming it with Milli-Q water (Refer to the Priming instructions under the Multidrop Micro section in Instrument Maintenance & Settings).

NOTE: Make sure to fill the trough 3 times when priming with Milli-Q water.  Every time the trough is nearly full, remember to empty its contents into a waste container.

  - Calibrate the Multidrop Micro (refer to Mutlidrop Micro under the Calibration Check section).

  - Empty the water in the microtubing back into the Milli-Q water reservoir.

  - Prime the Multidrop Micro containing the “Fosmid FW” cassette head with the FW primer BigDye cocktail and prime the Multidrop Micro containing the “Fosmid RV” cassette head with the RV primer BigDye cocktail.

NOTE: When dispensing FW cocktail, only use Multidrop Micro cassette head labeled as “Fosmid FW”; when dispensing RV cocktail, only use Multidrop Micro cassette head labeled as “Fosmid RV.”

  - Dispense 2μl of the cocktail into each well using the Multidrop Micro (Refer to the Dispensing instructions under the Multidrop Micro section in Instrument Maintenance & Settings).

NOTE 1: Monitor the 50-ml tube to ensure that there is enough cocktail for the plates and that the end of the Multidrop Micro tubing is submersed completely in the cocktail.

NOTE 2: If a clog develops or some other Multidrop Micro cassette failure occurs in the middle of dispensing, stop processing the batch, make a list of the plates in that batch that have not had liquid dispensed into them, and place a comment in the database regarding the new Multidrop Micro cassette’s unique identifying number. Continue processing the batch with a new cassette (after performing a calibration check on the new cassette). Note any changes in the MD Calibrations Excel Spreadsheet as well.

  - Visually check to make sure that the cocktail was dispensed into all 384 wells and that the well volumes are uniform.

Checkpoint: If 4 wells or less per plate are empty or low, manually add 2μl of cocktail to the empty wells with a pipette. If more than 4 wells are empty or overfilled, re-aliquot the plate with the hydra and run it through the Multidrop Micro again. If there is something wrong with a tip (e.g. clogged tip), submit a work request.

  - Place the plates in a centrifuge for a quick spin (1 minute at 1,000rpm).

  - Seal the 384-well plates with the Plateloc heat sealer.

  - Purge, clean, and remove the buffer cassette from the Multidrop Micro (Refer to the Purging, Cleaning cassette and trough, and Removing cassette for storage instructions under the Multidrop Micro Instrument Maintenance section).

NOTE: The Fosmid FW and Fosmid RV cassettes will be used daily until they fall out of calibration or a clog develops.

  - Save the leftover Fosmid BigDye cocktail after each batch by placing the tubes in the -20oC freezer.

  - In the Venonat database, click on RCA DB ( STEP 4B - Scan Plates for Cycle Sequencing:

  - Fill in the boxes: Dispense Operator, Dispense Instrument, FW Primer Lot, FW Kit Lot, RV Primer Lot, RV Kit Lot, Mix Operator, Days.

NOTE: If using different Multidrop Micro machines to dispense FW and RV cocktail, scan in FW and RV plates as separate batches. Select the appropriate Dispense Instrument when scanning plates in.

  - Scan in the plates according to which side of the thermocycler they will be loaded.

  - Click on SUBMIT.

  - Print this page and place it in the sequencing chemistry binder.

  - Load the plates on the correctly matched thermocyclers and close the lids.

  - Run the FOSMID program on the thermocyclers (refer to Instrument Maintenance & Settings for cycle program).

  - After the cycle sequencing is completed (approx. 6 hours), remove plates from thermocyclers and place them in bins labeled ABI Fosmids and date run.

  - Place the plates in the 4ºC fridge in Rm141 for the BET cleanup protocol.

# Reagent/Stock Preparation

BigDye v3.1 Cocktail, 0.5BD reaction for Plasmid Template

Per one 384-well Plate:

8.79 μl Milli-Q Water

316 μl BigDye v3.1

424 μl 5X Sequencing Buffer

19.2 μl Primer (250μM)

- When preparing fresh cocktail mix for dispensing with the Multidrop Micro, prepare 30% extra cocktail mix for batches of 19 source plates or less and 6 extra plates for batches of 20 or more source plates according to the cocktail volume per number of plates chart located in Sequencing Chemistry bay and also in Appendix A of this protocol. To calculate how much to prepare when using recycled Fosmid BigDye: Take the total volume you need and subtract the volume of leftover BigDye cocktail. Prepare cocktail based on the final volume.

Example: 20-source plate batch.

  - You need enough cocktail volume for 26 plates/FW primer and 26 plates/RV primer (total volume per primer = 19.968ml).

  - You have 5ml leftover BigDye cocktail.

  - You need to make 14.968ml fresh BigDye cocktail mix.

  - Prepare enough for a 20-source plate batch (total volume per primer = 15.360ml) and add to the leftover BigDye cocktail.

- Place two 50 ml corning tubes in a beaker filled with ice.  Affix a 2” piece of tape to the side of each tube with the date, direction, primer type, primer lot#, BigDye lot#, and buffer lot#.

- Add buffer, water, BigDye, primer, and any leftover cocktail to each 50-ml corning tube and mix by inverting the tube.

NOTE 1: Keep all reagents on ice while preparing cocktail and dispensing into the 384-well plates.

NOTE 2: If a tube of leftover cocktail has not been used within the last month, throw it out and make completely fresh mix.

70% Ethanol

- Pour 350 ml of 100% ethanol into a 500-ml bottle.

- Add 150 ml of Milli-Q water.

- Mix well by inverting bottle.

- Label bottle: 70% ethanol, date prepared, and your initials.

# Instrument Maintenance & Settings

Perkin Elmer 9700/Applied Biosystems 9700

Program FOSMID

95oC for 1 minute

95oC for 10 seconds

52oC for 5 seconds	  99 cycles

60oC for 2.5 minutes

Hold at 4oC

Volume = 6μl

Multidrop Micro

Switching on:

Switch the instrument on by using the main power switch on the left-hand side of the instrument. The green Power LED should light up.

Installing the cassette:

- Perform a visual check on the cassette and the cassette tubing to make sure that they are not damaged.

- Remove the tip protection from the dispensing cassette.

- Take the lower part of the cassette (with tips) into your right hand (the dispensing cassette tips should be pointing down) and the upper part in your left hand.

- Carefully place the eight tubes below the pump rotor and insert the lower part of the dispensing cassette into the lower part slots of the pump body. Check that the tubes are freely placed below the pump rotor and the tension limiting wires below the rotor shaft.

- Take a firm grip on the upper part of the dispensing cassette with your right hand and carefully pull the tubes around the pump rotor until the upper part reaches and fits into the upper part slots of the pump body.

- Check to make sure that both the upper and lower parts of the dispensing cassette are properly placed into their corresponding slots.

- Ensure that the tension limiting wires have a loose fit around the rotor shaft and tubes are evenly placed on the rotor needles (four tubes on each half of the pump rotor).

- Pull the rotor cover over the rotor.

- Remove the tubing weight protection.

Priming:

- After the cassette and priming trough are securely in place, place the tubing into a Milli-Q water reservoir.

NOTE 1: Fill fresh, clean, autoclaved bottles daily with Milli-Q water.

NOTE 2: Use separate wash bottles for FW and RV cassettes.

- Press the PRIME/DROP key until all eight channels are dispensing continuously into the priming vessel.

- Visually check to make sure that all of the tips are dispensing evenly.

- Press the EMPTY key until no water is left in the cassette tubing.

NOTE: Make sure to periodically empty the trough into a waste container.

- Place the tubing in the liquid that is to be dispensed and press the PRIME/DROP key until the liquid is dispensing evenly from all of the tips.

Dispensing:

- Check the control panel for proper selection of plate type (96 or 384 well plate), dispense volume, number of rows.

- Place the 384-well plate onto the plate adapter. Make sure that it is seated properly.

- Press the START key to begin dispensing.

- Observe that all tips are dispensing properly (i.e. no clogs, completely hitting inside the wells).

NOTE: If there is a clog in the tip, you must remove the cassette, submit a work request for the cassette, and install a new cassette.

- Remove the plate and perform a visual check to make sure that the Multidrop Micro has dispensed accurately and evenly across wells.

- Repeat steps 2 – 5 for each plate in the batch.

Purging:

- After all of the plates have been dispensed, press the EMPTY key to recycle any remaining liquid in the tubing.

- Remove the tubing from the reservoir.

Cleaning the cassette and trough:

- After you have finished dispensing, run Milli-Q water (trough full) through the tubing.

- At the end of the day, run Milli-Q water (trough full) followed by, 70% ethanol (trough full), and then followed by a separate bottle of Milli-Q water (trough full) through the tubing.

- Empty and wipe the trough with a Kimwipe to ensure that no liquid gets on the sensor.

Removing cassette for storage:

- Push the cover off of the rotor.

- Remove the cassette from the Multidrop Micro.

- Dry any liquid remaining on the cassette with a Kimwipe.

- Replace the tip protection cover on the cassette.

- Replace the tubing weight protection.

- Place the cassette back in its corresponding box.

Calibration Check:

Refer to the Calibration Check section on page 2.

# SOP Approval

| DEPARTMENT | APPROVED BY | DATE |
| --- | --- | --- |
| Lab Supervisor |  |  |
| Research & Development |  |  |
| Instrumentation |  |  |
| QC |  |  |
| Purchasing |  |  |
| EH & S |  |  |
| Informatics |  |  |
| Seq Assessment & Analysis |  |  |
| Dept Head of Prod Seq |  |  |

# Appendix A

1. BigDye v3.1 cocktail reagents per number of 384-well plates.

| ABI Sequencing Chemistry- Fosmid Sprint Prep |
| --- |
| When dispensing with the Multidrop Micro, prepare 30% extra mix for batches of 19 source plates or less and prepare mix for 6 extra plates for batches of 20 or more source plates. |
| No. of plates | Water (μl) | BigDye (ml) | 5X Buffer (ml) | Primer (μl) | Total in ea. Mix (ml) |
| 1 | 8.79 | 0.316 | 0.424 | 19.2 | 0.768 |
| 2 | 17.59 | 0.632 | 0.848 | 38.4 | 1.536 |
| 3 | 26.38 | 0.948 | 1.272 | 57.6 | 2.304 |
| 4 | 35.17 | 1.264 | 1.696 | 76.8 | 3.072 |
| 5 | 43.97 | 1.580 | 2.120 | 96.0 | 3.840 |
| 6 | 52.76 | 1.896 | 2.544 | 115.2 | 4.608 |
| 7 | 61.56 | 2.212 | 2.968 | 134.4 | 5.376 |
| 8 | 70.35 | 2.528 | 3.392 | 153.6 | 6.144 |
| 9 | 79.14 | 2.844 | 3.816 | 172.8 | 6.912 |
| 10 | 87.94 | 3.160 | 4.240 | 192.0 | 7.680 |
| 11 | 96.73 | 3.476 | 4.664 | 211.2 | 8.448 |
| 12 | 105.52 | 3.792 | 5.088 | 230.4 | 9.216 |
| 13 | 114.32 | 4.108 | 5.512 | 249.6 | 9.984 |
| 14 | 123.11 | 4.424 | 5.936 | 268.8 | 10.752 |
| 15 | 131.90 | 4.740 | 6.360 | 288.0 | 11.520 |
| 16 | 140.70 | 5.056 | 6.784 | 307.2 | 12.288 |
| 17 | 149.49 | 5.372 | 7.208 | 326.4 | 13.056 |
| 18 | 158.28 | 5.688 | 7.632 | 345.6 | 13.824 |
| 19 | 167.08 | 6.004 | 8.056 | 364.8 | 14.592 |
| 20 | 175.87 | 6.320 | 8.480 | 384.0 | 15.360 |
| 21 | 184.67 | 6.636 | 8.904 | 403.2 | 16.128 |
| 22 | 193.46 | 6.952 | 9.328 | 422.4 | 16.896 |
| 23 | 202.25 | 7.268 | 9.752 | 441.6 | 17.664 |
| 24 | 211.05 | 7.584 | 10.176 | 460.8 | 18.432 |
| 25 | 219.84 | 7.900 | 10.600 | 480.0 | 19.200 |
| 26 | 228.63 | 8.216 | 11.024 | 499.2 | 19.968 |
| 27 | 237.43 | 8.532 | 11.448 | 518.4 | 20.736 |
| 28 | 246.22 | 8.848 | 11.872 | 537.6 | 21.504 |
| 29 | 255.01 | 9.164 | 12.296 | 556.8 | 22.272 |
| 30 | 263.81 | 9.480 | 12.720 | 576.0 | 23.040 |
| 31 | 272.60 | 9.796 | 13.144 | 595.2 | 23.808 |
| 32 | 281.40 | 10.112 | 13.568 | 614.4 | 24.576 |
| 33 | 290.19 | 10.428 | 13.992 | 633.6 | 25.344 |
| 34 | 298.98 | 10.744 | 14.416 | 652.8 | 26.112 |
| 35 | 307.78 | 11.060 | 14.840 | 672.0 | 26.880 |
| 36 | 316.57 | 11.376 | 15.264 | 691.2 | 27.648 |
| 37 | 325.36 | 11.692 | 15.688 | 710.4 | 28.416 |
| 38 | 334.16 | 12.008 | 16.112 | 729.6 | 29.184 |
| 39 | 342.95 | 12.324 | 16.536 | 748.8 | 29.952 |
| No. of plates | Water (μl) | BigDye (ml) | 5X Buffer (ml) | Primer (μl) | Total in ea. Mix (ml) |
| 40 | 351.74 | 12.640 | 16.961 | 768.0 | 30.720 |
| 41 | 360.54 | 12.956 | 17.385 | 787.2 | 31.488 |
| 42 | 369.33 | 13.272 | 17.809 | 806.4 | 32.256 |
| 43 | 378.12 | 13.588 | 18.233 | 825.6 | 33.024 |
| 44 | 386.92 | 13.904 | 18.657 | 844.8 | 33.792 |
| 45 | 395.71 | 14.220 | 19.081 | 864.0 | 34.560 |
| 46 | 404.51 | 14.536 | 19.505 | 883.2 | 35.328 |
| 47 | 413.30 | 14.852 | 19.929 | 902.4 | 36.096 |
| 48 | 422.09 | 15.168 | 20.353 | 921.6 | 36.864 |
| 49 | 430.89 | 15.484 | 20.777 | 940.8 | 37.632 |
| 50 | 439.68 | 15.800 | 21.201 | 960.0 | 38.400 |

# Appendix B

AUDIT TRACKING

5/27/04 – Erika Lindquist and Sanjay Israni performed an audit between 7:16PM – 9:30PM. The operator was Megha Ambati.

PROCESS CHANGES

8/10/04 No longer taking individual well volumes per FW & RV plate per batch per Hydra Twister.

8/12/04 We started using 2 Multidrop Micros to dispense the chemistry simultaneously; one instrument for the FW direction & one for the RV direction.

8/11/04 Preparing fresh Fosmid BigDye cocktail at the beginning of every week, and recycling the leftovers throughout the week.

9/02/04 – Multidrop calibration weight range changed to 0.741 g- 0.875 g.

10/20/04 – HTS weight ranges have changed.  If weights fall outside 1.463 g– 1.55g it gets a Warning level, if it falls outside 1.441 g – 1.572g it gets a do not use level.