MegaBACE 4500 DNA Analysis System

Users Protocol

Version Number:		1.2 (Appended Ver2.4 of MB4000 User SOP)

Production Start Date:	11/11/04

Version 1.2 Date:		05/04/05

Author:			Chris Daum

Reviewed/Revised by:

# Summary

This protocol outlines the operation of the MegaBACE 4500 DNA Analysis System for high throughput capillary electrophoresis of DNA samples.

# Materials & Reagents

| Materials/Reagents/Equipment Disposables 384-well Nunc Plate 120μl, Clear 2.0ML Vial, Natural, PPCO Sterile, 1000/cs 2.0ML Vial, Blue, PPCO Sterile, 1000/cs Clear Adhesive Plate Sealers Reagents MegaBACE 4500 2.6% V2E LPA MegaBACE 10X Running Buffer Milli-Q water Equipment MegaBACE 4500 Sequencer Watson Marlow Pump Heat Block Vortex Centrifuge 5810 | Vendor VWR Fisher Scientific Fisher Scientific Edge Biosystems GE Healthcare* GE Healthcare* Millipore Super-Q System GE Healthcare* Watson Marlow VWR Scientific VWR Scientific Eppendorf | Stock Number 62409-074 342800-0020 642005-0680 48461 13575001 - - - - - - - |
| --- | --- | --- |

* GE Healthcare, formerly Amersham Biosciences

# EH&S

JGI employee performing this procedure must wear a lab coat and gloves.  In situation were there might be a chance of an accidental splash to the eyes, safety glasses must be worn.

# Procedure

NOTE: All reagents/stock solutions should be prepared prior to the start of the procedure.

1. Pre-Run Setup

  - Verify the run start time using the posted Run Schedule in Rm120.

  - Prior to the start time, assess how many instruments will be loaded and prepare the materials you will need for the run:

- Check the sample freezer in Rm121 to see if plates have been assigned for the run.  If no plates have been assigned, then pull plates from the sample freezer or BET Freezer201 in Rm140 and assign plates in FW & RV pairs to the same instrument for two runs.

- Ten minutes before loading sets of instruments, use the Watson-Marlow pump to fill the appropriate number of labeled translucent water tanks with 120ml of Milli-Q water and labeled yellow buffer tanks with 120ml of 1.5X running buffer.

i.	To operate Watson-Marlow pump, press the Start button twice to begin priming the tubing and dispensing 120ml per cycle.  To stop dispensing, press the Stop button once.

NOTE: The Watson-Marlow outlet tube is marked with lab tape.

Important! Run 240ml additional Milli-Q water through the pump after running 1.5X buffer through it in order to cleanse it and prevent buffer residue buildup.

Important! The dispense volume of the Watson-Marlow pump must be checked on a weekly basis. To do this, dispense into a graduated cylinder and measure the delivered volume; if it varies from 120ml + 5ml, notify MegaBACE instrumentation so that the pump may be calibrated.

NOTE: If the Watson-Marlow pump is non-functional, the water and buffer tanks may be manually filled using a graduated cylinder to measure out the 120ml dispense volumes.

- Set out a single buffer tank and water tank for each instrument in the loading set.

2. Sequencing Run

NOTE: A single sequencing run consists of loading sets of instruments. Refer to the posted Run Schedule for set information and start times. Make sure to fully complete the loading (matrix fill & sample injection) of one set before starting the next.

NOTE:  Refer to part 2 of Appendix B for an example of the loading scheme for instrument sets of 7.

  - Set out the matrix fill and pre-run materials you will need for each instrument in the set:

- One buffer tank containing 120ml of 1.5X running buffer per instrument in the set.

- Six pre-spun (2 minutes at 3,000rpm) tubes of 2.6% V2E LPA per instrument in the set.

  - Perform the Matrix Fill and Prerun Protocol for a set of instruments:

- In the Instrument Control window of the ICM software, with the Matrix Fill and Prerun protocol selected, click on START (perform this on each instrument in the set).

- Follow the instructions on the instrument display window to load the buffer tank into the left (cathode) side of each instrument followed by the matrix tubes into the right (anode) side of each instrument in the set; from the display window messages, verify that the drawers have been closed completely. Refer to Figures 1 & 2 in Appendix A.

Important!  Check the matrix tubes for bubbles and make sure there are no empty tubes; do not load these unsuitable tubes. If empty tubes are found, set them aside and label them with the LPA lot# and date found. If bubbles are still present in any of the tubes, re-centrifuge them for 2 minutes at 3,000rpm.

- After opening the anode and cathode drawers, wipe any spilled buffer, water or matrix with a damp paper towel.

- After loading the matrix tubes, listen for any high-pressure leaks. Inform MegaBACE instrumentation if any leaks occur.

  - Preparation of DNA samples for a set of instruments:

- After loading the matrix tubes and while the instrument is performing its automated fill and prerun, prepare the DNA sample plates for the set of instruments being loaded.

  - Pull the assigned sample plates from the freezer and thaw them on the heat blocks for 15 seconds at 90oC.

NOTE: Use provided timer for monitoring the thaw time of 15 seconds.

  - Vortex the sample plates on the Vortexer for 2 minutes on setting 7; loading plates in stacks of no more than 4 high.

NOTE: Ensure that plates are secure in Vortexer, but do not over-tighten the top cover screw-clamps or put excessive pressure on the top cover itself.

  - Quick spin the sample plates in the centrifuge for 1 minute at 1,000rpm.

Important! Once the sample plates have been centrifuged they must be handled gently as not to induce bubbles or produce air pockets at the bottom of wells.

- After preparing the DNA samples, the plates need to be scanned into the instrument on which they are to run; this will import the associated samplesheet record.

  - In the Plate Setup window of the ICM software, click on NEW.

  - In the Plate ID box, scan in the barcode on the sample plate that will be run on that instrument.

  - Verify that the correct values appear in the Electrophoresis Parameters: Sample Injection Voltage: 2 KV, Run Voltage: 5.9 KV, Sample Injection Time: 36 sec, Run Time: 215 min.

  - Perform the Inject Samples and Run Protocol for a set of instruments once the Matrix Fill and Prerun protocol has finished:

- Set out the materials you will need for each instrument in the set:

- One water tank containing 120ml of Milli-Q water.

- The prepared DNA sample plate assigned to each instrument.

- A 384-well NUNC Plate containing 85μl of 1.5X running buffer per well; place the plate on the Buffer Plate Adapter with well A1 located in the lower left-hand corner.

Important! Examine the buffer plate to ensure that each well has buffer in it and that the volume looks uniform across the plate.

- 6 pre-spun (1 minute at 1,000rpm) transparent blue micro-centrifuge tubes filled with 1.75ml of 1.5X running buffer.

Important! Examine the buffer tubes to ensure that each one has buffer in it, is bubble free, and that the volume looks uniform across all tubes.

- In the Instrument Control window of the ICM software, with the Inject Samples and Run protocol selected, click on START. The Select a Plate window will appear, scan in the barcode on the assigned sample plate to continue (perform this on each instrument in the set).

Important! It is crucial to sample performance that the following steps are completed as quickly as possible. Otherwise, the quality of the sequencing run may be compromised.

  - Follow the instructions on the instrument displays to load the water tank into the left (cathode) side of each instrument in the set; from the display window message, verify that the drawer has been closed completely. Refer to Figure 3 in Appendix A.

- Once the instruments finish their automated tip rinsing, the Confirm to Continue window will appear. The remaining steps will be performed on a subset of instruments at a time. Refer to part 2 of Appendix B for an example of the loading scheme of 7 instrument sets.

  - Click Enter on the Confirm to Continue window on the first instrument of the subset.

  - Follow the instructions on the instrument display to load the assigned sample plate on the adapter into the left (cathode) side of the first instrument in the subset; from the display window message, verify that the drawer has been closed completely. Refer to Figure 4 in Appendix A.

Important!  Ensure that the clear plate seal is completely removed and that no torn pieces of the seal remain on the plate. Also ensure that the plate is placed level on the adapter. Remaining pieces of seal or improperly placed plates may harm the instrument if loaded.

  - Repeat steps i. & ii. for each individual instrument remaining in the subset.

Important! Do not click Enter on all of the instruments in the subset at once; steps i. & ii. must be done individually for each one before moving on to the next.

  - Once the sample injection has completed on the first instrument in the subset, the Confirm to Continue window will appear.  Click Continue on the first instrument only.

  - Follow the instructions on the instrument displays to load the 85μl 1.5X buffer plate on the deep-well plate adapter into the left (cathode) side of the instrument followed by the blue 1.75ml 1.5X buffer tubes into the right (anode) side of the instrument. From the display window messages, verify that the drawers have been closed completely.

Important!  Examine the buffer plate and tubes before loading them into the instrument; ensure that the fill volumes are correct.

  - Repeat steps iv. & v. for each individual instrument remaining in the subset.

Important! Do not click Continue on all of the instruments in the subset at once; steps iv. & vi. must be done individually for each one before moving on to the next.

  - Repeat steps i. – vi. for all instruments in the second subset

  - Putting the instrument to sleep after the sequencing run (only to be done when instruments are to be slept for the night; refer to posted Run Schedule):

  - In the Instrument Control window, select the Sleep Parameters tab.

  - In the Sleep Time field, enter the number of hours the instrument is to sleep (1-16 hours).

  - Select the Sleep After This Run check box; the Instrument Control Manager will automatically store the capillaries in the run buffer tubes and buffer plate after the current run has finished.

Important! After the entered sleeping time has elapsed, it requires 3 full hours for the instrument to warm back up before it may be run again.

3. Monitoring and Assessing the Quality of the Run

  - Check the Fluorescence Image Display area to ensure that the scan images are being collected across the plate:

    - Click on the Run Image window of each instrument.

    - Click on the well buttons above the Fluorescence Image Display to view electropherograms for individual lanes. Initially, all 4 baselines should be present and at an intensity above zero. The lack of an image after 45 minutes (no sequencing data or baseline noise across all channels) may be indicative of instrument problems; notify MegaBACE instrumentation immediately.

  - Check the capillary current for uniform electrical current:

    - In the Options drop down menu select Current Monitor.

    - Examine the Current Monitor window for uniformity of current values across the plate; the values should be within a range of 6.5 – 8.5 μAmps. Check for regions of no current, abnormally high/low current, or erratic current that constantly changes values.

      - Individual zeros may indicate broken/clogged capillaries, empty buffer plate wells, or the presence of bubbles.

      - Groups of zeros or regions of abnormally high/low current may indicate more severe problems (i.e. broken capillaries or empty buffer tubes & empty buffer plate wells); notify MegaBACE instrumentation or administration immediately.

NOTE: Abnormal current values of individual capillaries may be referenced to the Bad Cap excel sheet located on the desktop of each instrument’s PC.

  - While monitoring MegaBACE runs, if any abnormal electropherograms are found which may include overloading, bubbles in the capillaries, sequencing artifacts, tandem repeats, or double traces, take a screen shot, paste it onto a PowerPoint slide (include: date, run #, instrument #, average intensity, the room’s temperature, and a brief description of the abnormal sequence) and email it.

  - While loading and monitoring runs in the MegaBACE Sequencing Lab, it is important to take note of and report any aberrations from the normal workflow. With as much detail as possible (documentation and screen shots if applicable) notify MegaBACE administration and instrumentation of the abnormality and send out an electronic depiction.  This would include any abnormalities, no matter how seemingly small, within the following areas:

    - MegaBACE Instruments – on-screen errors; software/hardware crashes; instrument malfunctions; poor current; failed or blinking lasers; difficulty in opening anode/cathode drawers; operator error while loading; aborting of runs.

    - Sample Plates – missing wells; damaged plates or barcodes; blank run images; run images with abnormal motifs.

    - Loading Runs – instruments down for service or PMs; evaporated buffer plates when pulling them out of instruments; loading sets at a different time than what is posted on the posted schedule; loading plates on the wrong MegaBACE instrument.

    - Venonat Database – outages; forms not working; scanning plates into the wrong MegaBACE field; choosing the wrong reagent lots, run number, or operator; plates that have not posted.

    - Reagents – changing of lots; empty or partially filled matrix tubes; buffer plates with empty wells; supply inventory; low on supplies.

    - Supporting Equipment & Robotics – issues with the Multidrop/Twister, Tecan, or Watson-Marlow pump; calibration checks; work requests.

    - Facilities – room temperature; running of the auxiliary AC; issues with the sink or Super-Q.

4. Post-Run

  - Scan in the barcodes of the sample plates that were loaded into the MegaBACE database:

- In the Venonat database, click on MegaBACE DB  Run Scan in Form MB4000s  enter in: run date, run #, your name as the operator, matrix lot, and buffer lot.

- Click on SUBMIT.

- Scan in the barcodes of the sample plate corresponding to the MegaBACE instrument on which it was loaded.

- Click on SUBMIT.

  - The loaded sample plates will be stored for a two-week period before disposal:

- After scanning the barcodes of the loaded sample plates into the Venonat database, tape the plates together and label with: Run Date, Run Number, Operator’s Initials.

- Store the plates in the Loaded Samples freezer in Rm121, placing them on the shelf that corresponds to the day of the week that they were loaded.

- To dispose of plates after a two-week storage period:

  - In the Venonat database, click on MegaBACE DB  Check Plate Status Before Disposing.

  - Scan the barcodes of the plates slated for disposal into the Enter Barcodes box.

  - Check the selection boxes next to Search for Yunian’s QC Reports and Display only runs that are not posted.

  - Click on SUBMIT.

  - An on-screen report will be generated, as long as a sample plate’s name is not displayed it may be thrown away in a lab garbage can. If a sample plate’s name is displayed, the plate has not yet posted. Save these plates and notify MegaBACE administration.

  - After loading the sets for a given run, the remaining water and buffer in the tanks used for the sequencing run must be disposed of and the tanks washed:

- Take the water tanks and buffer tanks to the sink.

- For Water Tanks:

- Dump and shake out the water into the sink and rinse the tanks with Milli-Q water only.

- Shake out the excess Milli-Q water and set the tank on the designated rack to dry.

- For Buffer Tanks:

  - Disposal of LPA/Buffer mixture – Turn on the sink’s hot tap water and allow it to run down the drain. Once the water is hot, dump and shake out each buffer tank into the sink.

  - Rinse each buffer tank with hot tap water to remove all of the remaining LPA/Buffer mixture and shake out the excess water. Once visible LPA strings no longer appear while shaking, rinse the tanks with Milli-Q water. Shake out the excess water and set the tank on the designated rack to dry.

  - After rinsing out the buffer tanks, allow the hot tap water to run for at least 2 minutes to prevent drain clogs from forming.

5. Aborting a Running Protocol

In the event of an instrument issue (i.e. on screen error messages, abnormal or extreme current indicating an empty buffer plate or matrix tube, laser shut downs, operator error, etc.), the following steps will abort a run:

  - In the Instrument Control window, click on STOP to end the running protocol. It will take the instrument ~30 seconds to abort the run and re-initialize itself.

NOTE: Should the Matrix Fill & Prerun protocol need to be aborted during the high-pressure fill, clicking STOP will not abort the protocol immediately; in order to abort immediately the instrument must be powered off using the power switch located on the right-lower side of the instrument.

  - If the issue is not serious in nature, the protocol may be restarted or the instrument may be left idle until the next run is ready to start. If there is not another run to start or the instrument requires service, it must be stored in water to prevent the capillaries from drying out (refer to Section 6).

  - If a run is aborted and not immediately restarted, the prepped sample plate that was assigned to run will need to be re-queued.

    - In the Venonat database click on MegaBACE DB  Scan in Reruns

    - From the drop down menu, select a reason for rerun

    - Scan the plate barcode(s) into the Rerun Plates field

    - Click on Submit, and return the plate(s) to the Rm121 sample freezer to be re-assigned

6. Restarting the Instrument & Rebooting the Computer

In the event of an instrument or software issue, the following steps will restart the instrument and reboot the computer and software:

  - While the instrument is stopped, close as many of the open software windows on the computer desktop as possible (some windows many not close if there is a software issue).

  - Turn the instrument’s power switch, located on the bottom right side, to the off position (a quick venting of the instrument’s pressure system may be heard).

  - Completely shut down the instrument’s computer and power it off.

  - Turn the instrument’s power switch to the on position, wait ~30 seconds for the system to initialize itself.

  - Once the instrument is initialized, turn the computer on and allow it to boot-up:

- While booting-up, the Notice to Users security screen will display, click OK and the computer will automatically log itself into the network (if the login is not automatic, refer to the User ID and Password posted on the top of the monitor to manually login).

- Once Microsoft Windows has started, the MegaBACE 4500 Host Scan Controller program will automatically start, allow it ~30 seconds to initialize.

  - Open the Instrument Control Manager program; its start icon is located in the MegaBACE folder on the computer’s desktop.

7. Leaving the Instrument Idle & Storing the Capillaries

Leaving the instrument idle means that the instrument power is on, but it is not running samples. If the instrument is left idle for 1 – 7 days, the capillaries need to be stored wet to prevent them from drying out. Leaving the instrument idle for more than 7 days requires the capillaries to be flushed and dried by MegaBACE instrumentation. The following steps will store the capillaries for a specified period:

  - Materials Required:

- Water tank with 120ml of Milli-Q water.

- Six clear 2mL micro-centrifuge tubes containing 1.75ml of Milli-Q water apiece.

  - In the Instrument Control window, select the Sleep Parameters tab; type the length of time the instrument will be idle in the Sleep Time field (1-168 hours).

  - In the Instrument Control window, with the Store Capillaries protocol selected, click on START. Follow the instructions on the instrument’s display window to load the water tank into the left (cathode) side of each instrument followed by the 6 water tubes into the right (anode) side of the instrument.

Important! After the entered sleeping time has elapsed, it requires 3 full hours for the instrument to warm back up before it may be run again.

8. Troubleshooting the MegaBACE 4500 DNA Analysis System

Refer to the MegaBACE 4000 DNA Analysis System Users Troubleshooting Guide as a reference for troubleshooting various issues that may be encountered while operating the MegaBACE 4500 instrument.

# Reagent/Stock Preparation

1.5X Running Buffer

- Add 300ml MegaBACE 10X Running Buffer to a 2-L bottle.

- Add 1.7L of Milli-Q water to the 2-L bottle.

- Mix well by inverting bottle.

- Label bottle: lot number, date prepared, and initials.

- Store in 4C deli.

Preparing the MegaBACE 4500 2.6% V2E Matrix (LPA)

  - Remove the matrix tubes from their heat-sealed tray that they were shipped in and transfer them into the gray tube-blocks.

NOTE: Save the tray and lid. Take them to the washroom (Rm 143) where they will be washed and sent back to GE Healthcare for re-use.

  - Centrifuge the matrix in the blocks for 2 minutes at 3,000rpm.

Important! After centrifuging the matrix tubes, check for any empty tubes or tubes that still have bubbles in them.  If empty tubes are found, set them aside and label them with the LPA lot# and date found.  If bubbles are still present in any of the tubes, re-centrifuge them for 2 minutes at 3,000rpm.

  - After centrifuging the matrix tubes, check for any empty or low level tubes. If any are found write the lot number and date found on the side of the tube, put them aside and notify MegaBACE administration.

  - Seal tubes with a clear plate seal and label: lot number, date prepared, and initials.

  - Store in 4C deli.

NOTE: Make sure to place the freshly made tubes in the back of the deli and bring the older tubes to the front.

Preparing the MegaBACE 384-well Nunc Buffer Plates

Refer to the Multidrop/Twister Operation Protocol for MegaBACE 4000 384-well Buffer Plates.

Preparing the 1.75ml 1.5X Buffer and 1.6ml Water Tubes used MegaBACE instruments

Refer to the Tecan Miniprep 60 Operation Protocol for MegaBACE Water & Buffer Tubes.

# SOP Approval

| DEPARTMENT | APPROVED BY | DATE |
| --- | --- | --- |
| Lab Supervisor |  |  |
| Research & Development |  |  |
| Instrumentation |  |  |
| QC |  |  |
| Purchasing |  |  |
| EH & S |  |  |
| Informatics |  |  |
| Seq Assessment & Analysis |  |  |
| Dept Head of Prod Seq |  |  |

# Appendix A

Figure 1. Buffer tank loaded onto the left (cathode) side of the MegaBACE 4500 instrument.

Figure 2. Six matrix tubes loaded onto the right (anode) side of the MegaBACE 4500 instrument.

Figure 3. Water tank loaded onto the left (cathode) side of the MegaBACE 4500 instrument.

Figure 4. Sample plate with adapter loaded onto the left (cathode) side of the MegaBACE 4500 instrument.

# Appendix B

MB4500 Loading Overview

1. Necessary Loading Reagents & Supplies:

1.1	1.5X running buffer

  - Diluted from 10X buffer stock – 300ml of 10X buffer & 1700ml of Super-Q water mixed in a 2-L bottle

  - Buffer Tank – 120ml of 1.5X run buffer dispensed by Watson-Marlow pump

  - Water Tank – 120ml of Super-Q water dispensed by Watson-Marlow pump

  - 2.6% V2E Matrix (LPA)

- Kits are located on the bottom left shelf of the deli in Rm121, the trays are individually sealed in foil bags

- Transfer tubes from trays to gray bases marked with orange tape, spin them down at 3000rpm for 2min.

  - 1.5X run buffer plate

- NUNC deep well plate filled with 85ul of 1.5X run buffer

- Prepare on Multidrop/Twister

    - In Configure Layout of Assist program change Plate Type to 384 NUNC Plate

    - Use MB4500_1.5X_buffer_plates_85ul program settings

    - Spin plates at 1000rpm for 1min

  - 1.5X run buffer tubes

  - 2ml Fisher Microcentrifuge tubes filled with 1.75ml of 1.5X run buffer, use gray bases marked with blue tape to put the tubes in.

  - Prepare on Tecan Miniprep 60

    - Use MB4500_1.75ml_Buffer_Tubes program settings

    - Spin tubes at 1000rpm for 1min

  - MB Sample Plate

- Injection Parameters: 2kV for 36sec

- Run Parameters: 5.9kv for 215min

2. Loading Scheme:

NOTE: A single sequencing run on the MB4500s consists of loading sets of 7 instruments in 3 parts.

Part 1: 7 instruments are started in a set according to Sections 2.1 through 2.4.b.i of the MB4500 Users SOP.

NOTE: Wait for all instruments in the set to finish their previous run prior to starting a new one.

Part 2: Sections 2.4.b.ii & iii of the SOP are then performed on a subset of 4 & 3 instruments respectively.

Part 3: Section 2.4.b.iv of the SOP is performed individually on each instrument of the subset.

Finally, Parts 2 & 3 are repeated on the 2nd subset of 3 instruments.

This is a descriptive outline of the MB4500 loading scheme:

  - Set out 120ml water reservoirs, 120ml 1.5X buffer reservoirs, and pre-spun V2E LPA for each of the 7 instruments in the Set.

  - Start the Matrix Fill and Prerun protocol on each of the 7 instruments:

- As prompted by the instrument display, load the 120ml 1.5X buffer tanks into the cathode of each instrument.

- As prompted by the instrument display, load 6 tubes of V2E LPA into the anode of each instrument.

  - Prepare the DNA samples assigned to the Set according to the MB4500 Users SOP, and scan them into their corresponding instrument.

- Sample Injection Parameters are 2kV for 36sec.

- Sample Run Parameters are 5.9kV for 215min.

  - Set out 85ul 1.5X NUNC run buffer plates and 1.75ml 1.5X run buffer tubes for each of the 7 instruments in the Set.  Start the Inject Samples and Run protocol on each of the 7 instruments:

- As prompted by the instrument display, load the 120ml water tanks into the cathode of each instrument.

- For the first 4 instrument in the set:

    - Click Enter on the Confirm to Continue window on the first instrument of the set and as prompted load the assigned sample plate on its adapter into the cathode; repeat this for the remaining 2 instruments in the subset.  Do not click Enter on all the instruments at once; this step must be done individually on each one.

    - Once sample injection has completed on the first instrument in the subset of 3, click Continue on the Confirm to Continue window and as prompted load the 85ul 1.5X run buffer plate and 1.75ml 1.5X run buffer tubes into the cathode and anode respectively; repeat this for the remaining 2 instruments in the subset.  Do not click Enter on all the instruments at once; this step must be done individually on each one.

- After fully completing steps 2.4.b.i & .ii on the first 4 instruments in the set, repeat these steps on the remaining 3 instruments of the 2nd subset.

# Appendix C

PROCEDURAL CHANGES

1/29/04 Stopping Blank Plates – Do not stop any more blank plates (no labeled DNA sample present), even if the plate will post with a very high failure rate. The reason for this policy change is that stopping blank plates skews overall performance results, making readlengths and pass rates appear higher. By letting these blank plates run and post, a more accurate representation of performance is gained and allows for better troubleshooting of the process line for areas that require improvement.

When a run that appears blank is encountered, continue the troubleshooting steps of checking the current monitor, HSC, and zooming in on the run electropherogram to see if the laser has cut off or if low intensity sequencing peaks may be seen. Additionally, continue to document these blank runs by taking screen shots of the run image and notifying the MegaBACE Maintenance Technician or Lead Operator; make note of it in your daily run log email as well.

4/6/2004 Lab Hygiene: Buffer Spills – Keep a damp towel handy while loading instruments or doing prep work in order to wipe up buffer spills as they occur; remember that wet buffer is much easier to clean-up than allowing it to dry which requires scrubbing.

5/17/2004 Saving HSC & HT Communication Data – Save the Host Scan Controller (HSC) and Hyper Terminal (HT) communication data from instruments that experience the following hardware/software issues:

  - HSC crashes (i.e. runs not starting/stopping, bad instrument status 51 & 52, and general loss of communication errors).

  - INTC/Board errors (i.e. CMON, TMPR, BEAM, ADAQ, SCAN, and EPHV board errors).

  - Any unusual/uncommon errors (i.e. instrument not going to sleep/HSC still acquiring data errors).

HSC communication data may be saved on any instrument when one of the above errors occurs. However, HT communication data may only be saved on those instruments with the program installed, which are MB115, 116, 122 - 135, & RNDMB1. Refer to the MegaBACE 4000 Comprehensive Users Protocol for instructions on saving HSC and HT communication data.

7/12/2004 Adhering to Run Schedules – Adhere to the posted MB run sets and loading times as noted in the daily Run Schedule, unless instructed otherwise by MegaBACE Administration or Instrumentation. Furthermore, completely load one set before moving on to the next, do not overlap sets (i.e. start the Matrix Fill protocol on one set while still performing the Inject Samples and Run protocol on the previous set). There are several reasons for the posted Run Schedule with set start times, instruments split up into loading sets, and breaks between some of the sets. The foremost of these are:

- Operator health and safety.

- Instrument well-being and prevention of catastrophic crashes/malfunctions.

- Prevention of operator error. Errors easily occur while rushing to load too many instruments at once or while loading overlapping sets, which can lead to harmful situations for both the operator and instrument.

- The ability to predict when instruments will be down, which is vital to the planning of instrument service and assigning priority samples.

- Minimization of variation. Introducing variation to any scientific/manufacturing process may lead to severe problems and be detrimental to performance; especially if the introduced variations are not properly noted and tracked.