Cycle Sequencing using BigDye v3.1

Performed on Rolling Circle Amplified DNA Template

1/8th Reaction (1.0BD)

Version Number:		2.1

Production Start Date:	10/01/03

Version 2.1 Date:		05/06/04

Authors:			Arshi Khan, Nancy Hammon, Catherine Gordon

Reviewed/Revised by:	Nahal Lalefar

# Summary

The BigDye v3.1 sequencing kit is used to cycle sequence plasmid DNA, which has been amplified by rolling circle amplification, in preparation for sequencing on the ABI 3730xl instruments.

# Materials & Reagents

| Materials/Reagents/Equipment Disposables 384-well PCR Microplate, Clear 125-ml Conical Tube 50-ml Centrifuge Tube PlateLoc Peelable Heat Seal LTS Pipette Tips (2, 10, 100, 200, 1000μl) Reagents BigDye Terminator v3.1 Cycle Sequencing Kit 5X BigDye Terminator v1.1/v3.1 Sequencing Buffer FW Sequencing Primer RV Sequencing Primer Coulter Clenz 100% Ethanol (200 Proof) Milli-Q Water Equipment 384-well Hydra Twister Hydra Wash Reservoir Cavro Dispensing Robot Pipettes (2, 10, 100, 200, 1000μl) GeneAmp PCR System 9700 PlateLoc Thermal Plate Sealer | Vendor Axygen/ISC BioExpress Falcon/VWR Corning/VWR Velocity 11 Rainin Applied Biosystems Applied Biosystems IDT IDT Beckman Coulter AAPER Alcohol (LBNL) Millipore Milli-Q System Robbins Scientific Robbins Scientific JGI Rainin Perkin Elmer (Applied Biosystems) Velocity 11 | Stock Number T-3059-1CS - 430290 06643-001 - 4337458 4336701 - - PN 8546930 6810-13289 - - - - - - - |
| --- | --- | --- |

# Calibration Check

Cavro 3

This should be performed prior to the dispensing run for each batch at Step 2.9.c.

- Make sure that the FW and RV tubing are in their respective Milli-Q water bottles.

- On the desktop located next to Cavro, double-click on the Cavro 3 Barcodes icon.

NOTE: Make sure only one program is open at a time.

- Under Select Stations, select BOTH.

- Under Select Process, select PRIME (Ensure that the feed tubes contain no air bubbles).

- Place an empty 384-well plate on the weighing balance and press the TARE button.

- Place the tared empty plate onto the bottom Cavro stage.

- Under Select Stations, select the BOTTOM station.

- Click on START.

  - Watch for droplets on the tips and missed wells on the plate.

  - Visually check the bottom of the plate to make sure that the volumes in all of the wells are uniform.

- Upon completion of dispensing, reweigh the plate to check the volume of water that is being dispensed and record the data in the Cavro notebook.

NOTE: If the weight ranges fall outside of 1.49g – 1.58g, submit a work request for calibration.

- Repeat steps 4 – 5.

- Place the tared empty plate into the top Cavro stage.

- Under Select Stations, select the TOP station.

- Repeat steps 8 – 9.

- Under Select Stations, select BOTH.

- Under Select Process, select RECYCLE.

Hydra Twisters

This should be performed on one FW and one RV plate within the first 10 plates per batch per Hydra Twister at step 1.11.

- Without disturbing the Hydra Twister, retrieve an empty FW destination plate from Stacker 4.

- Place the empty FW destination plate on the weighing balance and press the TARE button.

- Return the tared empty FW destination plate back to Stacker 4.

- After the tared FW destination plate has been aliquoted, remove it from Stacker 5 and replace it with an empty Axygen plate.

- Reweigh the aliquoted FW destination plate on the weighing balance.

NOTE: If the weight ranges fall outside of 0.3072g – 0.4608g, submit a work request for calibration.

- Enter the weight into the Plate Weight Excel Spreadsheet.

NOTE: To open excel spreadsheet click on Shortcut to Plate Weight on desktop. If shortcut is unavailable go to Network Places on desktop and click on Entire Network click on Entire Contents click on Microsoft Windows Network click on jgi_psf click on Octopus click on ProdSeq click on RCA click on Hydra Twister Calibration  click on Plate WeightData.

- Using the same aliquoted, weighed FW destination plate, measure the volume of a single well using the 2l pipette.

NOTE 1: If an accurate reading cannot be taken on the well, measure the volume of another well until an accurate reading can be taken. It may take up to 3 or 4 wells for an accurate reading.

NOTE 2: FW plates have a volume of approximately 1.2μl. RV plates have a volume of approximately 1.3-1.35μl.

- Enter the volume into the Plate Weight Excel Spreadsheet.

- Remove the empty Axygen plate from Stacker 5 on the Hydra Twister and replace it with the aliquoted weighed FW destination plate.

- Repeat steps 1 – 9 with a RV destination plate.

# Procedure

NOTE 1: Heat Inactivation (Step 5 of the Rolling Circle Amplification on Plasmid Glycerol Stock protocols) must be completed prior to aliquoting the sequencing template into the FW and RV plates.

NOTE 2: If you are resequencing plates, go to the Venonat database  click on RCA DB  click on Scan in RCA plates for Re-Sequencing only!  enter in: operator and direction  click on SUBMIT  print the barcodes and affix to 384-well PCR plates.

1. Aliquoting (Hydra Twister Robot)

NOTE: Make sure that the barcodes on the destination plates are marked with the correct primer association color.

  - Obtain the RCA source plates from the 4oC deli in the Sequencing Preparation room (Rm140) and remove the foil seals.

  - Match the source plates with the corresponding destination plates (i.e. The stacks of plates should be in the same order).

  - In the Venonat database, click on RCA DB  STEP 4A - Scan RCA Plates to Update Aliquot Information  enter in:

- Select the correct information from the drop-down menus: Hydra, Aliquot operator.

- Use the barcode scanner to scan in the RCA source plates.

- Click on SUBMIT.

- Scan in the corresponding FW and RV plates and click on SUBMIT.

Checkpoint: If there are any mismatched plates correct the mismatches before loading these plates onto the Hydra Twister stackers.

  - Place the RCA templates (source plates) in Stacker 1 and the destination plates (FW and RV pairs) in Stackers 3 and 4 (load stacker 4 up with a maximum of 32 plates before loading stacker 3). To distinguish Twister 2 plates from Twister 1 plates, add black dots to the FW and RV barcodes of Twister 2 plates.

  - At the beginning of each week, move the supply and drain Masterflex tubing to a new slot and write the date on the tubing.

NOTE: There are 4 slots on the tubing. When all slots have been filled twice, submit a work request to have the tubing changed (approx. every 2 months).

  - At the beginning of each day, open the Hydra Twister laptop monitors, log on, then push down and twist up each emergency stop button.

  - Open the Aliquot Autostart program on the laptop monitor.

  - Run the WASH_100μlx3 program.

  - After completion of the WASH_100μlx3 program, exit the program.

NOTE: Each time a task is completed in the Aliquot Autostart program, you must shut the program down and re-open it.

  - Re-open Aliquot Autostart and click on Aliquot Axygen 384’s.

Checkpoint: The program will ask if there is enough water. Check the supply container and if it is less than half full, press the START button located above the supply container. If the water container is greater than half full, click OK on the program window. This must be done for every batch.

NOTE 1: This program will aliquot 1μl per well from the RCA plates into each well of the corresponding FW and RV plates in preparation for sequencing.

NOTE 2: Periodically monitor the water trough to ensure water is filling and draining correctly. If water trough is not filling or draining properly or if the program was stopped abruptly then manually fill or drain trough by pushing the corresponding button towards manual on the Twister Hydra Pump Controller located next to the Hydra Twister. Make sure to switch the button back to automatic when finished.

  - Weigh one FW and one RV plate from each Hydra Twister (Refer to Hydra Twisters under the Calibration Check section).

  - After the program is complete and the pump has stopped running, close the window and push down and twist up the emergency stop button. This will move the robotic arm so that the RCA source plates may be removed.

  - Run the WASH_100μlx3 program in between each batch/run.

  - Visually check plates to make sure that all 384 wells contain RCA template.

Checkpoint: If 4 or more wells per plate are empty or overfilled, re-aliquot the plate manually with a Hydra. If there is something wrong with a tip (e.g. clogged tip), submit a work request. If the source plate is the cause of the missing wells, see the supervisor to make a determination on how to proceed.

  - For the source RCA plates:

- Place clear adhesive seals on all source RCA plates.

- From each RCA batch:

- Isolate 6 random plates.

  - Wrap the plates with labeling tape.

  - Label: RCA batch # and RCA date.

  - Put the plates in the 4oC deli in the QC Gels bin.

- Wrap the remaining plates in each batch with labeling tape.

- Label: ABI sequencing batch # and date.

- Put the plates in the 4oC deli on the correct shelf for that day.

  - For the FW and RV plates:

- Stack FW and RV plates into FW and RV sets.

- Place a clear adhesive seal on the top plate of each stack.

- Store the stacks in a 4oC deli until it is time to add the cocktail. If the plates are aliquoted one hour or more before the cocktail will be added, add a clear adhesive seal to each plate in the stack before storing the plates in the deli.

  - At the end of the day, place a trough with Coulter Clenz cleaning reagent in the destination area and run the Trough Wash program.

- Enter needle fill amount: 80%. Click OK.

- Enter the number of wash cycles: 5. Click OK.

- Enter the soak time: 60 sec. Click OK.

- Enter the number of bath wash cycles: 5. Click OK.

  - To shut down the Hydra Twister Robot for overnight storage:

- Click on End of the Day Fill to fill the needles with water.

- Turn off the robot.

- Log off of the laptop.

2. Sequencing Chemistry

  - Remove the BigDye v3.1 Terminator Kit, and FW/RV primers from the -20C freezer.

  - Remove the 5X buffer from the 4oC deli.

  - Let the reagents thaw on ice.

NOTE: The BigDye Kit usually takes 30 minutes to thaw but may take up to two hours.

  - Label two 250-ml bottles FW or RV, respectively, and place them on ice.

  - Prepare the BigDye cocktail (Refer to the Reagent/Stock Preparation section) using the two 250-ml bottles (one bottle for FW and one bottle for RV).

NOTE: When dispensing on Cavro 3 – make enough cocktail for 6 extra plates of FW and RV so the syringes don’t run dry. When dispensing by hand – make 20% extra cocktail for each.

  - Cap each FW and RV BigDye cocktail bottle, and mix by swirling gently and inverting each bottle 3 times.

  - Place two 125-ml conical tubes into two 600-ml beakers of ice labeled FW and RV.

  - Pour the FW cocktail into the 125-ml conical tube located in the FW beaker and pour the RV cocktail into the 125-ml conical tube located in the RV beaker.

  - Using Cavro 3, dispense 4μl of the FW and RV BigDye cocktail into each well of corresponding FW and RV 384-well plates:

  - Log onto the computer for Cavro 3.

  - Click on the Cavro 3 Barcodes icon located on the desktop.

  - Check the Cavro 3 calibration for each batch (Refer to the Calibration Check section).

  - Remove the FW Cavro tubing from the FW water bottle.

  - Replace the FW water bottle located on the FW Cavro stand with the FW beaker containing FW cocktail.

  - Place the FW tubing into the FW conical tube.

  - Make sure the FW stopper is at the bottom of the conical tube. If it is not, tape the glass rod of the Cavro tubing to the conical tube.

  - Repeat steps 2.9.d – 2.9.g for the RV Cavro and RV cocktail.

  - Under Select Process, select PRIME.

NOTE: Before priming, make sure that the troughs do not have water in them.

  - Dump the FW and RV trough contents back into corresponding FW and RV conical tubes of BigDye cocktail.

  - Remove the FW and RV plate stacks from the 4oC deli.

  - Place the FW or RV plate onto the corresponding FW or RV plate stand.

  - Click on START.

  - Upon completion of the dispensing run, take the plates off of the stands and visually check to make sure that the cocktail has been dispensed into each well.

Checkpoint: If 4 wells or less per plate are empty, manually add 4μl of cocktail to the empty wells with a pipette. If more than 4 wells are empty or overfilled, re-aliquot the plate with the hydra and run it through the Cavro again. If there is something wrong with a tip (e.g. clogged tip), submit a work request.

  - Place the plates on ice.

  - Add the next set of FW and RV plates to the FW or RV plate stand.

  - Click on the ENTER key on the keyboard.

  - Repeat steps 2.9.n – 2.9.q. until cocktail has been dispensed into all of the plates.

- Save the leftover BigDye cocktail after each batch:

  - Under Select Process, select PRIME.

  - Dump the remaining cocktail from FW and RV conical tubes and FW and RV troughs into corresponding FW and RV 50-ml tubes labeled with: the date, either FW or RV, batch number, primer type, and primer lot number.

  - Place the tubes in the -20oC freezer.

- Wash Cavro 3 after each batch.

  - Place the 250-ml water bottles back on their respective FW and RV stands.

  - Place the Cavro tubing back into the bottles.

  - Under Select Process, select PRIME and repeat 3 times.

NOTE: Empty the FW and RV troughs into waste containers after each wash.

  - Rinse out the FW and RV troughs by running them under Milli-Q water.

  - Add fresh Milli-Q water to each 250-ml bottle when the water levels get low.

  - Place the FW and RV troughs back into their slots.

- Place plates in a centrifuge for a quick spin (1 minute at 1,000rpm).

- Use the PlateLoc Thermal Plate Sealer to seal the plates with foil seals.

- In the Venonat database, click on RCA DB  STEP 4B - Scan Plates for Cycle Sequencing:

  - Fill in the boxes: Dispense Operator, Dispense Instrument, FW Primer Lot, FW Kit Lot, RV Primer Lot, RV Kit Lot, Mix Operator, Days.

  - Scan in the plates in FW and RV pairs according to which side of the thermocycler it will be loaded.

  - Click on SUBMIT.

  - Print this page and add it to the sequencing chemistry binder.

- Place the plates on the correctly matched thermocyclers.

- Run the BigDye program on the thermocyclers (refer to Instrument Maintenance & Settings for cycle program).

- After the program ends, remove the paired plates from the thermocyclers and place them in a bin labeled: ABI batch # and date.

- Place the plates in the 4oC fridge for the BET cleanup protocol.

- At the end of the day:

  - On Mondays, clean the Cavro by priming 70% Ethanol through the FW & RV tubes 3 times. Follow the ethanol wash with a water wash by priming Milli-Q water through the tubes 5 times.

  - Replace the two 250-ml water bottles on the FW and RV Cavro stands with clean bottles filled with fresh Milli-Q water.

  - Leave the Cavro tubing filled with water.

  - Close the Cavro 3 Barcodes program.

  - Log off of the computer.

# Reagent/Stock Preparation

BigDye v3.1 Cocktail, 1.0BD reaction for Plasmid Template (per one 384-well plate)

1.0 ml Milli-Q Water

400 μl BigDye version 3.1

200 μl 5X Sequencing Buffer

6.1  μl Primer (250μM)

- When preparing a fresh cocktail mix, prepare enough cocktail for the number of source plates in your sample plus 6 additional plates according to the cocktail volume per number of plates chart located in Sequencing Chemistry bay. To calculate how much to prepare: Take the total volume you need and subtract the volume of leftover BigDye cocktail. Prepare cocktail based on the final volume.

Example: 20-source plate batch.

  - You need enough cocktail volume for 26 plates/FW primer and 26 plates/RV primer (total volume per primer = 39.896ml).

  - You have 5ml leftover BigDye cocktail.

  - You need to make 34.896ml fresh BigDye cocktail mix.

  - Prepare enough for a 23-source plate batch (total volume per primer = 35.293ml) and add to the leftover BigDye cocktail.

- Add buffer, water, BigDye, primer, and any leftover cocktail to a 250-ml bottle and mix by swirling and inverting the bottle.

NOTE 1: Keep all reagents on ice while preparing cocktail and dispensing into the 384-well plates.

NOTE 2: If a tube of leftover cocktail has not been used within the last month, throw it out and make a completely fresh mix.

70% Ethanol

- Pour 350 ml of 100% ethanol into a 500-ml bottle.

- Add 150 ml of Milli-Q water.

- Mix well by inverting bottle.

- Label bottle: 70% ethanol, date prepared, and your initials.

# Instrument Maintenance & Settings

Perkin Elmer 9700/Applied Biosystems 9700

Program Big Dye

95oC for 1 minute

95oC for 30 seconds

50oC for 20 seconds	   25 cycles

60oC for 4 minutes

Hold at 4oC

Volume = 5μl

Cavro 3

Setup:

  - Log onto the computers located next to the Cavros.

  - Double-click on the Cavro 3 Barcodes icon.

Calibration:

Refer to the Calibration Check section on page 2.

Shut-down:

  - Close the Cavro 3 Barcodes program.

  - Log off the computer.