### Capillary Electrophoresis Process

### SOP: Production Capillary Electrophoresis on the MegaBACE 1000

Version Number:		4

Start Production Date:		3/02

Author:			Lisa Kegg/Metzalli Guel

Reviewed by:			Lyle Probst

SOP Change Document:

# Summary

To perform automated high throughput capillary electrophoresis using the MegaBACE 1000 Sequencing Instrument.

# Materials & Reagents

Materials/Reagents/Equipment			Vendor		Stock Number

##### Disposables

96 well cycle plate			Robbins Scientific		1055-90-0

96 well cycle plate			DOT-Sorenson			96MGB-50

Parafilm “M”				American National Cam

Clear Plate Sealers			Edge BioSystems		48461

##### Stock Solutions

Loading Buffer			Amersham			US79676

Labeled DNA samples		JGI – RCA Group

##### Reagents

Matrix (LPA)/Loading Buffer		Amersham			US79676

Milli-Q water				Milli-Q

##### Equipment

MegaBACE 1000			Molecular Dynamics

96 Plate Washer			Tecan

Plate Vortexer			VWR

Centrifuge 5417			Eppendorf

Centrifuge 5416			Eppendorf

96 well Heating Blocks		VWR

Ice Bucket				VWR

# Procedure

Prerun Set-up

Matrix (LPA) Preparation

- Place the Matrix (LPA) tubes in the gray PVC blocks designed for the Matrix tubes. If any matrix tubes are leaking or empty place them in the box marked “bad matrix”. These can be sent back to Amersham. Four blocks will fit into the centrifuge at one time.

- Centrifuge the Matrix (LPA) at 3000 rpm for 2 minutes.  Check the weight of the gray bases (written on the side of each) in order to balance them out in the centrifuge.

- Uncap the Matrix. The uncapping of Matrix should be divided evenly between all employees in the capillary group.  Each person should spread out the uncapping of their matrix through out their shift. Refer to breakdown of shift expectations posted in each lab.

- Seal with a plate sealer.

- Label with lot number and date uncapped.

- Place in 4 C refrigerator until use.

Sample Assigning and Set-up

- Note which MegaBACE’s are down.

- Select samples from the main sample freezer in Room 122. Priority samples should be selected first then production samples. Check the date and batch number and select the oldest production samples first.

- Assign a MegaBACE number to each sample and write it on the white part of the seal with a felt tip pen.  Forward and Reverses paired plates should be assigned to the same MegaBACE on consecutive runs.

- Place the samples back in the –20 freezer until they need to be prepared prior to loading the MegaBACE’s

Preparation of DNA Samples

- Samples can be prepared up to 30 minutes prior to a run.

- Remove the samples from the black bases.

- If the samples are frozen, place them on the VWR 96-well heating blocks for 30 seconds at a temperature of 90 C.

- Place up to 8 sequencing plates containing 10 uL of loading solution on the VWR Vortexer using the orange rubber separators between the samples.   Vortex on setting 6 for 2 minutes.

- Centrifuge up to 8 plates in the Beckmen at 1000 rpm for 1 minute.

- If a plate is dropped, centrifuge again at 1000 rpm for 1 minute. Check to make sure the volume is consistent in all wells.

- Place samples on the VWR 96-well heating blocks for 2 minutes at a temperature of 90 C.

- Place the samples on ice and cover with foil until loading.

MegaBACE Set-Up and Injection

1.	Check the sample to see if it has unusual run conditions (if nothing’s indicated on the sample, assume the normal run conditions apply and go to step 2). If you need to change the run conditions for any plate, do the following:

- On the “Plate Setup” screen of the ICM (Instrument Control Manager), select “Configure” from the top menu bar, and then click on “Edit mode.”

- Click on the “New” button near the bottom of the ICM, and type in the modified run conditions in the appropriate box(es).

- Scan in your sample, and if not automatically saved for you, click on the “Save” button near the bottom of the ICM.

2.	Scan the samples into the “Plate Setup” screen of the ICM on the appropriate MegaBACE and save.

3.	To start the run, go to the ICM and select the Instrument Control tab.

4.		Select matrix fill and prerun and press start.

5.	The LCD panel on the MegaBACE will direct the operator through the filling of LPA, pre-running, and injecting sample.

6.	While loading matrix tubes, buffer plates, and water troughs, be careful not to slam the drawers.  If drawers are hard to open/close most likely it is due to dried matrix and the drawer will need to be pulled out and wiped down with Milli-Q water by an instrumentation person.

7.	Clean the anode and cathode stages by wiping with a damp towel as needed.  Keeping the stages clean allows for smooth drawer movement and can prevent capillaries from sticking to the plates/stages/electrodes and breaking or bending.

8.	During the pre-run time the water troughs should be emptied, rinsed and filled with fresh Milli-Q water on each run.

- Electrophoresis will run for 160 minutes at 7kV.  Bring up the “Run Image” screen on each MegaBACE.  Also check the current monitor by going to the “Options” pull down menu and clicking on “Current Monitor.”  The current should be between 4 and 6 kV.

- After you have completed the loading process, scan your sample plates into the database in the Run scan in for MegaBACE 1000 to ensure that there will be a record of your running those samples.

- Use the main computer in Rm. 122 and scan your plates into the Excel “Run log.”

- Store your sample plates in the appropriate freezer (according to the room in which you loaded the samples) for a minimum of 5 days.

MegaBACE Problems

1.	If a machine experiences problems, report it to your Team Lead and change the MegaBACE status in the database when you scan in the plates.  Also, write on the MegaBACE itself, using the “Notepad” icon on the desktop, making sure to include the following information: date, time, brief description of problem, and your initials.

- If a machine goes down, store capillaries (SOP: Store Capillaries Protocol for MegaBACE 1000) in water on both anode and cathode sides. If a plate should get stuck or any time the anode or cathode is stuck in the down position, immediately notify a MegaBACE instrumentation member so that they remove the arrays before the matrix dries up. This will prevent clogged caps and allow us to continue the use of the arrays.

- Before bringing a machine back into production, make sure to check text messages on the computer screen to insure that the instrumentation person is finished working on it.

Instrument Parameters

Matrix Fill/High Pressure Time:  200 sec			Matrix Flush Time 1:  22 sec

Matrix Fill/Relaxation:  5 min				Matrix Flush Time 2:  7 sec

Prerun Time:  5 min					Low Pressure Time:  240 sec

Prerun Voltage:  10kV					User Input Time:  100 sec

Preinjection Voltage:  10kV				Preinjection Time:  15 sec

PMT1 Voltage:  750 V					PMT 2 Voltage:  750 V

Run Temperature:  44 C

Injection Conditions

Sample Injection Voltage: 1.6kV				Sample Injection Time: 36 seconds

Run Voltage: 7kV					Run Time: 160 minutes

Data Analysis

After completion of the run, a program will execute. This program automatically initiates base calling with Cimmaron 2.19.05 and then exports to .abd files. Once 96 files have been called, the program will rename the folder names, make contact with the FTP server and send all files. Once these are sent the raw and analyzed data files are moved from their folders and placed in the d: drive of the MegaBACE’s computer in a sent –raw/processed folder.

# Reagent/Stock Preparation

Buffer Preparation (1X)

- Measure out 2-50mL bottles of Running Buffer of the same lot number in a 1000mL graduated cylinder.  Final volume should be 100mL.  If volume is less than 100mL, adjust by adding additional buffer from the same lot number. If volume is greater than 100mL remove enough buffer to equal 100mL.

- Q.S. to 1000mL with Milli-Q water.

- Cover the graduated cylinder with parafilm and gently invert 5 times to mix completely.

- Fill bottles and store in refrigerator when not in use. Label bottles with lot number, date, time and initials of preparer.

- Prime the Tecan 96PW with buffer until buffer is being dispensed from all 96 tips.

- Change to the dispense setting on the Tecan 96PW.

- Check that the Tecan 96PW is set to dispense 200uL.

- Dispense loading buffer into a 96-well plate.

- Check that all 96 wells are being filled.

- Seal 96 well plate with a clear plate sealer.

- Repeat steps 10,11, and 12 as necessary.

- Centrifuge at 1000 rpm for 1 minute to remove bubbles.

- Sealed plates are stored in a 4 C fridge until ready for use.

- When finished with the Tecan 96PW prime (2-ten second cycles) and store with Milli-Q water.

- If the Tecan PW is not working correctly submit a work request.