- OFFICE OF RESEARCH OVERSIGHT

- GUIDANCE CHECKLIST FOR RESEARCH SAFETY

December 26, 2007

GENERAL INSTRUCTIONS:

The following checklist has been prepared as a self-assessment tool for the evaluation of VHA Research Safety Programs.  The items in this checklist can be attributed to various regulations and policies such as:

- 29 CFR 1910, Occupational Safety and Health Standards

- 29 CFR 1960, Basic Program Elements for Federal Employee Occupational Safety and Health Programs and Related Matters;

- VA Directive 7700, Occupational Safety and Health;

- VA Handbook 7700.1, Occupational Safety and Health Handbook;

- VHA Directive 7701, Occupational Safety and Health Program;

- VHA Handbook, 7701.1, Occupational Safety and Health Program Procedures;

- VHA Handbook 1200.06, Control of Hazardous Agents in VA Research Laboratories;

- VHA Handbook 1200.8, Safety of Personnel Engaged in Research;

- Biosafety in Microbiological and Biomedical Laboratories (BMBL), 5th edition;

- NIH Guidelines for Research Involving Recombinant DNA Molecules;

- 42 CFR 72 Possession, Use, and Transfer of Select Agents and Toxins; Final Rule (DHHS)

- 7 CFR 331 Agricultural Bioterrorism Protection Act of 2002 – Possession, Use, and Transfer of Biological Agents and Toxins; Final Rule (USDA: Plant and Plant Products);

- 9 CFR 121 Agricultural Bioterrorism Protection Act of 2002 – Possession, Use, and Transfer of Biological Agents and Toxins; Final Rule (USDA: Animal and Animal Products);

- NFPA 101 and

- Other applicable federal regulations, VHA Directives, VHA Handbooks, National Fire Protection Association (NFPA) standards, etc.

The checklist is comprehensive and intended to provide guidance for developing and monitoring research safety program compliance.  Facilities are encouraged to modify the checklist to address specific requirements for their individual research programs and are reminded that they are responsible for ensuring the veracity and appropriateness of all modifications.

- Always refer to the precise text of cited source documents for more specific requirements or when the meaning of any checklist element is unclear.

- The EPA regulations for hazardous waste must be considered in the context of individual state and local requirements, which may vary considerably.

- Each biosafety level includes the requirements for that level in addition to those for lower levels.

DOCUMENTS/SOURCE INFORMATION:

The following documents will assist with the self-review and are referenced by number in the “Documents/Source Information” column of the checklist for your convenience.

- Copies of last 2 years Annual Workplace Evaluations (AWE) conducted by the OSH Network Team (VHA Handbook 7700.1.5.a; 7700.1.11.e) – only information relating to research areas is needed.

- Copies of last 2 years safety inspection reports and hazard surveillance surveys (of research areas) conducted by the Facility Occupational Safety and Health Program (VHA Handbook 7700.1.6.a; 7700.1.11.f)

- Annual Vulnerability Assessment and any For-Cause Vulnerability Assessments of the Research Program conducted by Multi-Disciplinary Team

- Copy of any occupational injury and/or illness reports from the Research Laboratory to the Facility Safety and Occupational Health Management Office, if related to the Research Laboratories (information should be available through the ASISTS reporting system)

- Research Program Organizational Chart

- Research Safety Program and Research Service Safety Manual

- Medical Center Memoranda to establish various Subcommittees, if applicable

- R&D and SRS Committee Rosters with Membership Letters of Appointment (including terms)

- Minutes of the R&D, SRS, and other relevant Committees/Subcommittees for the last 12 months

- SRS Policies, Procedures, and Plans

- A list of all active protocols reviewed by the SRS for the last 12 months

- Safety Training Records for Investigators and all laboratory staff and any additional agent-specific training records, if applicable

- Emergency Plan/Incident Response Plan, training records, and related drills

- Research Chemical Hygiene Plan

- Hazardous Agents Control Program

- Copy of CDC/APHIS Certificate of Registration for Select Agents and Toxins, if applicable

- Biosafety Manual(s), if applicable

- BSL-3 Standard Operating Procedures/Manuals, if applicable

- Infection Control Plan

- Radiation Safety Program permits, if applicable

- Radiation Safety Manual, if applicable

- Copy of the inspection records conducted by external government agencies such as OIG, OSHA, CDC/APHIS and EPA, if applicable.

- All Applicable MOUs and CRADO approvals

Appendix A explains acronyms used in this document.

| A. RESEARCH SAFETY PROGRAMS |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Determine that the following items are appropriately and sufficiently implemented in your facility. Keep a completed version in the Research Service files. |  |  |  |  |  |  |  |
| 1. The VISN Safety Office conducts an Annual Workplace Evaluation (AWE) which includes: a. a comprehensive inspection using the Safety Automated Facility Evaluation (SAFE) software program; b. a program review; and c. a written report for the research program available for review by the facility. |  |  |  |  |  | VHA Directive 7701.3b; VHA Handbook 7701.1.11.e; 7701.1.11.f | 1 |
| 2. All research labs are included in the facility Occupational Safety and Health Program (OHSP) and they are inspected at least annually by the environmental rounds and research service personnel. |  |  |  |  |  | VHA Handbook 1200.8.3.a; 1200.8.4.a.2; 7701.1.11.e | 2 |
| 3. The Research Office maintains a Research Safety Program (RSP) that encompasses biological, chemical, and physical hazards in the research program and is administered by the Research Safety Coordinator (RSC - frequently the ACOS/R&D) who is appointed by the Research and Development Committee (R&D). |  |  |  |  |  | VHA Handbook 1200.8.3; 1200.8. 4.c.2.(d) | 5, 6, 7, 10 |
| 4. The Facility Director authorizes the RSC and appropriate research service supervisors to establish and implement the OSHA program in accordance with the recommendations of the corresponding AWE. |  |  |  |  | 29 CFR 1960.9 | VHA Handbook 7701.1.12.a.1; 1200.8.4.a.2 | 1 |
| ###### 5. Risks to visitors and other individuals with infrequent access to research hazards and/or areas (such as contractors, engineering, maintenance, and housekeeping personnel) are assessed and properly managed. |  |  |  |  | 29 CFR 1960.59(a) |  | 1, 2, 4, 6, 10, 22 |
| 6. The RSC and research service supervisory personnel are notified of unsafe working conditions in research areas and they are responsible for ensuring appropriate resolution or mitigation. |  |  |  |  | 29 CFR 1960.9 | VHA Handbook 7701.12.b.6 | 1, 2, 6, 9, 10 |
| 7. There is a Research Service safety manual (1200.8) and/or Safety Plan (1200.06) that includes biohazard controls and emergency procedures. - The Safety Manual is updated and approved annually by the Subcommittee on Research Safety (SRS) and R&D. - Individual labs adhere to all applicable federal safety and health requirements. - Non-exempt rDNA research adheres to the NIH Guidelines and is reviewed by a registered Institutional Biosafety Committee (IBC – can be an augmented SRS). - A Biosafety Officer (BSO) is appointed if non-exempt rDNA research involves Risk Group 3 agents or large scale production of viable rDNA molecules. - The Facility has a current permit from the VA Nuclear Health Physics Program (NHPP) for research involving radioactive compounds or ionizing radiation. - The Radiation Safety Officer (RSO) has developed written procedures (i.e., the Radiation Safety Plan) for use of radioactive materials in research laboratories. - The RSO and Radiation Safety Committee (RadSC) annually review and ensure the safe and effective use of ionizing radiation and radioactive materials used for research. - The RSO and RadSC have filed a written report with the NHPP concerning any incidents within the Research Service. - An emergency cascade phone list is included. - The Safety Plan includes regular (at least annual) safety inspections of the facility and drills or exercises conducted to test and evaluate the effectiveness of the Plan. |  |  |  |  | 10 CFR 20.1101; 29 CFR 1910.1450(c) | VHA Handbook 1105.1 3.k(5); 1105.1 3.m; 1105.1 4.a; 1105.1.4.1; 1200.06.7.a.9; 1200.06.7.h.4.(f); 1200.8.3.a.1; 1200.8.4.c.1; 1200.8.4.c.2.(a–e); Appendix E BMBL; NIH Guidelines Sec III, Sec IV-B-1-c; Sec IV-B-2-b Occupational Health Guidebook for Industrial Hygienist, Dept of Veterans Affairs, 1997, VA Directive 5019 | 6, 9, 10, 13, 20, 21, 22 |
| 8. High hazard areas of the medical center, including all research areas, are inspected quarterly by facility safety staff. In addition, all Research areas should be inspected annually by the Research Safety Coordinator and VISN staff. |  |  |  |  |  | VA Handbook 7700.1 6.a | 1, 2, 6, 10 |
| 9. There is a research chemical hygiene plan administered by the Research Chemical Hygiene Officer (CHO), who is appointed by the SRS. - Material Safety Data Sheets (MSDS) are available for all hazardous chemicals. - SRS reviews and approves chemical hygiene plan annually. - Facility Safety Officer/Industrial Hygienist reviews and approves a complete inventory of hazardous chemicals for each protocol prior to R&D review. - Chemical inventories are continuously updated, reviewed, and submitted to the CHO annually. - Chemical-specific training is provided as required by law (formaldehyde, ETO, etc.). |  |  |  |  | 29 CFR 1910.1450(e); 40 CFR 355 29 CFR 1910.0148 29 CFR 1910.1048 29 CFR 1910.1047 | VHA Handbook 1200.8. 3.a.2; 1200.8 4.4.3; 1200.8. 4.d.4; 1200.8. 4.d.14; 1200.8 4.e.1.(c); 1200.8 4.e.14.(b); 1200.06.7.g.1.(d) 29 CFR 1910.1450, Appendix A | 6, 9, 10, 11, 12, 14 |
| - All protocols using biological, chemical, physical, and/or radiation hazards are reviewed and approved by the SRS and R&D. There is a completed VA form 10-0398 for VA-funded protocols (or equivalent for non VA-funded protocols) on file. |  |  |  |  |  | VHA Handbook 1200.8. 4.c.2.(a); 1200.8. 4.d.1.(b); 1200.8.4.e.1 | 6, 9, 10, 11 |
| 11. The Facility Director (or designee) establishes clearance procedures for visitors and other individuals with an infrequent need to access research laboratory areas (chemical and biological agent disposal, laboratory cleaning, pest management services, security, etc.). |  |  |  |  | 29 CFR 1960.59(a) |  | 1, 2, 3, 6, 10, 13, 14, 16, 17, 18, 21, 22 |
| 12. A multidisciplinary team of local research personnel, a VA Police Service representative, the Facility Safety Officer, Safety Manager, RSC, Research CHO, Emergency Manager, RSO, and/or Industrial Hygienist conduct initial and annual vulnerability assessments of all labs. |  |  |  |  |  | VHA Handbook 1200.06.7.h.4.(j) 7700.1; 0730/1, paragraph 6. | 3, 6, 10 |
| 13. The same team conducts vulnerability assessments after any incident affecting high risk areas, materials, or physical security of research laboratories and reported to the SRS, R&D, and Responsible Official, if applicable. |  |  |  |  |  | VHA Handbook 1200.06.7.h.4.(j) 7700.1; 0730/1, paragraph 6. | 3, 6, 10 |
| 14. There is a Facility Emergency Preparedness Response Plan (EPRP) that includes the research program and which is: a. Reviewed annually and b. Practiced annually. |  |  |  |  |  | VHA Handbook 1200.06.7.a.9 1200.06.7.h 1200.06.7.h.4.(i) | 3, 6, 9, 10, 13 |
| 15. Biological fluids are handled using universal precautions according to bloodborne pathogen standards and in accordance with the Research facility’s Bloodborne Pathogen Exposure Control Plan. |  |  |  |  | 29 CFR 1910.1030(d) | VHA Handbook 1200.8.3.a.1.(a) | 6, 10, 19 |
| 16. All fume hoods, laminar flow chemical hoods, and biological safety cabinets (BSCs) are certified annually (semi-annually for airborne pathogens). Note: In addition to annual/semi-annual certification, portable BSCs must be recertified whenever they are moved. |  |  |  |  | 29 CFR 1910.1450(e)(3)(iii) | VHA Handbook 1200.8.4.e.5 29 CFR 1910.1450 Appendix A | 1, 2, 6, 10, |

| B. SUBCOMMITTEE ON RESEARCH SAFETY (SRS) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1. Committee Composition complies with VHA requirements. - At least 5 voting members nominated by the SRS and R&D Committee and appointed for a specified term by the Medical Center Director (MCD) including: - Chair – appointed by MCD for 1-year renewable term - Members with appropriate expertise in etiological agents, chemical hazards, and/or physical/radiation hazards, as applicable - Minimum of 2 non-affiliated members when non-exempt rDNA research is conducted - Biological Safety Officer (BSO), as applicable - Facility Safety Officer and IH, recommended as voting member(s) - Ex-Officio members including: - R&D Committee Liason (voting) - Research CHO (voting), as applicable - AO/R&D or R&D Office Representative (non-voting) - Employee Union Representative, as determined by the local union contract. |  |  |  |  |  | VHA Handbook 1200.8.4.d; VHA Handbook 1200.8. App. C | 6, 7, 8, 9, 10 |
| 2. Meets quarterly at minimum |  |  |  |  |  | VHA Handbook 1200.8.4.d.11 | 9, 10 |
| 3. Reviews and approves/disapproves all new research involving biological, chemical, physical, and radiation hazards at convened meetings with a quorum present |  |  |  |  |  | VHA Handbook 1200.8.4.d.1.(b) | 9, 10, 11 |
| 4. Annually reviews all active research protocols involving safety hazards |  |  |  |  |  | VHA Handbook 1200.8.4.d.3 | 9, 10, 11 |
| 5. Annually coordinates/provides research specific safety training for all laboratory personnel. |  |  |  |  | 29 CFR 1910.1450(f) | VHA Handbook 1200.8.4.d.5; 1200.8.4.d.10 | 9, 10, 11, 12 |
| 6. Ensures all deficiencies cited during safety inspections are properly abated. |  |  |  |  |  | VHA Handbook 1200.8.4.d.5.(a) | 1, 2, 9, 10 |
| 7. Reports operational problems and/or violations to RSC within 30 days, unless PI has done so. |  |  |  |  |  | VHA Handbook 1200.8.4.d.6 | 1, 2, 9, 10 |
| - Ensures proper health surveillance and exposure monitoring of personnel. |  |  |  |  | 29 CFR 1910.1450(c) | VHA Handbook 1200.8.4.d.7; 1200.8.4.d.13 | 1, 2, 4, 6, 9, 10, 11, 12, 19 |
| 9. Records and forwards minutes to the R&D for review and approval. |  |  |  |  |  | VHA Handbook 1200.8.4.d.8-9; 1200.8. App. D | 6, 9, 10 |
| 10. Use of an affiliate Safety Committee must be pre-approved by CRADO and include at least one qualified VA employee. |  |  |  |  |  | VHA Handbook 1200.8.4.c.1.(a-b) | 5, 6, 8, 23 |

| C. LABORATORY SAFETY - GENERAL |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1. The work area is free of unnecessary clutter. Fume hoods and BSC are not used for general storage and have vents/ductwork unobstructed. |  |  |  |  | 29 CFR 1910.1450(e)(3)(iii) | 29 CFR 1910.1450 Appendix A | 1, 2, 6, 10, 22 |
| ###### 2. Carpets and/or rugs are not used in labs or work areas where biological or chemical materials are handled. |  |  |  |  | 29 CFR 1910.1030(e)(3)(ii) |  | 1, 2, 6, 10, 14, 15, 17, 18, 22 |
| 3. For rooms equipped with sprinklers, all items on shelves have a minimum vertical clearance of 18-inches from sprinkler heads, heating pipes, and lighting fixtures. |  |  |  |  |  | National Fire Protection Association (NFPA) Standard 13 (NFPA 13-5.5.6) | 1, 2, 6, 10, 22 |
| 4. Functional fire extinguishers, appropriate for the hazard(s) present, are readily available and unobstructed. |  |  |  |  | 29 CFR 1910.157(c) & (d) | NFPA 10 (Standard for Portable Fire Extinguishers) NFPA 101-7.7.4.1 | 1, 2, 6, 10, 22 |
| 5. Personnel have access to emergency first aid services or kits. |  |  |  |  | 29 CFR 1910.151(a) & (b) |  | 1, 2, 6, 10, 22 |
| 6. Safety showers and eyewashes are located wherever hazardous agents are used. |  |  |  |  | 29 CFR 1910.151 (c) | BMBL BSL2.D.8; ABSL2.D.14 | 1, 2, 6, 10, 14, 17, 18, 22 |
| 7. Eyewash and safety showers are tested and results are documented in accordance with facility policy. |  |  |  |  | 29 CFR 1910.151 (c) | BMBL BSL2.D.8; ABSL2.D.14 | 1, 2, 6, 10, 14, 17, 18, 22 |
| 8. An occupational noise monitoring program is implemented when potentially hazardous noise exists. |  |  |  |  | 29 CFR 1910.95 |  | 1, 2, 6, 10, 22 |
| 9. Personnel who are required to use respirators have undergone medical evaluations and have been properly trained and fitted for their use. Personnel who are permitted to voluntarily use respirators are provided a copy of 29 CFR 1910.134 Appendix D |  |  |  |  | 29 CFR 1910.134(c), (e), (f), & (k) 29 CFR 1910.1450(i) |  | 1, 2, 6, 10, 12, 22 |
| 10. Use of extension cords is approved by RSC and minimized. |  |  |  |  | 29 CFR 1910.305(g)(1)(ii)(A); 29 CFR 1910.305(a)(2) |  | 1, 2, 6, 10, 22 |
| 11. Ground fault interrupter (GFI) electrical outlets are used in wet or high risk areas. |  |  |  |  |  | NEC (Article 210-8); NFPA 79; NFPA 70E | 1, 2, 6, 10, 22 |
| 12. Compressed gas cylinders are transported on cylinder carts and secured with chains or straps and capped when not in use. |  |  |  |  | 29 CFR 1910.101; | American Gas Association (AGA) Pamphlet | 1, 2, 6, 10, 22 |
| - Hazardous chemical storage is: - In properly identified and compatible containers and in a non-flammable storage room OR non-flammable cabinets; - Below eye level, but not on the floor; - Segregated according to compatibility (flammables separate from oxidizers, acids separate from bases, etc.); - Appropriate for flammable liquids (safety cans for volumes greater than 4-L and cabinets for volumes greater than 10-L); - Within appropriate cabinetry if flammable or corrosive; - In explosion-proof refrigerators when flammable cold storage is required; - Monitored for disposal prior to expiration – particularly for peroxide-forming reagents; - Examined periodically (at least annually) for replacement, deterioration, and container integrity. |  |  |  |  | 29 CFR 1910.106(d)(1)-(4); | NFPA 45 29 CFR 1910.1450 Appendix A | 1, 2, 6, 10, 14, 15, 22 |
| 14. All chemical and biological/medical lab wastes are collected and stored in compatible containers. Waste accumulation is minimized and limited to the lab area. Waste disposal is in accordance with federal and state regulations. |  |  |  |  | 40 CFR 262.43(c), subpart C (EPA) 29 CFR 1910.1030 (c)(4)(iii) (regulated biological/medical wastes) | 29 CFR 1910.1450 Appendix A | 1, 2, 6, 10, 14, 15, 22 |
| 15. Spill control programs are in place and may include spill kits. Training is provided when spill kits are available. |  |  |  |  |  | VHA Handbook 1200.06.7.h.4.b NIH Guidelines Sec IV-B-2-6-(6) Sec IV-B-3-c-(3) BMBL Appendix B | 1, 2, 6, 10, 12, 13, 14, 15, 19, 22 |
| 16. Controlled substances are properly secured and inventoried. |  |  |  |  | 21 CFR 1301.72-76 21 CFR 1304.11 | VHA Handbook 1108.1.14 | 1, 2, 6, 10, 22 |
| 17. Select carcinogens, reproductive toxins, and substances with a high degree of acute toxicity are handled in designated areas. |  |  |  |  | 29 CFR 1910.1450(e)(3)(viii) |  | 1, 2, 6, 10, 12, 14, 15, 16, 17, 18, 22 |
| 18. Eating, drinking, smoking, handling contact lenses, and/or applying cosmetics are prohibited in all labs and/or areas where chemicals are used. |  |  |  |  | 29 CFR 1910.1450(e)(3)(ii) | 29 CFR 1910.1450 Appendix A E.1.(d) | 1, 2, 6, 10, 14, 15, 22 |
| 19. A personal protective equipment (PPE) hazard assessment must be conducted for each lab. A copy of the certified hazard assessment must be maintained in each laboratory. All employees working in the lab must be trained in the PPE required for use in the lab. |  |  |  |  | 29 CFR 1910.132(d) |  | 6, 10, 11, 12, 14, 15, 17, 18, 19 |
| 20. PPE must be provided, used, and maintained in accordance with OSHA requirements and facility policies, (i.e., removal prior to leaving work area, cleaning and laundering, training, maintenance & care). |  |  |  |  | 29 CFR 1910.132(a) |  | 6, 10, 11, 12, 14, 15, 17, 18, 19 |

| D. BIOSAFETY (BSL) 1 AND ANIMAL BIOSAFETY LEVEL (ABSL) 1 LABORATORIES |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Standard Microbiological Practices |  |  |  |  |  |  |  |
| ###### 1. The Lab or Animal Facility Director (PI if nonexempt rDNA) controls access to labs or facilities in accordance with local institutional policies. |  |  |  |  |  | VHA Handbook 1200.06.4.k.2 BMBL Sec IV BSL-1 A.1; Sec IV BSL-1 A.11; Sec V ABSL-1 A.6; NIH Guidelines Appendix G-II-A-1-a | 3, 6, 10, 17 |
| ###### 2. Personnel are trained to follow safe work practices such as: ###### a. Hands are washed after work with cultures, glove removal, and before exiting the lab area. b. Eating, drinking, smoking, handling contact lenses, and/or applying cosmetics are prohibited in all labs. c. Mouth pipeting is prohibited and mechanical devices are used. d. Procedures minimize potential splashes or aerosols. e. Policies for the safe handling of sharps are enforced. f. Lab equipment and work surfaces are decontaminated at the end of the day and after any spill or splash. g. Laboratory cold storage is not used to store food items for personnel. |  |  |  |  |  | BMBL Sec IV BSL-1 A2; Sec IV BLS-1 A4; Sec IV BSL-1 A5; Sec IV BSL-1 A6; Sec IV BSL-1 A7; Sec V ABSL-1 A-8; Sec V ABSL-1 A-9; Sec V ABSL-1 A-10; Sec V ABSL-1 A-11; Sec V ABSL-1 A-12; Sec V ABSL-1 C.4; NIH Guideline Appendices G-II-A-1-b G-II-A-1-d G-ii-A-1-e G-II-A-1-f 29 CFR 1910.1450 Appendix A | 6, 10, 12, 17, 19 |
| ###### 3. Insect and rodent control programs are in place. |  |  |  |  |  | BMBL Sec IV BSL-1 A.10; Sec V ABSL-1 A.14; NIH Guidelines Appendix G-II-A-2-b | 1, 2, 6, 10, 17 |
| Safety Equipment (Primary Barriers) |  |  |  |  |  |  |  |
| ###### Lab coats, gowns, or uniforms are worn to minimize contamination of street clothes. No open-toed shoes or sandals are permitted. |  |  |  |  |  | BMBL Sec IV BSL-1 C.2; Sec V ABSL-1 C.2 29 CFR 1910.1450 Appendix A .E | 1, 2, 6, 10, 17 |

| ###### 5. Disposable gloves are available - specific types based on risk assessment. ###### Reusable gloves (utility, autoclave, cryogenic, etc.) may be needed and must be maintained in a clean and sanitary condition. |  |  |  |  |  | BMBL Sec IV BSL-1 C.4; Sec V ABSL-1 C.4 | 1, 2, 6, 10, 17 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ###### 6. Eye protection is provided and used for work that may involve splashes of hazardous materials. |  |  |  |  | 29 CFR 1910.1030 (c)(3) 29 CFR 1910.1450(e) 29 CFR 1910.132 and 133 | BMBL Sec IV BSL-1 C.3; Sec V ABSL-1 C.3 | 1, 2, 6, 10, 17 |
| Facilities (Secondary Barriers) |  |  |  |  |  |  |  |
| ###### 7. Labs have doors. |  |  |  |  |  | BMBL Sec IV BSL-1 D.1; Sec V ABSL-1 D.1 | 1, 2, 3, 6, 10, 17 |
| ###### 8. Labs/facilities have a sink available for hand-washing. |  |  |  |  | 29 CFR 1910.1030(d)(2)(iii) for BBP | BMBL Sec IV BSL-1 D.2; Sec V ABSL-1 D.2; Sec V ABSL-1 D.3; Sec V ABSL-1 D.4; Sec V ABSL-1 D.5 NIH Guidelines Appendix G-II-A-4-d | 1, 2, 6, 10, 17 |

| E. BIOSAFETY LEVEL (BSL) 2 AND ANIMAL BIOSAFETY LEVEL (ABSL) 2 LABORATORIES |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| All standard microbiological practices, safety equipment, and facility requirements for BSL/ABSL-1 apply in addition to the following: |  |  |  |  |  |  |  |
| Standard Microbiological Practices |  |  |  |  |  |  |  |
| ###### 1. Lab Director ensures personnel receive initial training on the potential hazards of specific agent(s) and annual updates or additional training as needed. |  |  |  |  |  | BMBL BSL-2 A.11; ABSL-2 A.3; NIH Guidelines Appendix G-II-B-2-m | 6, 10, 12, 13, 17, 19 |
| ###### 2. The Lab Director ensures lab personnel are proficient in applicable biosafety procedures. |  |  |  |  |  | BMBL BSL-2 B.5 | 6, 10, 12, 17 |
| ###### 3. Contaminated materials are placed in properly labeled containers that prevent leakage during collection, handling, processing, storage, transport, disposal, and shipping. |  |  |  |  | 29 CFR 1910,1030 (c)(4)(iii) for BBP | BMBL BSL-2 B.6; ABSL-2 A.15 | 1, 2, 6, 10, 17 |
| ###### 4. Potentially contaminated waste is decontaminated by an approved method before disposal. |  |  |  |  | 29 CFR 1910.1030(e)(1) | BMBL Sec IV BSL-1 A.8; Sec V ABSL-1 A.15; NIH Guidelines Appendix G-II-A-1-c 29 CFR 1910.1450 Appendix A | 1, 2, 6, 10, 17 |
| Safety Equipment (Primary Barriers) |  |  |  |  |  |  |  |
| ###### 5. BSCs are properly maintained and used for lab procedures involving: ###### a. Aerosols and/or splashes; or ###### b. High concentrations and/or large volumes of infectious agents. Note: Current certification is a maintenance indicator. |  |  |  |  |  | BMBL BSL-2 C.1; ABSL-2 C.1 NSF 49 | 1, 2, 6, 10, 17 |
| ###### 6. Appropriate personal protective equipment (PPE), as determined by risk assessment, is used and left in the laboratory area before exiting. |  |  |  |  | 29 CFR 1910.1450(e)(3)(iii) 29 CFR 1910.132 | BMBL BSL-2 C.2-5; ABSL-2 C.2-4; NIH Guidelines Appendix G-II-B-2-f | 1, 2, 6, 10, 17 |

| Facilities (Secondary Barriers) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ###### 7. When appropriate, biohazard signs are posted with the following information: ###### Agent(s) in Use ###### Biosafety Level ###### PPE and Immunization requirements ###### PI and Lab Supervisor’s Name with contact information ###### Entry and Exit procedures |  |  |  |  | 29 CFR 1910.1030(e) & (g) 29 CFR 1910.145 | BMBL Sec IV BSL-2 A.9; Sec V ABSL-2 A.5 | 1, 2, 6, 10, 13, 17 |
| ###### 8. Lab doors are self-closing with locks that conform to local institutional policies. |  |  |  |  |  | BMBL BSL-2 D.1; ABSL-2 D.1 NFPA 101-7.2.1.8.1 | 1, 2, 3, 6, 10, 17 |
| ###### 9. Lab is constructed and equipped to facilitate cleaning and decontamination. |  |  |  |  |  | BMBL BSL-2 D.3-4; ABSL-2 D.3-4 | 1, 2, 6, 10, 17 |
| ###### 10. BSCs are located away from doors, windows that open, and heavily traveled areas. |  |  |  |  |  | BMBL BSL-2 D.6; ABSL-2 D.11 | 1, 2, 6, 10, 17 |
| ###### 11. Adequate illumination is provided to minimize reflections and glare that impede vision. |  |  |  |  |  | BMBL ABSL2.D.10 | 1, 2, 6, 10, 17 |
| ###### 12. Bench tops are impervious to water and resistant to damage by moderate heat and/or chemicals. |  |  |  |  |  | BMBL BSL-2 D.4a.; ABSL-2 D.4; NIH Guidelines Appendix G-II-B-4-b | 1, 2, 6, 10, 17 |
| ###### 13. Lab furniture is covered with a non-fabric material that can easily be decontaminated. |  |  |  |  |  | BMBL BSL-2 D.4b; ABSL-2 D.4 | 1, 2, 6, 10, 17 |
| Lab Specific Practices |  |  |  |  |  |  |  |
| ###### 14. Lab Director (PI if nonexempt rDNA) is responsible for assessing lab-associated risk(s) and establishing entry/exit requirements for personnel. |  |  |  |  |  | BMBL BSL-2 B.1; ABSL-2 A.6; NIH Guidelines Appendix G-II-B-1-a; B-2-b; B-2-c. | 3, 6, 10, 17 |
| ###### 15. Lab Director incorporates biosafety procedures into lab SOPs or utilizes a lab-specific research biosafety manual. |  |  |  |  |  | BMBL BSL-2 B.4; ABSL-2 A.2; NIH Guidelines Appendix G-II-B-2-m | 6, 10, 17 |

| ###### 16. Lab personnel receive appropriate immunizations and/or clinical testing for agents of concern. ###### Note: Medical surveillance program is based on risk assessment. |  |  |  |  |  | BMBL BSL-2 B.2; ABSL-2 B.1 | 1, 2, 4, 6, 10, 17 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ###### 17. Baseline serum samples are collected and stored for all at-risk personnel, as appropriate. |  |  |  |  |  | BMBL BSL-2 B.3; ABSL-2 B.1; NIH Guidelines Appendix G-II-B-2-l | 1, 2, 4, 6, 10, 17 |
| ###### 18. The lab and lab equipment are routinely decontaminated in accordance with applicable local, state and/or federal regulations and articles are contained or decontaminated prior to removal from the facility. |  |  |  |  | 29 CFR 1910.1030(c)(2)(xiv) for BBP 29 CFR 1910.1450(e)(3)(viii) | BMBL BSL-2 B.7; ABSL-2 B.4 | 1, 2, 6, 10, 17 |
| ###### 19. Incidents involving infectious materials are: ###### Reported immediately to the Lab Director; ###### Followed by medical monitoring and appropriate care of personnel at risk; and ###### Documented by a written report or record. |  |  |  |  |  | BMBL BSL-2 B.8; ABSL-2 B.5; NIH Guidelines Appendix G-II-B-2-k | 1, 2, 4, 6, 10, 17 |
| ###### 20. Animals and plants not involved in research cannot enter the laboratory. |  |  |  |  |  | BMBL BSL-2 B.9; ABSL-2 A.13 | 1, 2, 6, 10, 17 |

| F. BIOSAFETY LEVEL (BSL) 3 AND ANIMAL BIOSAFETY LEVEL (ABSL) 3 LABORATORIES |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| All standard microbiological practices, safety equipment, and facility requirements for BSL/ABSL-1 and BSL/ABSL-2 apply in addition to the following: |  |  |  |  |  |  |  |  |  |  |
| Standard Microbiological Practices |  |  |  |  |  |  |  |  |  |  |
| ###### 1. Facility has method(s) to decontaminate all infectious lab waste - preferably within the lab. |  |  |  |  |  |  |  |  | BMBL BSL-3 D.11; ABSL-3 B.5; ABSL-3 D.12; NIH Guidelines Appendix G-II-C-4-h | 1, 2, 6, 10, 15, 17, 18 |
| ###### 2. Materials for decontamination outside the lab are packaged in accordance with applicable local, state, and federal regulations before removal. |  |  |  |  |  |  | DOT HM 181 (Hazardous Materials Regulations) |  | BMBL BSL-3 A.8b; NIH Guidelines Appendix G-II-C-2-b | 1, 2, 6, 10, 15, 17, 18 |
| Safety Equipment (Primary Barriers) |  |  |  |  |  |  |  |  |  |  |
| ###### 3. Standard PPE includes solid front lab coats, wrap-around gowns, or coveralls and reusable PPE is decontaminated before laundering. |  |  |  |  |  |  |  |  | BMBL BSL-3 C.2; ABSL-3 C.2 | 1, 2, 6, 10, 15, 17, 18 |
| ###### 4. Gloves are worn when handling infectious agents, animals, and/or contaminated equipment and disposed after use, when overtly contaminated, or when the integrity is compromised. |  |  |  |  |  |  |  |  | BMBL BSL-3 C.4; ABSL-3 C.4 | 1, 2, 6, 10, 15, 17, 18 |
| ###### 5. Properly selected and fit tested respiratory and/or face protection is used in rooms with infected animals. Employees wearing respiratory protection receive required training. |  |  |  |  |  |  |  |  | BMBL BSL-3 C.5; ABSL-3 C.3 | 1, 2, 6, 10, 15, 17, 18 |
| ###### 6. Manipulations of infected material(s) are conducted in Class II or III BSC. BSC or other physical containment devices are used for procedures that may produce aerosols. |  |  |  |  |  |  |  |  | BMBL BSL-3 C.1; ABSL-3 C.1; NIH Guidelines Appendix G-II-C-3-a | 1, 2, 6, 10, 15, 17, 18 |
| Facilities (Secondary Barriers) |  |  |  |  |  |  |  |  |  |  |
| ###### 7. Lab areas are separated from general building traffic and entered through an anteroom separated by 2 self-closing and locking doors. |  |  |  |  |  |  |  |  | BMBL BSL-3 D.1; ABSL-3 D.1 | 1, 2, 3, 6, 10, 15, 17, 18 |
| ###### 8. Each lab contains a hands-free sink for hand-washing near the exit. If the lab has segregated zones, each must have a sink. |  |  |  |  |  |  |  |  | BMBL BSL-3 D.2; ABSL-3 D.2; NIH Guidelines Appendix G-II-C-4-e | 1, 2, 6, 10, 15, 17, 18 |
| ###### 9. An eyewash station is available and located in the laboratory. |  |  |  |  |  |  |  |  | BMBL BSL-3 D.13; |  |
| ###### 10. The laboratory must be designed so that it can be easily cleaned and decontaminated. All windows are sealed. Carpets and rugs are not permitted. Seams, floors, walls, and ceiling surfaces should be sealed. Spaces around doors and ventilation openings should be capable of being sealed to facilitate space decontamination. |  |  |  |  |  |  |  |  | VHA Handbook 1200.06.7.e; BMBL BSL-3 D.5; ABSL-3 D.5; NIH Guidelines Appendix G-II-C-4-f | 1, 2, 3, 6, 10, 15, 17, 18 |
| ###### 11. Vacuum lines are protected by disinfectant traps and HEPA filters or equivalent. |  |  |  |  |  |  |  |  | BMBL BSL-3 D.7; NIH Guidelines Appendix G-II-C-2-o | 1, 2, 6, 10, 15, 17, 18 |
| ###### 12. Facility Management or Engineering Services documents and verifies BSL-3 facility design, procedures, and function prior to operation and annually at minimum. |  |  |  |  |  |  |  |  | BMBL BSL-3 D.15; ABSL-3 D.14 | 1, 2, 6, 10, 15, 17, 18 |
| ###### Lab Specific Practices |  |  |  |  |  |  |  |  |  |  |
| ###### 13. All personnel entering the lab are advised of potential hazards and meet specific entry/exit requirements (limited access). |  |  |  |  |  |  |  |  | BMBL BSL-3 B.1; ABSL-3 A.6; NIH Guidelines Appendix G-II-C-2-d | 3, 6, 10, 12, 15, 16, 17, 18 |
| ###### 14. All personnel receive appropriate medical evaluation, surveillance, and treatment as documented by written records. |  |  |  |  |  |  |  |  | BMBL BSL-3 B.2; ABSL-3 A.4 | 4, 6, 10, 13, 15, 17, 18, 19 |
| ###### 15. The Lab Director develops a lab-specific biosafety manual to address specific hazards, and ensures laboratory personnel have read and proficiently follow specific practices and procedures. |  |  |  |  |  |  |  |  | BMBL BSL-3 B.4; ABSL-3 A.2; NIH Guidelines Appendix G-II-C-2-s | 6, 10, 12, 13, 15, 16, 17, 18 |
| ###### 16. Equipment and work surfaces are disinfected regularly, after work w/agents, and/or after spills or contamination have occurred. |  |  |  |  |  |  |  |  | BMBL BSL-3 B.7; ABSL-3 A.12 | 6, 10, 15, 17, 18 |
| ###### 17. Spills and accidents involving infectious materials are immediately reported to the Lab Director, Facility Safety Officer, Project Officer, and Facility Director. |  |  |  |  |  |  |  |  | BMBL BSL-3 B.8; ABSL-3 B.7; NIH Guidelines Appendix G-II-C-2-q | 3, 4, 6, 10, 13, 15, 17, 18, 19 |
| ###### 18. The SRS reviews BSL-3 SOPs annually. |  |  |  |  |  |  |  |  | VHA Handbook 1200.06.7.a.8 | 6, 10, 15, 16, 17, 18, 22 |

| ABSL | Animal Biosafety Level |  | NEC | National Electrical Code |
| --- | --- | --- | --- | --- |
| ACOS/R&D | Associate Chief of Staff – Research and Development |  | NFPA | National Fire Prevention Association |
| AGA | American Gas Association |  | NHPP | Nuclear Health Physics Program |
| APHIS | Animal and Plant Health Inspection Service (USDA) |  | NIH | National Institutes of Health (DHHS) |
| ASISTS | Automated Safety Incident Surveillance and Tracking System |  | NSF | National Sanitation Foundation |
| AWE | Annual Workplace Evaluation |  | NRC | U.S. Nuclear Regulatory Commission |
| BMBL | Biosafety in Microbiological and Biomedical Laboratories, 5th Edition |  | OIG | Office of Inspector General |
| BSC | Biological Safety Cabinet |  | ORD | Office of Research and Development |
| BSL | Biosafety Level |  | OSH | Occupational Safety and Health |
| BSO | Biosafety Officer |  | OSHA | Occupational Safety and Health Administration |
| CFR | Code of Federal Regulations |  | OSHP | Occupational Safety and Health Program |
| CHO | Chemical Hygiene Officer |  | PI | Principal Investigator |
| CDC | Centers for Disease Control and Prevention (DHHS) |  | PPE | Personal Protective Equipment |
| CRADO | Chief of Research and Development Officer |  | RadSC | Radiation Safety Committee |
| DHHS | U.S. Department of Health and Human Services |  | R&D | Research and Development Committee |
| DOT | U.S. Department of Transportation |  | rDNA | recombinant DNA (deoxyribonucleic acid) |
| EPA | U.S. Environmental Protection Agency |  | RSC | Research Safety Coordinator |
| EPRP | Emergency Preparedness and Response Plan |  | RSO | Radiation Safety Officer |
| GFI | ground fault interrupter |  | RSP | Research Safety Program |
| HEPA | High Efficiency Particulate Air |  | SAFE | Safety Automated Facility Evaluation |
| IBC | Institutional Biosafety Committee |  | SOPs | Standard Operating Procedures |
| L | liter |  | SRS | Subcommittee on Research Safety |
| MCD | Medical Center Director |  | VA | U.S. Department of Veterans’ Affairs |
| MOU | Memorandum of Understanding |  | VHA | Veterans Health Administration |
| NHPP | Nuclear Health Physics Program |  | VISN | Veterans Integrated Service Network |
| MSDS | Material Safety Data Sheets |  | USDA | U.S. Department of Agriculture |