Validation

Flag   	Sub	Flag Description

A		Sampler adjustment or maintenance.

A1	Sampler audit during sample period.

A2	Sampler cleaned.

A3	Particle size cut device regreased or replaced.

B		Field Blank.

D		Sample dropped.

D1	Sample dropped after sampling.

D2	Filter dropped during unloading.

F		Filter damaged or ripped.

F1	Filter damaged in the field.

F2	Filter damaged when removed from holder.

F3	Filter wrinkled.

F4	Filter torn due to over-tightened filter holder.

F5	Teflon membrane separated from support ring.

F6	Pinholes in filter.

G		Filter deposit damaged.

G1	Deposit scratched or scraped, causing a thin line in the deposit.

G2	Deposit smudged, causing a large area of deposit to be displaced.

G3	Filter returned to lab with deposit side down in PetriSlide.

G4	Evidence that part of deposit has fallen  off filter.

G5	Finger touched filter in the field (without gloves).

G6	Finger touched filter in the lab (with gloves).

H		Filter holder assembly problem.

H1	Filter misaligned in holder - possible air leak.

H2	Filter holder loose in sampler - possible air leak.

H3	Filter holder not tightened sufficiently - possible air leak.

H4	Filter support grid upside down.

I		Inhomogeneous filter deposit.

I1	Evidence of impaction - deposit heavier in center of filter.

I2	Random areas of darker or lighter deposit on filter.

I3	Light colored deposit with dark specks.

Validation

Flag   	Sub	Flag Description

I4	Non-uniform deposit near edge - possible air leak.

I5	Discoloration on sample.

J		Filter misaligned, possible leakage.

K		(reserved)

L		(reserved)

M		Sampler malfunction.

M1	Dichotomous sampler assembled with virtual impactor 180 out of phase; only PM10 data reported.

N		Foreign substance on deposit.

N1	Insects on deposit, removed before analysis.

N2	Particles on deposit much larger than inlet device allows.

N3	Metallic particles observed on deposit.

O		Sampler operation error.

O1	Pump was not switched on after changing samples.

O2	Timer set incorrectly.

O3	Filter loaded in wrong port.

P		Power failure during sampling.

Q		Flow rate error.

Q1	Initial or final flow rate differed from nominal by > ±10%.

Q2	Final flow rate differed from initial by > ±10%.

Q3	Initial or final flow rate not recorded, used estimated flow rate.

Q4	Nominal flow rate assumed.

Q5	Initial flow readjusted prior to sampling.

Q6	Initial flow exceeded ±15% of nominal flow.

R		Replacement filter used.

R1	Filter that failed flow rate or QC checks replaced with spare.

R2	Filter sampling sequence changed from order designated on field data sheet.

Validation

Flag   	Sub	Flag Description

S

S1	Sample "suspect", but not invalid.

S2	Sample "suspect"; care should be taken when using data for statistical or modeling analysisi.

T		Sampling time error.

T1	Sampling duration error of > ±10%.

T2	Sample start time error of > ±10% of sample duration.

T3	Elapsed time meter reading not recorded or recorded incorrectly.  Sample duration estimated based on readings from previous or subsequent sample.

T4	Nominal sample duration assumed.

T5	Sample ran during prescribed period, plus part of next period.

T6	More than one sample was run to account for the prescribed period.

U		Unusual local particulate sources during sample period.

U1	Local construction activity.

U2	Forest fire or slash or field burning.

V		Void sample, Invalid.

W		Wet Sample.

X		No sample was taken this period, sample run was skipped.

Y		(reserved)

Z

Z1	Teflon and quartz filters were loaded reversely in SFS.

Z2	PM2.5 and PM10 filter pack switched.

Z3	Fine and Coarse filters were loaded reversely in dichotomous sampler.

a	Flagged samples are categorized as valid, suspect, or invalid.  Any flags except 'S1', 'S2' and 'V' or no flags at all indicate valid samples.  The 'S1' or 'S2' flag in combination with any other flags indicates samples of suspect validity.  The 'V' flag in combination with any other flags indicates invalid samples.  Notes from Field Data Sheets not covered by the flags are recorded in the memo field of the database.