Thermal Shroud Specifications

Thermal shroud shall…

- be capable of being used with an existing water-based chiller system operating in a 10-30 C temperature range.  Chiller is Lydall R-series chiller model # RAA-003B-BE01CBD3.  Water flow rate up to 1.6 gpm at 30 psi.

- fit into a 48” diameter vacuum chamber.

- be 6’ long, without end caps.

- have an assembled inner diameter at least 42” in diameter.

- be made in 3 separate pieces to allow it to fit into chamber.

- have all three cooling tubes connected to form a single cooling loop.

- have all cooling tube fittings accessible so they can be connected to each other after installation into the vacuum chamber.

- have the cooling tubes on the exterior of the shroud (facing out)

- be made from aluminum

- be painted or anodized black

- rest on the existing standoffs in the chamber.  Chamber drawing is attached.

- use ½” Swagelok fittings on the ends of the tubes

The goal is to maintain an isothermal environment in the vacuum chamber to within 0.1 degrees C per hour.  The heat input into the chamber by the test articles is minimal, less than 10 watts intermittently.