Additional Technical Information Supporting the Sole Source Procurement of a 40 inch Diameter PTFE Integrating Sphere from Sphere Optics:

It is recommended that the 40 inch diameter PTFE Integrating sphere be purchased from Sphere Optics Corporation.  Specifications for the Sphere include the following:

1. The sphere must be a 40 inch diameter aluminum shell lined with polytetrafluoroethylene (PTFE). The 40 inch sphere diameter is needed to accommodate higher operating temperatures, foregoing the need to cool the source and to provide a stable output with shorter warmup period.  The aluminum shell is required to enable the sphere to reach equilibrium color temperature stability in less than 1 hour at 412 nm.  PTFE is required due to its high reflectance across the ultraviolet, visible, near infrared and shortwave infrared wavelength regions, its durability, and its non-hygroscopic properties. Provided it is kept clean, PTFE does not require replacement or recoating. In the event of a contamination event, PTFE can be easily cleaned.

2. The non-uniformity of the output of the sphere must be 1% or less over the output aperture. Moreover, given our uniformity specifications and in order to accommodate facility customers with instruments with wide field of vision, a spherical PTFE interior which conforms to the shape of the interior spherical aluminum shell, is essential. Based on our laboratory measurements, approaches which attempt to approximate a spherical interior using a series of flat PTFE panels are not acceptable.

3. The lamp ports must be of modular design to allow conversion to lamps of other wattage and types (e.g. from tungsten halogen to xenon).

4. Long term (i.e. hours after warmup) radiance stability of the sphere must be equal to or better than 0.2% at 412 nm.

5. The sphere must provide radiance output equal or greater than 2500 W/m2 sr μm at its peak wavelength output (i.e. at an approx. wavelength of 1 μm).

6. The sphere must include an external lamp housing with an attenuator, to allow near-continuous radiant output between sphere lamp radiance levels.