November 19, 2002

Item:  	4341B Thoracentesis Catheters in Thoracentesis Trays Distributed by Allegiance Healthcare Corporation

General Information:  	ECRI, a medical device safety organization, issued a critical alert regarding the potential threat caused by some thoracic catheters in Pharmaseal Thoracentesis Trays distributed by Allegiance Healthcare Corporation.  When contacted by the National Center for Patient Safety, Allegiance Healthcare provided a recall notice (see Attachment 1).  They also confirmed that at least 28 VA medical centers are using this product.

Specific Information:	The catheters in the following lot numbers may be brittle creating the potential for breakage during use. These catheters, which are used to aspirate fluid from the lungs, are only exposed after they are inserted into the patient; therefore they cannot be examined for flexibility or strength prior to insertion.  Possible complications resulting from fragmented catheters are foreign-body reactions, infection, and the need for surgery to remove the fragments.

| Catalog No. | Lot Numbers |
| --- | --- |
| 4341B | L1J035X L1J088 L1K042 L1K058 L1K070 L1L057X L1N018 L1N031 L1N039X L1N094 L1P017 L1P045 L1S015 L2A012 L2A021 L2A048 L2A075 L2B039 L2B047 L2B073 L2C012 L2C038 L2C083 L2D041 L2D067 L2E036 |

Action:	1.  Remove all stock with the identified lot numbers from the inventory.

2.  Follow product return instructions provided on Attachment 1 and Attachment 2.

Contact:	Joe DeRosier or Bryanne Patail at the VA National Center for Patient Safety, (734) 930-5890.

Attachment 1 of 2

Recall Notice from Allegiance Healthcare Corporation

## November 15, 2002

# Product Description:  4341B THORACENTESIS TRAY WITH CATHETER

| Catalog Number | Lot Numbers |
| --- | --- |
| 4341B | L1J035X L1J088 L1K042 L1K058 L1K070 L1L057X L1N018 L1N031 L1N039X L1N094 L1P017 L1P045 L1S015 L2A012 L2A021 L2A048 L2A075 L2B039 L2B047 L2B073 L2C012 L2C038 L2C083 L2D041 L2D067 L2E036 |

Dear Valued Customer,

Allegiance Healthcare has identified that a number of the 4341B thoracentesis catheters, within the above listed thoracentesis tray, in the Lot Numbers listed above, may be brittle, thus creating the potential for breakage during use.  In our effort to provide the highest level of quality products to our customers, Allegiance Healthcare has issued a voluntary recall of these products.  The Food and Drug Administration has been made aware of our action.

Our records indicate that your facility may have purchased the above referenced products.  Please search your inventory and follow the steps outlined below.

STEP 1	Inspect your inventory and immediately segregate any affected product containing the above referenced lot numbers.

STEP 2	For product return and instructions on how to obtain a no cost replacement please contact Allegiance Healthcare Customer Service.  	If you are a Hospital, call:  800.964.5227

All other customers call:    888.444.5440

To avoid a stock out situation, only product with the affected lot numbers will be replaced.

STEP3	Please fax (847.689.9101) the attached Acknowledgment Form regardless of if you have found affected product or not.  This document helps to ensure that our customers have received the recall notification.

STEP 4	If you are a distributor, please follow steps 1 through 3.  In addition, you must notify all customers who received affected product distributed by you.  All product returns and product replacements should be handled by you the distributor.

Allegiance Healthcare has investigated this situation and has initiated appropriate corrective action to ensure patient safety.  Please accept our apology for any inconvenience this may have caused you and your staff.  If you should have any questions regarding this action, please contact Quality Systems-Professional Services at 1.800.292.9332.

Sincerely,

Patricia Sharpe-Gregg

Director, Quality Systems

Allegiance Healthcare Corporation

File No. 2002466

#### Attachment 2 of 2

Recall Notice from Allegiance Healthcare Corporation

a Cardinal Health company

CUSTOMER ACKNOWLEDGMENT FORM

ALLEGIANCE HEALTHCARE

RECALL NOTICE DATED:  NOVEMBER 15, 2002

This is to acknowledge that we have received your communication, concerning the above referenced action.  We have reviewed the documentation and are providing you with the following checked information:

INVENTORY SECTION

We have no affected inventory within our possession.

We have affected inventory within our possession.

The following quantities of the 4341B Thoracentesis Tray have been identified in our facility:

| Lot Code | Qty | Lot Code | Qty | Lot Code | Qty | Lot Code | Qty | Lot Code | Qty |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| L1J035X |  | L1N018 |  | L1S015 |  | L2B047 |  | L2D067 |  |
| L1J088 |  | L1N031 |  | L2A012 |  | L2B073 |  | L2E036 |  |
| L1K042 |  | L1N039X |  | L2A021 |  | L2C012 |  |  |  |
| L1K058 |  | L1N094 |  | L2A048 |  | L2C038 |  |  |  |
| L1K070 |  | L1P017 |  | L2A075 |  | L2C083 |  |  |  |
| L1L057X |  | L1P045 |  | L2B039 |  | L2D041 |  |  |  |

SIGNATURE SECTION

Print name of person completing this form

Print Institution Name, City, State, Zip

Signature of person completing this form: 							______

If your facility is a member of a health system or group operation, please identify on this line:

_

# Please return this form via fax to 1.847.689.9101

File No. 200246