Veterans Health Administration Warning System

Published by VA Central Office

April 28, 2004

Item:  	Implantable Cardioverter Defibrillator (ICD), models Micro-Jewell II 7223Cx and GEM DR 7271 manufactured by Medtronic, Inc. and implanted in 1997 and 1998.  Of 6268 units originally implanted, an estimated 1800 remain in use worldwide.

Specific Incident:	During a cardiac event requiring cardioversion or defibrillation, internal capacitors may take longer to charge and can cause a delay or non-delivery of appropriate shock therapy.

Action:  	1.	Immediately check patient records in CPRS to identify patients with the 	ICD implants described in this notification.

2.   Contact your local Medtronic sales representative (or Medtronic 	Technical Services, listed in the Contact section of this notification if 	you cannot reach the local representative) for assistance with 	determining whether the identified implant requires follow-up action as 	identified by Medtronic.

- If follow-up action is necessary, see the attached letter from Medtronic describing required actions.

Source:	Manufacturer and FDA

Contact:	Medtronic Technical Services: 1-800-723-4636

Edmund Keung, MD -- Director, VA Western Pacemaker Surveillance 				Center and National ICD Surveillance Center: (415) 750-2077

Ronald Jones -- Pacemaker Engineer, VA Eastern Pacemaker Surveillance 			Center: (202) 745-8504

Letter from Medtronic (PDF Format)