Veterans Health Administration Warning System

Published by VA Headquarters

May 30, 2002

Item: 	Philips/Agilent/Hewlett Packard Viridia information centers; central station monitors using the Hewlett Packard Vectra VL400 personal computer. This PC is used in models M3150A, M3151A, M3153A and M3150AU#A01 central station monitors. All units shipped between Feb 1, 2001 and Feb 21, 2000 are affected. Bedside monitors are NOT affected.

Specific Incident: 	The hard disk drive installed in the units may fail, causing loss of central station monitoring and alarming. Bedside monitoring and alarms are not affected.

Action: 	1. If you have the affected units and did not already receive and respond to Philips Medical Systems letter (attached), immediately (within 24 hours) check your central station monitoring serial numbers against the list in the attached letter from Philips Medical.

2. If affected, contact Philips Medical at (800) 548-8833 to arrange replacement of the hard drives.

Additional Information:

Some other Philips systems have this same problem, but these failures do not affect patient safety. These units are also identified in the attached letter. Please review the letter and contact Philips as needed. Also, please contact Paul Sherman at CEOSH via e-mail if you are affected by this recall and did not receive Philips’ letter.

Source: 		CEOSH and Manufacturer

Contact: 	Philips Technical Support at (800) 548-8833. You can also contact Paul Sherman, VA Center for Engineering & Occupational Safety and Health (CEOSH) at (314) 543-6700.

###### Philips Medical Systems

300 Minuteman Road

Andover, MA 01810

April 2002

Dear Biomedical Engineering Manager,

Through our product quality surveillance system, Philips Medical Systems has become aware of an issue with medical devices used for central monitoring, obstetrical care, enhanced ECG analysis and documentation management. These devices may have been distributed with Hewlett-Packard Vectra VL400 PCs containing disk drives with a higher than expected failure rate.

Voluntary Medical Device Recall

Philips Medical Systems has initiated a voluntary recall for the following products which may be affected by premature hard disk drive failure: Agilent and Viridia Information Centers, model numbers M3150A, M3151A, M3153A and M3150AU#A01.

Voluntary Medical Device Correction

Philips Medical Systems has initiated a medical device correction for the following products which may be affected by premature hard disk drive failure but do not pose a risk to patient safety: Agilent Technologies OBTraceVue, model numbers: M1382B, M1383B, M1384B, M1385B, M1383C and M1384C; Agilent Technologies MIDA System, model# M2040A; Agilent Technologies Inter-SDN Gateway, model# M1272A

You are being contacted because our shipping records indicate that your facility has received an Agilent or Viridia Information Center. Attached to this letter you will find the Medical Device Recall that provides instructions on how to identify the affected units and what actions are required of your facility.

If you feel you may have other products affected by the medical device correction program for which instructions are not attached to this letter, please contact the Philips Medical Support line at 1-800-548-8833.

We sincerely regret the inconvenience that this will cause you. Your satisfaction with Philips Medical products and with our response to this problem is very important to us.

Sincerely,

Elisabeth George

Vice President of Quality and Regulatory

Philips Medical Systems

Cardiology and Monitoring Systems

Philips Medical Systems 		URGENT: MEDICAL DEVICE RECALL

300 Minuteman Road			Agilent and Viridia Information Center

Andover, Ma 01810 			M3150A, M3151A, M3153A, M3150AU#AOI

April 2002

| AFFECTED PRODUCTS | Products: Information Center System and Upgrades M3150A Agilent or Viridia Information Center M3151A Agilent or Viridia Information Center Client M3153A Agilent or Viridia Surveillance Center M3150AU#AO1 Agilent or Viridia Information Center Upgrade Units Affected: Systems distributed with Hewlett-Packard Vectra VL400 personal computers ("PC") with serial numbers listed on page 3 of this notice. Systems were shipped between February 1, 2001 through February 21, 2002 Manufactured and Distributed by: Agilent Technologies Healthcare Solutions Group and Philips Medical Systems*. * Agilent Technologies Healthcare Solutions Group was acquired by Royal Philips Electronics on August 1. 2001. |
| --- | --- |
| REASON FOR RECALL | Hewlett-Packard Vectra VL400 PCs (HP Part# P3715W) with the listed serial numbers may contain disk drives with a higher than expected failure rate. |
| HAZARD INVOLVED | Loss of central patient monitoring and central patient alarming There is no prior indication or warning of the disk drive failure. Bedside monitoring and alarming is unaffected. |
| HOW TO IDENTIFY | See the attached photographs for system identification |
| REQUIRED ACTION | Using the directions on the following pages, identify the PCs that will require the hard disk drive replacement. You will receive a call from a Philips Medical Systems representative to schedule a repair. If you have questions about this recall notice, please call the Philips Medical Support line at 1-800-548-8833. |

M3150A, M3151A, M3153A, M3150AU#AOI

System Identification

| How to identify if you have an Agilent or Viridia Information Center subject to the recall: Locate the display and the PC for the Information Center The hard drive to be replaced resides within the Hewlett-Packard Vectra VL400 PC. If your Information Center PC is not a Hewlett- Packard Vectra VL400, it is not affected by this recall. Once you have identified the PC serial number, refer to the next page to determine if your unit needs a replacement drive | Information Center Display The Information Center display will look similar to this Information Center PC The PC may or may not be located with the display The PC will have a label located on the top that identifies the unit and the PC serial number The PC serial Number can also be found on the back of the unit |
| --- | --- |

| If your device's PC serial number is listed here, then it requires a replacement drive. A Philips Representative will be calling you shortly to schedule an appointment to perform the replacement | Serial Numbers of Affected PC's (Hewlett-Packard Vectra VL400, HP Part# P371SW) US1 0204040 through US1 071 0329 US1 0811665 through US10811789 US10902444 US10902817 throughUS10902887 US11304210 US11304576 through US11304670 US11307358 US11307974 through US11500994 US11608246 US11608393 throughUS11703817 US11704443 US11709038 through US11709052 US11800005 US11800291 through US11904733 US121 04028 US121 04584 through US121 04621 US121 06958 US121 07167 through US121 07365 US12203988 US 12203996 through US 12204354 US1221 0341 US1221 0494 through US1221 0535 US1221 0537 US1221 0539 through US1221 0575 US1221 0579 US1221 0581 US1221 0583 through US1221 0622 US12403504 US12403528 through US12408065 US12408067 US12408070 throughUS12408085 US12608928 US12609515 through US12609752 US12609757 through US12609883 US12618799 US12618806 through US12618807 |
| --- | --- |

M3150A, M3151 A, M3153A and M3150AU#AO1 Recall Notice

# System Identification