Veterans Health Administration Warning System

Published by VA Central Office

August 12, 2004

Item:  	Shiley Tracheosoft XLT Extended Length Tracheostomy Tube and Cannula FDA Class I Recall.

### General Information:	FDA and Nellcor/Tyco have issued a Class I (most serious) recall of the Shiley Tracheosoft XLT Extended Length Tracheostomy Tube and Cannula. The outer cannula may separate from the hub and neck flange allowing the outer cannula to travel farther into the patient’s airway leading to obstruction of the airway and significantly interfering with breathing and ventilation.

Specific Incident:	Nellcor/Tyco has received two reports in which the cannula separation was allegedly associated with patient deaths. This recall affects 73,355 disposable units that the company has shipped to the U.S. and international customers over the last four years.

Actions:	1.	Patients with the affected models (see attached) should be contacted and arrangements made to exchange the equipment.

- If you have the affected devices in stock work with your materials management (AMMS) to make sure that a recall package has been received. Follow the instructions in the package to receive credit for the recalled products.

- If you have these devices in stock and did not receive a recall package from the company, contact Nellcor/Tyco Technical Services Department at 1-800-635-5267.

- See attached information on comparable replacements from Smiths Medical Bivona® to the recalled products.

Source:   		FDA and Tyco Healthcare.

# Contact: 		Bryanne Patail at the VA National Center for Patient Safety (NCPS) phone: (734) 930-5852.

FDA Home Page ¦ CDRH Home Page ¦ Search ¦ CDRH A-Z Index ¦ Contact CDRH

| ## Medical Device Recalls Class 1 Recall: Shiley Tracheosoft Tracheostomy Tube and Disposable Inner Cannula |  |
| --- | --- |

| Date Recall Initiated: |  | July 8, 2004 |  |
| --- | --- | --- | --- |
| Product: |  | Shiley Tracheosoft XLT Extended Length Tracheostomy Tube and Disposable Inner Cannula |  |
| Use: |  | The device is an extended length cannula (tube) with accessory components used to provide an artificial airway to assist in the treatment of a variety of respiratory diseases and airway management in adults. The tube is inserted into a tracheostomy incision in a patient’s neck and trachea. The tracheostomy tube is secured in place through the tube’s hub and flange assembly with the use of a holder or neck strap. |  |
| Recalling Firm: |  | Nellcor/Tyco Healthcare 4280 Hacienda Drive Pleasanton, CA 94588-2719 |  |
| Reason for Recall: |  | The outer cannula may separate from the hub and neck flange allowing the outer cannula to travel farther into the patient’s airway and significantly interfere with breathing and ventilation. This recall affects 73,355 disposable units that the firm has shipped to U.S. and international customers over the last four years. |  |
| Public Contact: |  | James Bonds Senior Director, Regulatory Affairs 925-463-4371 james.bonds@tycohealthcare.com |  |
| FDA District: |  | San Francisco |  |
| FDA Comment: |  | Failure of the tracheostomy tube can allow the tube to migrate, leading to obstruction of the airway and subsequent lack of ventilation. Airway obstruction or failure to ventilate can lead to permanent neurological injury or death. Class I recalls are the most serious type of recall and involve situations where there is a reasonable probability that use of the product will cause serious injury or death. |  |

Pleasanton, CA – August 6, 2004 – On July 8, 2004, Nellcor/Tyco Healthcare, initiated a voluntary recall of the Shiley® TracheoSoft® XLT Extended Length Tracheostomy Tube. The Company has shipped 73,355 units over the last four years throughout the U.S. and Canada. This product can be used in the hospital or home-care environment to provide tracheal access for airway management. The recall was prompted by reports of the outer cannula separating from the swivel neck plate. While the number of customer reports of cannula separation is very low, if this failure mode does occur, the outer cannula could travel into the patient’s airway, potentially interfering with breathing or ventilation; which could result in serious injury or death.

Accordingly, in the interest of patient safety, our recall extends to all lots of the following products:

| Model Number | Description |
| --- | --- |
| 72110-050 | Size 5, Proximal Extension, Uncuffed |
| 72110-060 | Size 6, Proximal Extension, Uncuffed |
| 72110-070 | Size 7, Proximal Extension, Uncuffed |
| 72110-080 | Size 8, Proximal Extension, Uncuffed |
| 72120-050 | Size 5, Proximal Extension, Cuffed |
| 72120-060 | Size 6, Proximal Extension, Cuffed |
| 72120-070 | Size 7, Proximal Extension, Cuffed |
| 72120-080 | Size 8, Proximal Extension, Cuffed |
| 73110-050 | Size 5, Distal Extension, Uncuffed |
| 73110-060 | Size 6, Distal Extension, Uncuffed |
| 73110-070 | Size 7, Distal Extension, Uncuffed |
| 73110-080 | Size 8, Distal Extension, Uncuffed |
| 73120-050 | Size 5, Distal Extension, Cuffed |
| 73120-060 | Size 6, Distal Extension, Cuffed |
| 73120-070 | Size 7, Distal Extension, Cuffed |
| 73120-080 | Size 8, Distal Extension, Cuffed |
| 77100-050 | Size 5, XLT, Disposable Inner Cannula |
| 77100-060 | Size 6, XLT, Disposable Inner Cannula |
| 77100-070 | Size 7, XLT, Disposable Inner Cannula |
| 77100-080 | Size 8, XLT, Disposable Inner Cannula |

The Company has received two reports in which the cannula separation was allegedly associated with two patient deaths. Those TracheoSoft XLT Tracheostomy Tubes currently in use should be exchanged, when practical, with an alternative model at the discretion and judgment of the patient’s physician.

Institutions with affected product have received packages outlining the recall process. As of this date, 80 percent of those customers have responded to the recall action. The U.S. Food and Drug Administration (FDA) and officials in other affected countries have been apprised of this recall action.

This action does not affect any other model of Shiley tracheostomy products. We apologize for any disruption this action may cause physicians and their patients. Please be assured that Nellcor/Tyco Healthcare has taken appropriate steps to prevent recurrence of this problem and ensure patient safety.

Clinicians and patients with inquiries should contact our Technical Services Department at 1-800-635-5267, option 3.

About Tyco Healthcare
Tyco Healthcare, a business segment of Tyco International Ltd., is a leading manufacturer, distributor and servicer of medical devices worldwide. Its broad portfolio includes disposable medical supplies, monitoring equipment, medical instruments and bulk analgesic pharmaceuticals, sold under such names as Auto Suture, Kendall, Mallinckrodt, Nellcor, Puritan Bennett, United States Surgical, Valleylab and others.

Updated August 9, 2004