Veterans Health Administration Warning System

Published by VA Central Office

Item:  	Smiths Medical, Fast Flow Fluid Warmer Models 250, 500 and 1000.

Specific Incident:	Hospitals in Australia report serious patient injury from intravascular air embolisms introduced while using the Smiths Medical fluid warmers identified in this Alert.

Action:  	1. Promptly (within one week) check local inventory to identify all units affected by this Alert and institute recommended practices to minimize risk (attached) for all users of Smiths Medical Fast Flow Fluid Warmers, models 250, 500 and 1000.

2.  Procure and install the Smiths Medical accessory Air Detector/Clamp, model H-31 that is designed to reduce the risk of air embolism introduction while using the Smiths Medical Fast Flow Fluid Warmers, models 250, 500 and 1000.

Addl. Information: 	All new models of Smiths Medical fluid warmers include the air detector/clamp as a standard component.   Also, while we are unaware of other sites without this accessory, check your inventory and contact your Smiths Medical representative to update their list and order the H-31 detector/clamp.

Source:	VA Center for Engineering & Occupational Safety and Health

Smiths Medical and ECRI

Contact:	Paul Sherman (CEOSH):  314-543-6700

- Smiths Medical recommendations to minimize

- risk of air embolism while using their fluid warmers:

(1) 	Remove all air from fluid bags and lines before connecting to patients.

(2) 	Follow all instructions for priming the infusion line and set.

(3) 	Monitor fluid lines, and check to ensure that they are free of air bubbles.

(4) 	Do not open ports in the infusion system unless precautions are taken to prevent the entry of air.

(5) 	Do not reconnect partially emptied fluid bags. Disconnected bags may contain a significant volume of air, which can enter the bags when they are first disconnected.

(6) 	Do not use autotransfusion bags with warming systems unless the warmer is clearly indicated for this use.

List of VA sites with H1000 fluid warmers without the H-31 Air Detector/Clamp - as identified by Smiths Medical:

| City | State | Street Address | Serial Number |
| --- | --- | --- | --- |
| Chicago | IL | 2030 West Taylor St | J02087 |
| Clarksburg | WV | 1 Medical Center Dr | 20020120 |
| Clarksburg | WV | 1 Medical Center Dr | J02062 |
| Cleveland | OH | 10701 East Boulevard | 20030492 |
| Gainesville | FL | 1601 SW Archer Rd | J01886 |
| Huntington | WV | 1540 Spring Valley Dr | 20030151 |
| Indianapolis | IN | 1481 West 10th Street | 20010209 |
| Lexington | KY | 2250 Leestown Rd/Bldg 12 | J01386 |
| Lexington | KY | 1101 Veterans Dr | 20020139 |
| Lexington | KY | 1101 Veterans Dr | 20020135 |
| Lexington | KY | 1101 Veterans Dr | 20020136 |
| Los Angeles | CA | 1103 Wilshire Blvd | 20000648 |
| Louisville | KY | 800 Zorn Ave | 20010033 |
| New York | NY | 423 East 23rd St | 19990447 |
| Philadelphia | PA | University Ave | J01791 |
| Pittsburgh | PA | University Dr | 20000379 |
| Shreveport | LA | 510 East Stoner Ave | J01725 |
| Temple | TX | 1901 S 1st St | 20030300 |
| Tucson | AZ | 3601 S 6th Ave | 19990286 |
| Wichita | KS | 5500 East Kellogg | 20030482 |