1

DEPARTMENT OF HEALTH AND HUMAN SERVICES

FOOD AND DRUG ADMINISTRATION

CENTER FOR BIOLOGICS EVALUATION AND RESEARCH

BLOOD PRODUCTS ADVISORY COMMITTEE

80th Meeting

This transcript has not been edited or corrected, but appears as received from the commercial transcribing service:  Accordingly the Food and Drug Administration makes no representation as to its accuracy.

Thursday, July 22, 2004

8:00 a.m.

Holiday Inn Gaithersburg

Two Montgomery Village Avenue

Gaithersburg, Maryland

2

PARTICIPANTS

James R. Allen, M.D., MPH, Acting Chair

Linda A. Smallwood, Ph.D., Executive Secretary

MEMBERS

Kenneth Davis, Jr., M.D.

Donna M. DiMichele, M.D.

Samuel H. Doppelt, M.D.

Jonathan C. Goldsmith, M.D.

Harvey G. Klein, M.D.

Suman Laal, Ph.D.

ACTING CONSUMER REPRESENTATIVE

Katherine E. Knowles

NON-VOTING INDUSTRY REPRESENTATIVE

Michael D. Strong, Ph.D.

TEMPORARY VOTING MEMBERS

Liana Harvath, Ph.D.

Matthew J. Kuehnert, M.D.

Susan F. Leitman, M.D.

Keith C. Quirolo, M.D.

George B. Schreiber, Sc.D.

Donna S. Whittaker, Ph.D.

3

C O N T E N T S

PAGE

Welcome, Statement of Conflict of Interest,

Announcements

Linda Smallwood, Ph.D.                           5

James R. Allen, M.D.                            11

Committee Updates

FDA Current Thinking on TRALI:

Leslie Holness, M.D.                          13

Donor Blood Pressure Determination:

Alan Williams, Ph.D.                          23

Open Public Hearing

TRALI:

Kay Gregory, AABB, ABC                       32

Michael Fitzpatrick, Ph.D., ABC              40

Donor Blood Pressure Determination:

Kay Gregory, AABB, ABC                       50

I. Dating of Irradiated Red blood Cells

Introduction and Background:

Ping He, M.D.                                   60

Presentation:

Gary Moroff, Ph.D.                              80

Presentation:

Larry Dumont                                   114

Presentation:

Dean Elfath, M.D.                              130

Presentation:

Jessica Kim, Ph.D.                             137

Open Public Hearing

Allene Carr-Greer, AABB                        159

Michael Fitzpatrick, Ph.D.                     162

Richard Davey, M.D.,

New York Blood Centers                       174

4

C O N T E N T S (Continued)

PAGE

FDA Current Thinking and Questions for the

Committee:

Jaro Vostal, M.D., Ph.D.                       177

Committee Discussions and Recommendations                184

II.  New Standard for Platelet Evaluation

Introduction and Background:

Salim Haddad, M.D.                             231

Presentation:

James AuBuchon, M.D.                           283

Presentation:

Edward Snyder, M.D.                            218

Open Public Hearing

Allene Carr-Greer, AABB                        308

Michael Fitzpatrick, Ph.D.                     308

Larry Dumont, Gambro BCT Inc.                  308

FDA Current Thinking and Questions for the

Committee:

Jaro Vostal, M.D., Ph.D.                       310

Committee Discussion and Recommendations                 312

III.  Experience with Monitoring

of Bacterial Contamination of Platelets

Introduction and Background:

Jaro Vostal, M.D., Ph.D.                       342

Summary of ACBSA Meeting: Bacterial Contamination:

Jerry A. Holmberg, Ph.D.                       371

Presentation:

Steven Kleinman, M.D.                          388

Open Public Hearing

Boris Rotman, Ph.D., BCR Diagnostics           420

Committee Discussion and Recommendations                 430

5

1                      P R O C E E D I N G S

2            Welcome/Statement of Conflict of Interest

3             DR. SMALLWOOD:  Welcome to the 80th

4   meeting of the Blood Products Advisory Committee.

5             I am Linda Smallwood, the Executive

6   Secretary.  At this time, I will read the conflict

7   of interest statement that applies to this meeting.

8             This announcement is part of the public

9   record for the Blood Products Advisory Committee

10   meeting on July 22nd/23rd, 2004.

11             Pursuant to the authority granted under

12   the Committee Charter, the Director of FDA's Center

13   for Biologics Evaluation and Research has appointed

14   the following individuals as temporary voting

15   members:  Drs. Liana Harvath, Blaine Hollinger,

16   Matthew Kuehnert, Susan Leitman, Keith Quirolo,

17   George Schreiber, Donna Whittaker, Ms. Katherine

18   Knowles.

19             To determine if any conflicts of interest

20   existed, the agency reviewed the agenda and all

21   relevant financial interests reported by the

22   meeting participants.

6

1             For Agenda Topics I, II, III, and V, the

2   Food and Drug Administration has prepared general

3   matter waivers for the special government employees

4   participating in this meeting who required a waiver

5   under Title 18, United States Code 208.

6             Because general topics impact on so many

7   entities, it is not prudent to recite all potential

8   conflicts of interest as they apply to each member.

9   FDA acknowledges that there may be potential

10   conflicts of interest, but because of the general

11   nature of the discussions before the committee,

12   those potential conflicts are mitigated.

13             Based on a review of the agenda, all

14   relevant financial interests reported by the

15   meeting participants, and on the FDA draft guidance

16   on disclosure of conflict of interest for special

17   government employees participating in an FDA

18   product-specific advisory committee meeting, there

19   are no meeting participants who required a waiver

20   under Title 18, United States Code 208 for

21   discussions on hepatitis B virus nucleic acid

22   testing for donors of whole blood.

7

1             We would like to note for the record that

2   Dr. Michael Strong is participating in this meeting

3   as the Non-Voting Industry Representative acting on

4   behalf of regulated industry.  Dr. Strong's

5   appointment is not subject to Title 18, United

6   States Code 208.

7             He is employed by the Puget Sound Blood

8   Center and Program and thus has a financial

9   interest in his employer.  He also is a researcher

10   for two firms that could be affected by the

11   committee discussion.  In addition, in the interest

12   of fairness, FDA is disclosing that his employer

13   Puget Sound Blood Center has associations with

14   regional hospitals and medical centers.

15             With regard to FDA's invited guest

16   speakers, the Agency has determined that the

17   services of these guest speakers are essential.

18   There are interests that are being made public to

19   allow meeting participants to objectively evaluate

20   any presentation and/or comments made by the

21   guests.

22             For the discussions of Topic I related to

8

1   the Dating of Irradiated Blood, Dr. Gary Moroff is

2   employed by the American Red Cross Holland Labs.

3             For the discussions of Topic II on a New

4   Standard for Platelet Evaluation, Dr. Edward Snyder

5   is employed by the Yale-New Haven Hospital Blood

6   Bank.  He also has associations with clinical

7   trials that involve red blood cells.

8             Dr. James AuBuchon has grants and/or

9   contracts with firms that could be affected by the

10   discussions.  He is also a scientific advisor for

11   several affected firms.

12             For the discussion of Topic III on

13   Experiences with Monitoring of Bacterial

14   Contamination of Platelets, Dr. Steven Kleinman

15   receives consulting fees from two firms that could

16   be affected by the committee discussions.

17             Dr. Jerry Holmberg has a financial and

18   professional interest in several firms that could

19   be affected by the committee discussions.

20             In addition, there are regulated industry

21   and other outside organization speakers making

22   presentations. These speakers have financial

9

1   interests associated with their employer and with

2   other regulated firms.  They were not screened for

3   these conflicts of interest.

4             FDA members are aware of the need to

5   exclude themselves from the discussions involving

6   specific products or firms for which they have not

7   been screened for conflicts of interest.  Their

8   exclusion will be noted for the public record.

9             With respect to all other meeting

10   participants, we ask in the interest of fairness

11   that you state your name, affiliation, and address

12   any current or previous financial involvement with

13   any firm whose products you wish to comment upon.

14   Waivers are available by written request under the

15   Freedom of Information Act.

16             At this time, I am asking if there are any

17   further declarations that have not been mentioned

18   that need before this meeting proceeds.

19             [No response.]

20             DR. SMALLWOOD:  Hearing none, thank you.

21             I would also just like to announce that

22   there is a new procedure and that for each day, and

10

1   also maybe for specific topics, there will be

2   another reading of a conflict of interest

3   statement.  That is new, but just to let you know

4   that that is what is taking place.

5             Also, with regard to those speakers that

6   will be speaking in the open public hearing, there

7   will be a statement read by the chairman for each

8   open public hearing to remind you to make the

9   declaration of your name and affiliation and to

10   reveal any association that is pertinent to that

11   discussion.

12             At this time, I would like to make a few

13   announcements.

14             There will be a workshop on plasma

15   standards scheduled August 31st through September

16   the 1st, 2004.  It will be held on the NIH campus,

17   and there is an announcement on the FDA web site.

18             Additionally, the next meeting of the

19   Blood Product Advisory Committee is tentatively

20   scheduled for October 21st/22nd, 2004 at this

21   hotel.  There will be further announcements.

22             At this time, I will introduce to you the

11

1   members of the Blood Products Advisory Committee.

2             Today, Dr. James Allen will be the Acting

3   Chairman in the absence of Dr. Kenrad Nelson, who

4   is expected to join us tomorrow.  Dr. Allen, would

5   you please raise your hand. Thank you.

6             As I call your names, would you please

7   raise your hand.

8             Dr. Kuehnert.  Dr. Harvath.  Dr. Klein.

9   Dr. Goldsmith.  Dr. Leitman.  Dr. Doppelt.  Dr.

10   DiMichele.  Dr. Davis.  Dr. Laal.  Dr. Quirolo.

11   Dr. Whittaker.  Dr. Schreiber.  Ms. Knowles.  Dr.

12   Strong.

13             Thank you.

14             As indicated on the agenda, we do have

15   times indicated for the speakers.  We would ask

16   that you would adhere to that.  Our Acting Chairman

17   says he will enforce that and we have a timer.

18             At this time, I would like to turn over

19   the proceedings of this meeting to the Acting

20   Chairman, Dr. James Allen.

21             DR. ALLE                                  N:  Thank you, Dr.

Smallwood.

22             Good morning and welcome to the meeting.

12

1   We have a very full agenda with a lot of important

2   items.  I don't think in my experience on the

3   committee I have ever seen so many questions being

4   asked in one meeting, so it is important that we

5   get the information before us from the speakers as

6   succinctly as possible, that we maximize the time

7   that we have for committee discussion and questions

8   of the speakers, and discussion among ourselves

9   before deciding to vote.  So, I really would like

10   to ask people, please, to keep your presentations

11   to the point and move along properly.

12             We have got two committee updates

13   initially and then we will follow that by an open

14   public hearing.  There are comments during the open

15   public hearing that will be addressing both of the

16   updates, but we will have both updates first with

17   time for questions of the speakers.

18             At this point, let's move into the first

19   committee update, Dr. Leslie Holness from the Food

20   and Drug Administration will give an update on

21   Transfusion Related Acute Lung Injury (TRALI).

22                        Committee Updates

13

1                  FDA Current Thinking on TRALI

2                       Leslie Holness, M.D.

3             DR. HOLNESS:  Thank you, Dr. Allen.

4             Good morning.

5             [Slide.]

6             The FDA Fatality Program receives reports

7   of fatalities that occur as a complication of

8   transfusion or donation.  We have seen a steady

9   rise in fatalities due to TRALI since the first FDA

10   report in 1992.

11             [Slide.]

12             This slide covers reported fatalities for

13   three fiscal years.  Between 2001 and 2003, the

14   three principal causes reported in terms of numbers

15   are TRALI, ABO hemolytic reactions primarily caused

16   by clerical errors, and bacterial contamination.

17             In Fiscal 2001 and 2003 TRALI led in the

18   number of fatality reports received.  In Fiscal

19   2002, reports of fatalities from bacterial

20   contamination of products were most numerous.

21   Other transfusion related fatality causes were

22   non-ABO, antibodies, and mishandling of products.

14

1   In this category, the transfusion may or may not

2   have contributed to the recipient's death.

3             In this category, the fatalities were not

4   transfusion related, and there are donor fatalities

5   and the total fatalities at the bottom of the

6   slide.

7             [Slide.]

8             If we look at the average of the key

9   causes for the last three years, TRALI leads with

10   16.3 percent followed by ABO hemolytic transfusion

11   reactions at 14.3 percent, and bacterial

12   contamination at 14.1 percent.

13             [Slide.]

14             So, the FDA Fatality Program reports that

15   TRALI was implicated in 16 to 22 percent of total

16   fatalities reported in each of the last three

17   years, and it was the most common cause of

18   transfusion related fatalities reported to the FDA

19   in 2003.

20             The majority of deaths were associated

21   with fresh frozen plasma followed by red blood

22   cells and apheresis platelets.

15

1             [Slide.]

2             Dr. Kathleen Sazama, of M.D. Anderson

3   Cancer Center at the University of Texas, looked at

4   20 years of FDA fatality reports from 1976 to 1995,

5   and found respiratory deaths as a percentage of

6   total reported deaths to be 15 percent, and many of

7   these are probably due to TRALI.

8             [Slide.]

9             This slide is a bar graph of TRALI

10   fatalities reported to the FDA and the total

11   fatalities reported to the FDA from 1995 to 2003.

12   There has been a steady increase in total fatality

13   reports to spike in 1998 and also a steady increase

14   in the TRALI fatalities.

15             [Slide.]

16             These are the TRALI fatalities broken out.

17   There is a slowing in 1999 and 2000, but all

18   together there is a steady increase in TRALI

19   fatalities up to 2003.

20             [Slide.]

21             Some of the fatalities are associated with

22   HLA or granulocyte antibodies, and they are sent in

16

1   with the fatality reports.

2             This is a graph of the number of

3   fatalities due to TRALI that were reported to the

4   FDA in these various years, and these are the

5   number of cases where HLA or antigranulocyte

6   antibodies were found.  In most cases, antibodies

7   were found in over 50 percent of the TRALI

8   fatalities.

9             [Slide.]

10             This slide shows the preliminary results

11   of a consensus conference held in Toronto, Canada,

12   in April of this year, 2004.  The conference was

13   sponsored by Canadian Blood Services, Hema-Quebec,

14   and the International Society for Blood

15   Transfusion, ISBT.

16             It was a two-day conference with over 19

17   speakers.  There are preliminary results.  More

18   detailed results will be published at the beginning

19   of next year.  So, the magnitude of the TRALI risk

20   is unknown.  Depending on the studies, the

21   estimates are between 1 in 5,000 to 1 in 10,000

22   transfusions.

17

1             There is evidence for two mechanisms for

2   TRALI, and there is insufficient evidence for

3   screening tests and for donor exclusion measures at

4   this time.

5             [Slide.]

6             In April of 2003, the NHLBI convened a

7   working group of TRALI experts to develop a clear

8   definition to be used to clinical investigation and

9   patient care.  The definition with limitations is

10   as follows:

11             In patients with no acute lung injury

12   prior to transfusion, the diagnosis of TRALI is

13   made if there is new acute lung injury and an onset

14   during or within 6 hours after the end of a

15   transfusion of one or more plasma containing blood

16   products, and there are no other risk factors for

17   acute respiratory distress syndrome.  This

18   definition is still being worked on.

19             [Slide.]

20             These are FDA actions taken in 2001.  The

21   issue was presented to the Blood Products Advisory

22   Committee on June 15th of 2001.  We will see the

18

1   results in the next slide.

2             CBER has published a Health Alert in the

3   form of Dear Colleague letter to the blood

4   community in October of 2001.  It was to remind

5   physicians to include TRALI in a differential

6   diagnosis of a patient in respiratory distress

7   during or following a transfusion.

8             Pre-storage leukocyte reduction of blood

9   products was recommended to help prevent formation

10   of leukocyte antibodies in recipients.

11             We recommended voluntary Med Watch

12   reporting of non-fatal TRALI cases, and there were

13   several poster presentations to raise clinician

14   awareness of TRALI.

15             [Slide.]

16             This slide shows the BPAC vote on June 15,

17   2001. The question to the committee was:  Should

18   the FDA consider regulatory action at this time to

19   identify donors and donations at increased risk to

20   producing TRALI in a recipient?

21             The votes were:  1 Yes, 13 No, and there

22   were no abstentions.

19

1             [Slide.]

2             One member thought it was prudent to

3   identify and defer donors implicated in multiple

4   TRALI cases.

5             BPAC agreed that this should be the

6   responsibility of each establishment.

7             The committee also recommended research to

8   define the scope of the syndrome and a prospective

9   epidemiologic study to establish incidence, donor

10   and recipient risks.

11             [Slide.]

12             The further recommendations from the

13   committee.

14             The role of HLA, leukocyte antibodies and

15   other potential causative mechanisms need to be

16   investigated.  A careful evaluation of cases in

17   which the donor can be linked with the reaction.

18             A multi-center study to assess and

19   evaluate acute pulmonary reactions and lung

20   problems in the transfusion setting using a

21   standardized protocol, and the surveillance of

22   recipients of IVIG for TRALI reactions.

20

1             [Slide.]

2             These are possible future regulatory

3   strategies that are being discussed at the FDA at

4   this time.

5             Diversion of plasma from female donors to

6   components other than fresh frozen plasma.  This

7   does not involve a new question and fresh frozen

8   plasma is most often involved in TRALI.  This is

9   being tried in the UK at this time, but there have

10   been no impressions of the results yet.

11             Our problem is that the plasma in other

12   components are ignored and that shortages of FFP

13   may occur.

14             [Slide.]

15             Preventive antibody testing and

16   questioning of donors, female donors, on parity,

17   followed by plasma product diversion and red blood

18   cell loss from donors at risk.

19             The problem here is that samples and

20   testing are not standardized.  All white blood cell

21   antibodies may not be equal in their ability to

22   cause TRALI in recipients.

21

1             [Slide.]

2             Defer donors implicated in a single unit

3   or in more than one multiple unit TRALI case

4   regardless of antibody status.

5             This allows the first case of TRALI to

6   occur which may be fatal, and it depends on

7   accurate case reports and donor tracing.

8             That's it.  With that, I end my

9   presentation.

10             DR. ALLEN:  Thank you, Dr. Holness.

11             Questions from the committee members?

12             Obviously, this is very important data,

13   but it is limited in that it is reporting only of

14   fatalities.  Do you have other information in terms

15   of how well the research community has responded to

16   this issue?  Are there any recommendations coming

17   out of the recent meeting that you think should

18   come before the committee at least today or in the

19   near future?

20             DR. HOLNESS:  I think that probably the

21   best thing is to wait until the full report of the

22   committee is out before we make recommendations.

22

1             DR. ALLEN:  Okay.  Other questions?  Yes.

2             DR. SCHREIBER:  From your graph it looks

3   like we are seeing an increased frequency of TRALI,

4   but it is probably due to more awareness, don't you

5   think, of the reporting, particularly since all of

6   the activities that started around '99?

7             DR. HOLNESS:  That is true.

8             DR. SCHREIBER:  My other question is on

9   one of the slides from the Toronto, you had an

10   incidence, I think it was 1 in 5,000 to 1 in

11   100,000, and that is the incidence of TRALI

12   reactions, TRALI-type reactions, but the mortality

13   rate is somewhere closer to 1 in 750,000, I

14   believe.

15             DR. HOLNESS:  I think you are right on

16   that, yes.

17             DR. SCHREIBER:  Thank you.

18             DR. ALLEN:  On that slide, it said 1 in

19   5,000 and 1 in 100,000, but I think you read 1 in

20   5,000 and 1 in 10,000.  Which is the correct

21   number, the 10,000 or 100,000?

22             DR. HOLNESS:  1 in 100,000.  It is my

23

1   mistake, I am sorry.

2             DR. ALLEN:  Thank you very much.

3             We will move on to the second committee

4   update, Donor Blood Pressure Determination

5   presented by Dr. Alan Williams.

6                Donor Blood Pressure Determination

7                       Alan Williams, Ph.D.

8             DR. WILLIAMS:  Thank you, Jim, and good

9   morning.

10             As you will note from some of the

11   statements from the blood and plasma community that

12   have been distributed, FDA has been asked to

13   restate and reconsider its position with respect to

14   blood pressure determination as a criterion for

15   blood donation and plasma donation eligibility.

16   That is what I intend to do very briefly this

17   morning.

18             [Slide.]

19             The FDA regulatory position is stated

20   quite clearly in two regulations.  21 CFR

21   640.3(b)(2) requires donor's systolic and diastolic

22   blood pressure are within normal limits, unless a

24

1   physician, after examining the donor, is satisfied

2   that the donor is otherwise qualified.

3             This needs to be considered in conjunction

4   with another regulation, 21 CFR 606.100(b)(2),

5   which states that a blood collection facility

6   include in its Standard Operating Procedures

7   methods of performing donor qualifying tests and

8   measurements, including minimum and maximum values

9   for a test or a procedure when a factor in

10   determining acceptability.

11             [Slide.]

12             When reviewing Standard Operating

13   Procedures presented by licensed blood collection

14   establishments, in fact, we do look for SOPs that

15   define both an upper and a lower range of normal

16   blood pressure, and, in addition, if outside the

17   normal range, a donor must be medically evaluated

18   for donation eligibility.

19             Not only do we do that in current

20   submissions, but we, in fact, did a randomized look

21   at prior approvals of SOPs, and in all of the

22   licensed establishments that we looked at, they had

25

1   both lower and upper limits included.

2             FDA has not historically specified the

3   cutoff values to be used for a lower limit.  This

4   is, in fact, controversial as to what the

5   predictive value of the lower limit is and what the

6   lower limit of normal, in fact, should be,

7   certainly a subject for future discussion.

8             But I will note that while there are some

9   studies which have been cited by some of the

10   position statements, what probably is lesser known

11   is that FDA has received some isolated reports of

12   severe vasovagal reactions in donors who were

13   found, upon review of the record, to have had

14   abnormally low blood pressures at the time of

15   donation.

16             [Slide.]

17             To summarize, I think what the basis is of

18   the industry request for policy clarification, on

19   the fact that the predictive value of a single low

20   blood pressure determination has not been finally

21   established, I think you can see a range in the

22   literature, and in some of the cases, particularly

26

1   the case-controlled studies, you can see that blood

2   pressure on a univariate analysis emerges as a

3   factor, but may not stand up to a multivariate

4   analysis.

5             This is evidence that it may not be an

6   independent factor, but, in fact, may be tied up in

7   interaction with demographic or other variables.

8   So, analyses of some of these studies require both

9   large studies and rather complex multivariate

10   analysis to determine what the interaction effects

11   and other potential impact might be.

12             The European community, particularly the

13   UK, in their blood collection procedures do not

14   determine a blood pressure value at all, although

15   if the donor has a history of reactions or of

16   hypertension, they maintain the equipment available

17   to make the determination, but, in short, in the

18   UK, the blood pressure is not determined.

19             The 2004 EU directive does not include a

20   blood pressure determination requirement, and the

21   current Council of Europe guide includes only an

22   upper blood pressure limit.

27

1             [Slide.]

2             And though not necessarily scientifically

3   based, the observation has been made that the

4   voluntary industry standards for blood collection,

5   which originally required both an upper and lower

6   blood pressure value, were modified to remove the

7   lower level requirement some time ago, in 1987, and

8   that some blood establishment SOPs may current

9   omit, or may have historically omitted, a lower

10   blood pressure cutoff value.

11             As I stated, licensed establishments are

12   reviewed for having an SOP that includes this

13   requirement, it is possible that some of the sites

14   that don't may be registered facilities which

15   should be following the regulation, but whose SOPs

16   are not reviewed by FDA.

17             [Slide.]

18             So, in summary, FDA strictly adheres to

19   the existing regulations, but FDA does not

20   recognize the need for scientific consensus on the

21   value of donor blood pressure determinations and

22   considers its regulations to require that they be

28

1   scientifically based.

2             So, I think some of the uncertainties that

3   may be there in the published literature should be

4   looked at further.

5             Under the HHS Blood Action Plan, FDA

6   intends to propose rulemaking that will

7   comprehensively address donor eligibility

8   requirements including blood pressure, and as part

9   of this process, there will be an opportunity for

10   data presentation and comment to any proposed rule

11   that might emerge.

12             Thank you.

13             DR. ALLEN:  Thank you, Dr. Williams.

14             Questions or comments from the committee?

15   Yes.

16             DR. GOLDSMITH:  What were the blood

17   pressures in the FDA reports for the severe

18   reactions, how low were they really?

19             DR. WILLIAMS:  I do not remember the

20   actual values, but they were lower than what the

21   original industry standard was, which I believe the

22   lower limits were 90 and 50 for the systolic and

29

1   diastolic.

2             DR. KLEIN:  Alan, I am sure you are aware

3   that there may be a little bit of a reporting bias

4   in those reports you have.  There has been an

5   extensive literature on vasovagal reactions and

6   donor vital signs, and to the best of my knowledge,

7   there has never been any correlation between blood

8   pressure and vasovagal reactions, and knowing how

9   vasovagal reactions generally occur, I am not sure

10   that there should be.

11             DR. LEITMAN:  You quote a study, which is

12   an excellent one, a multi-center study published in

13   Transfusion in '99, where Trend and colleagues

14   looked at the effect of blood pressure and other

15   factors pre-donation on the incidence of vasovagal

16   during donation, and in a univariate analysis, low

17   blood pressure was associated with vasovagal

18   reactions, but in a regression analysis, which is

19   very important, when you put in the variables of

20   age, weight, and donor status prior donations, that

21   fell out.

22             So, you really have to go with the best

30

1   scientific data you have, I think, in a complex

2   analysis like this, and that is very helpful for me

3   to look at this data.

4             DR. WILLIAMS:  Yes, I agree.  There was

5   probably a most sophisticated analysis to address

6   this particular subject, but, you know, one could

7   argue was that study large enough to pick up a

8   potential interaction effect between the

9   demographic variables and blood pressure.

10             In fact, on univariate analysis, there was

11   quite a difference, 3 percent incidence of

12   reactions with the lower blood pressures versus 1

13   percent of the control group, so I think it just

14   bears a further look with more sophisticated

15   analysis.

16             DR. ALLEN:  Other questions or comments?

17   Okay.  Thank you very much.

18             At this point, we will move on the open

19   hearing, to the public hearing.

20             Before we get started on that, I need to

21   read an open public hearing announcement for

22   general matters meetings.

31

1             Both the Food and Drug Administration and

2   the public believe in a transparent process for

3   information gathering and decisionmaking.  To

4   ensure such transparency at the open public hearing

5   session of the Advisory Committee meeting, FDA

6   believes that it is important to understand the

7   context of an individual's presentation.

8             For this reason, FDA encourages you, the

9   open public hearing speaker, at the beginning of

10   your written or oral statement to advise the

11   committee of any financial relationship that you

12   may have with any company or any group that is

13   likely to be impacted by the topic of this meeting.

14   For example, the financial information may include

15   a company's or a group's payment of your travel,

16   lodging, or other expenses in connection with your

17   attendance at the meeting.

18             Likewise, FDA encourages you at the

19   beginning of your statement to advise the committee

20   if you do not have any such financial

21   relationships.  If you choose not to address this

22   issue of financial relationships at the beginning

32

1   of your statement, it will not preclude you from

2   speaking.

3                       Open Public Hearing

4             DR. ALLEN:  Let's go ahead with the public

5   statements on TRALI.  I have a request from the

6   American Association of Blood Banks.

7             MS. GREGORY:  Thank you.  My name is Kay

8   Gregory and I am the Director of Regulatory Affairs

9   for the AABB, and I have only financial

10   arrangements with them and no other companies.

11             AABB is an international association

12   dedicated to advancing transfusion and cellular

13   therapies worldwide.  Our members include more than

14   1,800 hospital and community blood centers and

15   transfusion and transplantation services as well as

16   approximately 8,000 individuals involved in

17   activities related to transfusion, cellular

18   therapies, and transplantation medicine.

19             For over 50 years, AABB has established

20   voluntary standards for, and accredited

21   institutions involved in, these activities.  AABB

22   is focused on improving health through the

33

1   advancement of science and the practice of

2   transfusion medicine and related biological

3   therapies, developing and delivering programs and

4   services to optimize patient and donor care and

5   safety.

6             The AABB believes that TRALI is a

7   significant transfusion safety concern that merits

8   increased awareness and research.  In an effort to

9   educate our members about the clinical and

10   laboratory features of TRALI, AABB has issued

11   guidelines for the management of TRALI, and our

12   association considers this a priority transfusion

13   safety matter.

14             We commend the FDA for alerting physicians

15   to the risk of TRALI from transfusion of

16   plasma-containing blood products in 2001, however,

17   we are disappointed that the Federal Government has

18   not done more to advance needed research regarding

19   this important transfusion safety issue since the

20   Blood Products Advisory Committee last addressed

21   TRALI in 2001.

22             In order to allow for the most effective

34

1   and meaningful research and clinical understanding

2   of this condition, the AABB proposed that a

3   standard uniform definition of TRALI be established

4   and adopted by the medical community and

5   policymakers, including the FDA.

6             Earlier this year, Canadian Blood Services

7   and Hema-Quebec hosted a valuable consensus

8   conference, bringing together the leading experts

9   to discuss the current state of knowledge regarding

10   TRALI.

11             At the end of this conference, the group

12   recommended definitions of TRALI and "possible

13   TRALI," and we have attached to our written

14   statement our current understanding of those

15   definitions.

16             In general, the group recommended that

17   TRALI should be diagnosed in patients with no acute

18   lung injury prior to transfusion who, during or

19   within six hours after transfusion, experienced

20   certain specific criteria.  They distinguished

21   "possible TRALI" cases, which would involve

22   patients with the same criteria who also had one or

35

1   more temporally associated ALI risk factors.

2             The AABB endorses the definitions set

3   forth during the consensus conference and urges the

4   FDA to adopt these definitions as well.  Emerging

5   data and research regarding TRALI should be

6   carefully monitored to determine if refinements to

7   these definitions are necessary over time.

8             Using the uniform definitions, AABB

9   recommends that additional research be conducted to

10   define the scope of the problem and its mechanisms

11   or pathophysiology.  As we proposed to BPAC in

12   2001, AABB continues to advocate a prospective

13   epidemiologic study to establish the incidence of

14   TRALI.  For example, we propose a multi-center

15   study of acute lung problems in the transfusion

16   setting to assess, evaluate, and analyze all

17   pulmonary reactions using a standardized protocol.

18             The AABB also continues to recommend that

19   the NHLBI establish a multi-center study to lead to

20   a better understanding of the mechanisms that cause

21   TRALI.  Once the mechanisms of TRALI are better

22   understood, the risk factors in donors and

36

1   recipients may become apparent.

2             The AABB continues to believe that more

3   data are needed before establishing donor deferral

4   criteria or other regulatory strategies for TRALI.

5   When a severe clinical reaction has occurred, an

6   antibody has been identified in the donor and the

7   recipient has the corresponding antigen, the

8   preventive measure is relatively clear.

9             In such cases, it is generally agreed that

10   blood from that donor should never again be

11   transfused to the same recipient.  However, it is

12   not so clear that such a donor should be

13   permanently deferred from donating any blood

14   component.

15             The appropriate preventive measures for

16   TRALI are even less obvious for the majority of

17   pulmonary reactions that occur in the transfusion

18   setting.

19             It is also important to understand what

20   proportion of the donor population would be

21   affected by proposed deferral criteria or other

22   regulatory strategies, so that the potential impact

37

1   on the blood supply can be evaluated. These data

2   are especially critical, as we already too

3   frequently face blood shortages in regions across

4   the country.

5             A careful and thorough analysis of the

6   risks and benefits of any donor deferrals or any

7   other regulatory strategy must be completed before

8   taking steps that could unnecessarily hinder

9   patient access to life-saving blood components.

10             Thank you.

11             DR. ALLEN:  Thank you very much.

12             Questions or comments from the committee

13   in response?  Yes.

14             DR. KLEIN:  We have heard on a couple of

15   occasions now about the Canadian Consensus

16   Conference and clearly it's an important one, but

17   the results haven't been published yet, and I would

18   certainly caution the FDA about the definition that

19   has been proposed.  It's a preliminary definition.

20             Many of the patients that we take care of

21   are in intensive care units, they are on

22   respirators, they do have some kind of underlying

38

1   lung disease, and they get a lot of blood

2   transfusions.  By the definition that has been

3   proposed, should any of them have what looks like

4   TRALI, they would be excluded under the proposed

5   definition.

6             I am not sure that is the permanent

7   definition.  I think we ought to wait before

8   adopting anything to see what the publication says.

9             DR. ALLEN:  Thank you.  I think that is

10   very good advice.

11             Other questions or comments?  Yes.

12             DR. HARVATH:  I would like to just address

13   a couple of points about the recommendation for

14   supportive research from the NHLBI perspective.

15   There have been a number of ways we have been

16   trying to stimulate this during the past several

17   years.

18             One of the ways that we are going about

19   doing this is through the transfusion medicine

20   hemostasis clinical trial network, which is a

21   multi-center, 17 clinical centers throughout the

22   United States.

39

1             We have discussed with that committee

2   looking at prospectively any study which involved

3   the transfusion of components, and almost every

4   study does, looking prospectively to find any

5   evidence of TRALI in the patients in those studies,

6   so it will be the opportunity to look at both a

7   platelet transfusion study, which is our first

8   study, and possibly a second study that would

9   potentially involve FFP, and these would be

10   randomized studies, and the work would be done

11   prospectively.

12             The second point is that the NHLBI also

13   funds a multi-center acute respiratory distress

14   network, which involves the pulmonary specialists,

15   and they have become interested in this area, so

16   there are investigators who are also interested in

17   looking in that patient population.

18             So, these are existing clinical trial

19   networks where this would be possible to integrate

20   this type of research, and also to add that NHLBI

21   welcomes any investigator-initiated studies to come

22   forward to the institute and to let us know what

40

1   kinds of research investigators or groups of

2   investigators would like to pursue.  So, we are

3   very open to that.

4             DR. ALLEN:  Thank you.  It is certainly

5   helpful to have both lung and blood in the same

6   institute from that perspective, I am sure.

7             Other questions or comments?

8             Okay.  We will move on to the next

9   statement on blood pressure lower limits by the

10   AABB.  I am sorry, excuse me, we do have one

11   additional statement on TRALI, Dr. Fitzpatrick from

12   the America's Blood Centers.

13             DR. FITZPATRICK:  Mike Fitzpatrick, Chief

14   Policy Officer for America's Blood Centers, and I

15   am employed by them.

16             America's Blood Centers, or ABC, is an

17   association of 76 not-for-profit, community-based

18   blood centers that collect nearly half of the U.S.

19   blood supply from volunteer donors.  ABC thanks

20   FDA's Center for Biologics Evaluation and Research

21   for the opportunity to make public comments before

22   the Blood Products Advisory Committee.

41

1             Our members share FDA's concerns about

2   transfusion related acute lung injury.  While rare,

3   this is a serious and sometimes fatal

4   transfusion-associated event.  We know that TRALI

5   is a complex phenomenon, and there is no agreement

6   in the published literature about the major

7   mechanisms of disease.

8             This was clearly documented at the

9   Canadian Consensus Conference that we have heard

10   about.

11             At least two mechanisms appear to play a

12   role, one involving antibodies to leukocytes, the

13   other involving biologically active mediators.

14   Interestingly enough, in the paper published by

15   Silliman--I won't quote the source here, but you

16   have got that--most of the TRALI events appear to

17   be related to biologically active mediators and

18   only one of the 90 reactions studied involved a

19   plasma unit.

20             Most reactions, 74, involved whole blood

21   derived and apheresis platelets.  Kopko has

22   indicated that many units implicated in TRALI

42

1   reactions carry antibodies to white blood cells.

2   However, she concluded from her studies that HLA

3   antibodies in a donor corresponding to HLA antigens

4   in a recipient are not sufficient to cause TRALI in

5   all recipients.

6             She also noted that based on lookback

7   studies, donors implicated in TRALI reactions can

8   cause TRALI in other recipients, regardless of

9   antigen-antibody correlations.  While presentations

10   also indicated a higher rate of female plasma

11   donors who have been pregnant carry anti-HLA

12   antibodies, data is lacking that would establish a

13   definitive link between gender and/or anti-HLA

14   antibodies and TRALI.

15             Dr. Holness from FDA presented the FDA

16   fatality data at that conference and a summary of

17   the data today here.  He showed an apparent

18   increase in TRALI associated fatalities in recent

19   years.  He also indicated that the majority of the

20   49 fatalities that occurred between 2001 and 2003

21   were associated with plasma transfusions.  The

22   number or percent was not indicated.

43

1             The donor data presented did not include

2   donor gender or prevalence of antibodies to

3   leukocytes, so we cannot estimate the impact of the

4   three preventive strategies enumerated by FDA:

5   only transfuse plasma containing components from

6   male donors, perform preventive antibody testing,

7   defer donors implicated in TRALI cases.

8             We agree that FDA should review and

9   consider interventions to address the issue of

10   TRALI.  The impacts of such strategies must also be

11   considered by asking the following questions:

12             How many TRALI associated fatalities will

13   be prevented by the implementation of each

14   strategy?  What blood components should be included

15   in the strategy?  TRALI has been associated with

16   all blood components, including red blood cells,

17   apheresis platelet units, which contain as much or

18   more plasma than a unit of fresh frozen plasma.

19             What impact will this have on the

20   availability of components?  Are there other

21   strategies that could be considered?

22             The data presented by FDA, the current

44

1   literature, the recommendations made by BPAC in

2   2001 and the conclusions of the Canadian Consensus

3   Conference, while not yet published, that were

4   summarized at the meeting, do not provide a clear

5   basis for any of the regulatory strategies listed.

6   Whole blood, whole blood derived platelets,

7   apheresis platelets, and plasma have all been

8   implicated in TRALI.  Why restrict the approach the

9   plasma, what about apheresis platelets?

10             We carried out a survey to assess the

11   impact of using only male plasma and platelet

12   apheresis products among ABC members.  Forty-two

13   centers collecting a total of almost 4 million

14   whole blood and apheresis units a year responded.

15             Based on the gender distribution of ABC

16   donors, we estimate that a ban on female plasma and

17   apheresis platelets would lead to the loss of

18   113,000 donors and 275,000 donations in one year.

19   If we double this estimate to include collections

20   by the American Red Cross, 550,000 donations would

21   be lost in the U.S.

22             Females represent about 44 percent of all

45

1   apheresis donors. Our members indicated that they

2   could not effect these changes without seriously

3   impairing product availability.  When our members

4   were asked whether they could provide male plasma

5   only to their hospitals, 55 percent responded yes.

6             However, they indicated that it would take

7   them between 18 and 24 months to implement the

8   changes, including software modifications, and that

9   the change would create serious shortages of type

10   specific plasma, particularly type AB.

11             ABC members disagree with FDA's point of

12   view that strategy number 3, deferral of donors

13   implicated in TRALI incidents, is inadequate

14   because it allows for the first incident to occur

15   before donor deferral is instituted and does not

16   eliminate TRALI.

17             Unfortunately, all the proposed strategies

18   suffer from this deficiency because of the myriad

19   causes of TRALI.  Strategy 1 addresses an

20   undetermined fraction of TRALI cases and has more

21   serious consequences for blood availability.

22             At the present time and with the present

46

1   knowledge, regulatory action should be restricted

2   to donors implicated in TRALI episodes, as stated

3   in the third strategy.

4             FDA also needs to support effective

5   training of physicians and other hospital personnel

6   for early recognition of TRALI, based on the case

7   definition being considered by an NHLBI task force,

8   which was not discussed this morning, under the

9   leadership of Dr. Pearl Toy.  This may be more

10   efficient in the prevention of fatalities than any

11   of the proposed strategies.

12             The implementation of a global strategy

13   such as the deferral of male donors may have other

14   adverse consequences.  It may convey to the medical

15   community and to the public the erroneous

16   impression that the problem of TRALI has been

17   addressed and resolved, leading physicians to

18   consider other diagnoses and prescribe

19   inappropriate therapy.

20             Finally, we will have to deal with the

21   frustration of female donors when they learn that

22   their donations are not good for transfusion.

47

1             ABC members thank FDA and the BPAC for the

2   opportunity to comment.

3             Thank you.

4             DR. ALLEN:  Thank you very much.

5             Questions or comments with regard to Dr.

6   Fitzpatrick's presentation?  Jay.

7             DR. EPSTEIN:  I just want to comment that

8   this was an informational update, and the intent

9   was more to get the issue on everybody's radar

10   screen than to propose action at this point in

11   time.  I think that Dr. Holness' presentation made

12   clear that we are aware of all the uncertainties

13   and the ambiguities.

14             It is also true that the UK, faced with

15   the same uncertainties and ambiguities, felt that

16   action should be taken and it has its pros and

17   cons, so this is not rush to judgment and I

18   appreciate all the cautionary notes that have been

19   sounded, but I think that, you know, we have been

20   living with awareness of TRALI without effective

21   intervention for some time, and the idea here is to

22   provoke ourselves to think about could we be doing

48

1   more and what should that be.  So, this is just an

2   early stage of thinking.

3             DR. FITZPATRICK:  We understand, Jay, and

4   we just appreciated the opportunity to make some

5   comments and present some questions.

6             DR. ALLEN:  Other questions or comments

7   from the committee members?

8             Dr. Davis.

9             DR. DAVIS:  I would like to speak as

10   somebody that treats a lot of people with acute

11   lung injury.  TRALI is not something that is on

12   most of our radar screens.  Most of the people that

13   have acute lung injury, if you look at the list of

14   risk factors, most of those people, as Dr. Klein

15   alluded, get multiple transfusions for a lot of

16   other reasons.

17             The other thing that I really haven't

18   heard, and I don't know if there is an answer to

19   the question, is what is the survival rate.  I mean

20   we have heard what the fatality is, but how many

21   people get TRALI and actually survive.

22             I think it is going to be hard to isolate

49

1   those kinds of isolated transfusion-related

2   injuries.  I am not sure how many clinicians are

3   actually aware of TRALI.

4             DR. ALLEN:  I think those are very good

5   points and certainly go right along with what Dr.

6   Harvath was saying about the need for prospective

7   multi-center studies.

8             The comment was made earlier today also

9   about the definition of TRALI, and it sounds to me,

10   with two proposed mechanisms in place, that we may

11   actually be dealing with multiple different

12   clinical events that need to be teased apart and

13   separated, and it sounds to me as though there is a

14   lot of research that needs to be done.

15             It is an important issue given the

16   relative incidence in terms of serious events

17   related to transfusion. I am sure that we are not

18   ready for any regulatory consideration at this

19   point.

20             We will look forward to additional updates

21   and research findings.

22             Other questions or comments?

50

1             [No response.]

2             DR. ALLEN:  Okay.  Thank you.

3             We will move on the statements with regard

4   to blood pressure lower limits.  AABB.

5             Could I ask you not to read the first

6   paragraph, please, and just to move on with the

7   statement itself. Thank you.

8             MS. GREGORY:  Thank you.  I had every

9   intention of doing that.

10             I also wanted to make the committee aware

11   that I am speaking, not only on behalf of the AABB,

12   but I am also speaking on behalf of America's Blood

13   Centers.  Your written statements don't reflect

14   that simply because of the need to get it in

15   quickly, so that you could have it ahead of time,

16   but I am speaking for both organizations.

17             Neither the AABB nor ABC supports the need

18   for a lower limit for blood pressure for blood

19   donors.  Blood collection facilities have had only

20   upper limits for blood pressure in place for many

21   years.

22             The AABB Standards for Blood Banks and

51

1   Transfusion Services requires that the blood

2   pressure be 180 systolic and 100 diastolic.  These

3   levels have been the requirement since 1987.  This

4   particular standard was reviewed again in 2002 and

5   again in 2003, and the Blood Banks and Transfusion

6   Services Standards Program Unit found no scientific

7   evidence to warrant changing the standard.

8             I also want to explain the difference

9   between the AABB standards and the AABB technical

10   manual because the written materials that you

11   received talked about statements that are in the

12   technical manual.

13             The AABB standards are where the

14   requirements are stated, and they include

15   requirements for both quality management and

16   technical requirements.  The technical manual is

17   published to provide background material, some

18   guidance, and methods and procedures, but does not

19   include requirements.

20             The technical manual may provide practices

21   that will assist facilities in implementing

22   standards, but the standards is the definitive

52

1   document.

2             Another reason why we do not see a need

3   for a lower limit for blood pressure is that we

4   know that blood pressure is not a requirement for

5   donor qualification in the latest European Union

6   Commission directive.

7             The Council of Europe Guide states:  If

8   pulse and blood pressure is tested, then the pulse

9   should be regular and between 50 and 100 beats per

10   minute.  It is recognized that recording the blood

11   pressure may be subject to several variables, but

12   as a guide, the systolic blood pressure should not

13   exceed 180 millimeters of mercury and the diastolic

14   pressure 100 millimeters.

15             A review of medical textbooks revealed

16   that there is no consistency about what is

17   considered to be hypertension in asymptomatic

18   individuals, and that a low blood pressure is not a

19   matter of great concern or interest outside of the

20   emergency room or intensive care settings.

21             A number of researchers have published

22   articles in peer-reviewed journals showing a lack

53

1   of correlation between low pre-donation systolic or

2   diastolic blood pressure and adverse donor

3   reactions.

4             A 2002 study of 72,059 whole blood

5   donations at the American Red Cross showed no

6   statistical association between low pre-donation

7   systolic or diastolic blood pressure and adverse

8   reactions.

9             In addition, the American Red Cross

10   reviewed pre-donation blood pressure on all donors

11   with adverse reactions that resulted in

12   hospitalization from January of 1999 to December of

13   2002.  This review showed no over-representation of

14   low blood pressure in those donors.

15             Finally, a review of donor fatality

16   reports obtained under the Freedom of Information

17   Act showed no low pre-donation blood pressure

18   either.

19             There are two Code of Federal Regulation

20   requirements that FDA has quoted as the rationale

21   for adding a lower limit for blood pressure.  21

22   CFR 640.3(b)(2), which states that systolic and

54

1   diastolic blood pressure must be within normal

2   limits, and 606.100(b)(2), which states that the

3   standard operating procedures for donor-qualifying

4   tests and measurements must specify maximum and

5   minimum values.

6             It is unclear why FDA has recently chosen

7   to selectively enforce this particular requirement

8   for blood pressure.  There are other

9   donor-qualifying tests and measurements that do not

10   have both upper and lower limits. For example,

11   temperature has only an upper limit, and weight,

12   hemoglobin, and age only a lower limit.

13             We have already noted the lack of uniform

14   agreement as to what constitutes a low blood

15   pressure in asymptomatic individuals.  In short,

16   while there may be a regulation that can be cited

17   as justification for this change in policy, the

18   regulation has not been enforced in the past and a

19   change in policy is unnecessary.

20             A key element of the FDA's 2004 strategic

21   action plan is efficient risk management.  This

22   plan states that in all of its major policies and

55

1   regulations, FDA is seeking to use the best

2   biomedical science, the best risk management

3   science, and the best economic science to achieve

4   health policy goals as efficiently as possible.  A

5   change to the requirement for donor blood pressure

6   does not meet these criteria.

7             DR. ALLEN:  Thank you.

8             Questions or comments?  Yes.

9             DR. DiMICHELE:  Thank you for that.  It

10   seems to me that the question at hand here is

11   whether a low blood pressure is physiologic for the

12   individual or whether it might represent

13   dehydration or for vasomotor tone and inability to

14   vasoreact in the face of acute volume reduction,

15   those kinds of issues.

16             It appears based on the epidemiologic data

17   that most of it isn't.  However, when you speak

18   about asymptomatic hypotension or asymptomatic

19   blood pressure, low blood pressure, do you--remind

20   me, I should know this because we have looked at

21   those criteria and those questions time and time

22   again--but do you ask a question in the pre-donor

56

1   screening about symptomatic hypotension or

2   symptomatic low blood pressure in the pre-donation

3   screening questionnaire?

4             MS. GREGORY:  We don't ask that

5   specifically, but we do ask things like are you

6   being treated by a doctor, things that we think

7   would elicit that information, but not that

8   specific question.

9             DR. DiMICHELE:  And the second question I

10   have is again, it is obvious that if the blood

11   pressure is physiologic for the individual, it is

12   probably not going to tend to be pathologic in any

13   way in donation, so do you have a way for multiple,

14   when you have repeat donors, to actually track

15   their blood pressures over time and to be able to

16   identify a low blood pressure that might be

17   unphysiologic for that individual?

18             MS. GREGORY:  The blood pressure is

19   recorded at each donation, and we do keep those

20   records, so I think there probably would be a way

21   to track that if we needed to.

22             DR. DiMICHELE:  So, really, basically, a

57

1   decrease in routine blood pressure that wasn't

2   previously hypotensive or certainly symptomatic

3   hypotension might be ways of picking up symptoms

4   without necessarily initiating a lower limit.

5             MS. GREGORY:  Thank you.

6             DR. ALLEN:  Let me just clarify, though, I

7   don't think when a person comes in to donate blood,

8   the information is obtained for that donation only,

9   and I don't think they go back and look at a

10   sequence of past blood pressure determinations.

11             Certainly, a change in laboratory values

12   might be noted, but I don't think that they would

13   go back and look at the pre-, you know, they don't

14   have pulled up on a computer screen or a paper

15   record that would show what the blood pressure

16   determinations were at the last two or three

17   donations.

18             DR. DiMICHELE:  That is what I was asking,

19   if that information was readily available.

20             MS. GREGORY:  That's right, we would not

21   look at it right then, but we would have the

22   ability to look at it should we think there is a

58

1   need to look at it for some reason.

2             DR. ALLEN:  Thank you.  Other questions or

3   comments?  Yes, Dr. Williams.

4             DR. WILLIAMS:  Kay, a comment and a

5   question.

6             This was characterized as a change in

7   regulatory policy.  I think that perhaps isn't

8   correct.  It might be a rift in communication

9   particularly with respect to the industry voluntary

10   standards, but the question is what is the risk

11   side of the equation.

12             We take the point that regulation should

13   be scientifically based, but what is the impact on

14   blood collection?  I would pose the same question

15   to the PPTA speaker.  What is the donor loss, what

16   are the operational implications of recording a

17   lower blood pressure?  What is the impact?

18             MS. GREGORY:  I think the operational

19   limitations are that we already record everything

20   under the sun, and recording one more thing might

21   not seem like it would be that difficult, but it is

22   one more chance to record it wrong and you have to

59

1   keep track of it all, and it is not that it is

2   impossible to do, it is just we would like to be as

3   efficient as we possibly could, and we don't think

4   there is a reason for recording this.

5             DR. ALLEN:  Thank you.

6             We have a written statement also from the

7   Plasma Protein Therapeutics Association.  Is there

8   a need to read that, do we have a speaker or a

9   proposed speaker?  Okay. Thank you.  I will just

10   note for the record that there is a written

11   statement from PPTA also.

12             The open public hearing is now closed.

13             We will move on to the next item on the

14   agenda, which is an open committee discussion of

15   Topic I, Dating of Irradiated Red Blood Cells.  We

16   have a number of speakers and we will plan to spend

17   the rest of the morning on this discussion,

18   concluding with a series of questions for the

19   committee.

20             The first speaker with Introduction and

21   Background is Dr. Ping He with the Food and Drug

22   Administration.

60

1             I. Dating of Irradiated Red Blood Cells

2                   Introduction and Background

3                          Ping He, M.D.

4             DR. HE:  Good morning.  I am going to talk

5   about the introduction and background of dating

6   period for gamma irradiated red blood cells.

7             [Slide.]

8             Why do we have the irradiated blood

9   product?  The answer is that the irradiation of

10   blood products can prevent transfusion associated

11   graft-versus-host disease, GVHD.  GVHD occurs when

12   viable cytotoxic allogeneic lymphocytes are

13   transfused to a recipient unable to reject them and

14   cause disease.

15             Patients at risk are neonates,

16   immunocompromised patients for different reasons,

17   and the recipient is genetically related to the

18   blood donors.

19             [Slide.]

20             Here is the clinical pathological features

21   of GVHD.  GVHD is a rare batch of very fatal

22   complication of transfusion associated disease.

61

1   All cells with HLA antigens are affected.  The

2   cause of disease starts with the lymphocytes from

3   donor, from the transfused blood and graft into

4   recipient.

5             These transfused donor lymphocytes can

6   then proliferate and damage the target organs, such

7   as bone marrow, skin, gastrointestinal tract, and

8   liver.  The symptoms may appear two to 30 days

9   after blood transfusion with skin rash, diarrhea,

10   liver enzyme elevation and pancytopenia.

11             The occurrence is about 0.1 to 1 percent

12   with no effective therapy.  The mortality is high,

13   usually close to 100 percent.

14             [Slide.]

15             Fortunately, GVHD can be prevented by

16   gamma irradiation of blood products prior to

17   transfusion to inactivate the donor lymphocytes

18   that cause the disease.

19             The blood products can be irradiated if

20   the recipient is immunocompromised or the blood

21   donor is genetically related to the recipient.

22             [Slide.]

62

1             The advantage of gamma irradiation of

2   blood products is that it can prevent GVHD,

3   however, the disadvantage of irradiation of red

4   blood cells can cause the decrease of

5   post-transfusion of red blood cell recovery and

6   increase the leaking of intracellular potassium.

7   This raises the concerns for the safety and

8   efficacy of the irradiated red blood cells.

9             Therefore, the irradiated red blood cells,

10   the storage period should be limited.  This issue

11   has received attention in the past and it was the

12   subject of a 1992 NIH workshop, and in 1994, a BPAC

13   discussion.

14             In July 1993, a memorandum was issued by

15   FDA and it recommended that the irradiation dosage

16   for RBC should be 2,500 centigray in the center

17   portion of the container and 1,500 centigray in the

18   other point.  The dating period for RBC should be

19   not more than 28 days from the date of the

20   irradiation, but not more than the dating period of

21   the original product.

22             This means that if an adequate solution

63

1   allows RBCs to be stored up to day 42.  The RBCs

2   can be irradiated anytime from day 1 to day 42

3   after collection.  If the irradiation happens from

4   day 1 to day 14, then, the irradiated RBCs will

5   have an additional 28 days for the storage.

6             [Slide.]

7             If the irradiation happened after the day

8   14 of the collection, then, the irradiated RBCs can

9   be stored up to day 42.

10             [Slide.]

11             Here is the BPAC 1994.  FDA proposed

12   changing the dating period of irradiated red blood

13   cells based on updated data from American Red Cross

14   and NIH obtained since 1993. However, the committee

15   at that time recommended no change from 1993

16   memorandum, but the committee did suggest that the

17   criteria might be reconsidered should future data

18   become available.

19             In February 2000, FDA issued a Guidance

20   for Industry titled as Gamma Irradiation of Blood

21   and Blood Components: A Pilot Program for

22   Licensing.

64

1             This guidance recommended the same dating

2   period as it was stated in 1993 memorandum.

3             [Slide.]

4             Today, we would like to revisit this issue

5   again for following reasons:  In view of potential

6   safety and efficacy concerns for RBCs irradiated in

7   new anticoagulant and additive solutions, FDA

8   requests that sponsors should submit data to

9   support licensure.

10             Recent data from the sponsors raised

11   concerns on the efficacy of RBCs irradiated in the

12   new anticoagulant and additive solutions.  I am

13   going to briefly summarize some data to explain our

14   concern.

15             In addition, we also have concerns about

16   the FDA's current criteria for in vivo RBC

17   recovery.

18             [Slide.]

19             Here is the FDA current recommendations

20   for the in vivo RBC recovery evaluations.  This

21   evaluation is to test the ability of RBC products

22   to circulate after autologous infusion of

65

1   radiolabeled RBCs.

2             The viability of RBC is assessed by

3   determining the RBC recovery 24 hours after

4   infusion of autologous cells.  Usually, we look for

5   equal or greater than 75 percent recovery.

6             The current criteria recommends that the

7   study site should be 20 to 24 in more than one site

8   with standard deviation of less or equal to 9

9   percent.  If the sample mean is equal or greater

10   than 75 percent, then we will have a 95 percent

11   lower confidence limit greater than 70 percent.

12             This criteria will be further discussed by

13   Dr. Kim, the mathematical statistician from FDA

14   later this morning.

15             [Slide.]

16             Here is the study results of sample

17   failure 28 days after irradiation from Manufacturer

18   A.  Sample failure is defined as less than 75

19   percent recovery after reinfusion.

20             The sponsor conducted two groups of

21   studies.  The first group, the RBCs irradiated on

22   day 1, evaluated on day 28.  The second group, the

66

1   RBCs irradiated on day 14, and evaluated on day 42.

2             The study was carried out at three

3   different sites.  For Site 1 and Site 2, the RBCs

4   were irradiated for 500 centigray, and for Site 3,

5   the RBCs were irradiated for 3,000 centigray.

6             As you can see, the RBCs irradiated on day

7   1 and evaluated on day 28, there was 3 out of 22

8   failures, about 14 percent failure, but RBCs

9   irradiated on day 14 and evaluated on day 42, there

10   were 5 out of 21 failures, about 24 percent

11   failures.

12             [Slide.]

13             Here is another example of sample failure

14   28 days after irradiation from Manufacturer B.  In

15   this study, the sponsors only did studies of RBCs

16   irradiated on day 14, and they were evaluated on

17   day 42.

18             The studies were carried out in two

19   different centers, and all the RBCs are irradiated

20   at 2,500 centigray.  Two groups of studies done,

21   one is for tests, the other group is for control.

22   The test units, as you can see, there were about 21

67

1   percent failure, and for control group, there were

2   11 out of 24, about 46 percent of failure.

3             [Slide.]

4             So, from this data, we find out the longer

5   the storage or the later the irradiation, the

6   higher the failure rate.

7             [Slide.]

8             Here is the dating period of irradiated

9   red blood cells in Council of Europe Guide that was

10   published in the year 2003 just for the reference

11   here.

12             Here is the direct quote.  "Red cell

13   products may be irradiated up to 14 days after

14   collection and thereafter stored until the 28th day

15   after collection."

16             "In view of the increased potassium leak

17   post irradiation, intrauterine or massive neonatal

18   transfusion should be used within 48 hours of

19   irradiation."

20             [Slide.]

21             Here are the issues for which we seek

22   advice from the committee.

68

1             [Slide.]

2             I am going to present the questions right

3   now just to highlight the issues we are going to be

4   focusing on this morning.  These questions will be

5   re-presented during the open committee discussion.

6             Questions on Dating Period of Gamma

7   Irradiated Red Blood Cells.

8             Question No. 1.  Do the committee members

9   agree that the current recommendations regarding

10   the dating period for gamma-irradiated red blood

11   cells should be modified?

12             [Slide.]

13             Question No. 2.  If you do agree to

14   modify, please comment whether the available

15   scientific data support the following candidate

16   modifications to FDA's current guidance on

17   irradiated RBCs.

18             a.  For RBC products that are irradiated

19   within the first 26 days after the date of

20   collection, the products should not be stored more

21   than 28 days from the date of collection.

22             b.  For RBC products that are irradiated

69

1   on or after 26 days from the date of collection,

2   the post-irradiated products should be stored no

3   more than 48 hours after irradiation.

4             [Slide.]

5             Question No. 3.  Does the committee have

6   any alternative modifications to FDA's current

7   guidance regarding the dating period for

8   gamma-irradiated red blood cells that should be

9   considered.

10             [Slide.]

11             Question No. 4.  Here are the questions on

12   RBC in vivo recovery acceptance criteria.  Please

13   comment on the following options:

14             a.  Keep the current criteria, which is

15   sample mean equal or greater than 75 percent,

16   standard deviation equal or less than 9 percent,

17   and 95 percent lower confidence limit for the

18   population mean above 70 percent.

19             b.  Proposed new criteria 1:  We propose

20   sample mean equal or greater than 75 percent,

21   standard deviation less or equal than 9 percent,

22   and a 95 percent one-sided lower confidence limit

70

1   for the population proportion of successes greater

2   than 70 percent.  The success is defined as greater

3   or equal than 75 percent of RBC recovery.

4             c.  Proposed new criteria 2:  A 95 percent

5   one-sided lower confidence limit for the population

6   proportion of successes greater than 70 percent and

7   a minimum individual recovery of all samples equal

8   or greater than 60 percent.  These criterias will

9   be further discussed by Dr. Kim also in the later

10   session.

11             That is all I am going to say now.

12             DR. ALLEN:  Thank you, Dr. He.

13             Comments or questions relative to the

14   presentation?  Yes.

15             DR. KLEIN:  I have two questions.  The

16   first is on your slide about the Council of Europe,

17   that first recommendation, since I am not sure

18   other speakers will be addressing that, do you know

19   whether their standard is based on any data or on

20   what data their standard is based?

21             My second question is you are raising this

22   issue at this time for the committee.  Is there a

71

1   problem with the current standard, and, if so, what

2   is it?

3             DR. HE:  Well, I am not very sure about

4   whether the Council of the European Guide is best

5   on data, but that is how their recommendation, it

6   is just for the reference.

7             The second, even though we don't have any

8   adverse event report from the gamma-irradiated

9   product, however, from one of the control, you can

10   say that there was almost 46 percent of the

11   failure, which is of quite concern for the efficacy

12   and safety, probably mostly efficacy concerns.  We

13   also feel that the data that is going to be

14   presented later today probably will explain some of

15   the concerns we have.

16             DR. ALLEN:  Yes.

17             DR. KUEHNERT:  I am a little confused

18   about what you just said.  You said a 46 percent

19   failure, but isn't that the control group?

20             DR. HE:  Yes, that is control.  That

21   control is actually, it's a blood bag with

22   anticoagulant additive solutions cleared by FDA

72

1   years ago.  That clearance, at that time, it was

2   cleared for collecting of the whole blood and for

3   the processing of the red blood cells.

4             At that time, the bag, during the

5   clearance, there was no gamma-irradiated studies

6   done, but the blood centers just used that bag to

7   collect the blood and to irradiate the blood, sort

8   of like historical without any study at that time.

9             DR. KUEHNERT:  But the control group,

10   there was no irradiation, hence, the term control

11   group, and I guess I am confused about the concern

12   there.  Is the concern about the anticoagulation,

13   anticoagulant that is used or is it--it's not about

14   the irradiation because that was the control group

15   actually, right?

16             DR. HE:  Perhaps not because of

17   irradiation, perhaps because of any unit, no matter

18   if it's the old anticoagulant or new anticoagulant

19   or in--we have kind of a collecting bag, that

20   irradiation should not happen later than day 14 or

21   should not be stored up to 42 days.

22             DR. VOSTAL:  Maybe I can help out here.

73

1   The reason we are looking at this issue right now

2   is that we received an application from

3   manufacturers that brought to us novel combination

4   of anticoagulants that were never tested after

5   irradiation.

6             So, we looked at that data and we realized

7   that we have never seen those before.  We looked at

8   the combinations.  We realized we had never seen

9   the data with irradiated cells, so we requested

10   that.  The data came in and we are in the process

11   of evaluating them right now.

12             Part of that submission, we looked at the

13   control cells, which actually I think were

14   irradiated, and we looked that there is a high rate

15   of failure in the control cells, control cells

16   being currently approved products.

17             So, we are looking at this in terms of

18   whether the current standard is appropriate or

19   whether it should be changed.  We are going to see

20   that control data again when the company presents

21   their data on their own.

22             DR. ALLEN:  Let me just clarify.  The

74

1   control cells had been irradiated?

2             DR. VOSTAL:  It is my understanding that

3   they were irradiated cells.  You are comparing a

4   currently approved irradiated product to a novel

5   combination irradiated product.

6             DR. ALLEN:  Maybe that will be clarified

7   when we get the presentation further.

8             Dr. Leitman.

9             DR. LEITMAN:  This 75 percent standard

10   applies not only to products that are modified by

11   irradiation post collection, but applies to the FDA

12   evaluation of any new storage vehicle, new bag, new

13   anticoagulant.  So, you gave us three statistical

14   options to choose from.

15             That doesn't only apply to radiation, to

16   everything.  It's a separate issue, how does the

17   FDA decide that any modification to collection or

18   storage vehicle or treatment of the component is

19   okay.  Is that what we are being asked to look at?

20             DR. VOSTAL:  That's correct.  The standard

21   for red cell performance is 75 percent recovery at

22   24 hours of radiolabeled cells.

75

1             Now, the options that we have presented

2   are designs of the studies that are used to

3   evaluate those products, and you can either look at

4   the average performance of the study volunteers and

5   average it out to see if they are greater than 75

6   percent, and the issue there is if you allow that,

7   sometimes you see data that has a number of

8   failures, but still meets the average greater than

9   75 percent.

10             Some of these studies that we have been

11   looking at pointed to us that you have studies that

12   4 out of 20 fail even though they meet the 75

13   percent average, and we were wondering whether that

14   is appropriate.

15             The other way of looking at it is you

16   could fix the percentage of those donors that meet

17   the criteria, you know, fix the proportion, and

18   those are some of the other options.

19             You can say 80 percent of those donors

20   have to meet the 75 percent criteria, you know,

21   instead of looking at their average.  So, those are

22   the options that we are going to be discussing.

76

1             DR. ALLEN:  Dr. Klein.

2             DR. KLEIN:  If I could just follow up on

3   my question.  Since one of the options you are

4   going to ask us to look at is harmonization with

5   the European standard, do you know what data those

6   are based on, and anticoagulant bags, all of the

7   various things that we are concerned about since we

8   obviously wouldn't want to harmonize something that

9   doesn't make any sense?

10             DR. VOSTAL:  Right.  I also don't know the

11   data for that decision in the Europe.

12             Could I just make one more point?  We

13   would like to just make sure that you understand

14   that non-irradiated products do meet the 75 percent

15   criteria, so we are not questioning the 75 percent,

16   only the design of the studies that is used to

17   evaluate that.

18             DR. ALLEN:  Dr. Strong.

19             DR. STRONG:  Could I ask where the 75

20   percent number came from, is that based on

21   scientific evidence?

22             DR. VOSTAL:  I think it was decided back

77

1   in 1982, and that was a consensus decision, and it

2   was more or less an arbitrary decision.

3             DR. ALLEN:  I will just comment that I was

4   astounded as I looked through the materials in

5   preparation for the meeting at the extraordinarily

6   wide range.  You know, 75 percent as a mean may or

7   may not have scientific validity, but for any given

8   donor, there was an extraordinarily wide range, and

9   that surprised me.  I had not anticipated that.

10             DR. LEITMAN:  This is deja vu.  This

11   conversation was held in 1994 at the BPAC

12   committee, and remember the consensus agreement

13   then was the discussion about the 75 percent, what

14   if you have 73 percent, 71 percent, 68 percent, how

15   much is enough for recovery of a red cell

16   component, and the comment was that if you don't

17   treat the components in that manner, and the option

18   is graft-versus- host disease, then, 68 percent is

19   good, it's acceptable.

20             So, a product treated in such a manner

21   that it avoids a fatal reaction, at two-thirds, I

22   think that comment was are available for the

78

1   patient, that's a good outcome, so it is quite

2   arbitrary.

3             DR. EPSTEIN:  There is another reverse

4   side to that argument, Susan, what you say is

5   certainly correct, but how often do you need to

6   irradiate an older unit and then hold it a long

7   time.

8             I mean as a practical matter, most units

9   are irradiated earlier in their shelf life, and

10   what we are really saying is if you have a need to

11   irradiate an older unit, just don't store it a long

12   time, and that will not be a frequent situation.

13             What we are really saying is we can

14   prevent a potential harm, we can improve the

15   quality of the delivered product.  We don't really

16   think that it is going to compromise product

17   availability or the availability of a given unit.

18             So, I understand what you are saying and I

19   agree, but the reverse side is that this is not

20   impractical.  There will be those who comment that

21   it is a recordkeeping issue and relabeling dating

22   is nightmarish, and we are going to hear that

79

1   argument.

2             DR. ALLEN:  Thank you.  Go ahead.

3             DR. GOLDSMITH:  Do we really know the

4   implications of a safety and efficacy point of view

5   of infusing these units that have lower than 75

6   percent recovery?  Do we actually have clinical

7   information on outcomes?

8             DR. VOSTAL:  I think that is very

9   difficult to assess.  I mean in a single

10   individual, transfusing cells that would not meet

11   the 75 percent probably would not make that much of

12   a difference unless there was a really damaged

13   unit.

14             But I think we are trying to apply this

15   across the whole population, and if you would

16   accept the percentage, a lower recovery for all

17   products, you would end up having decreasing

18   availability of the blood supply, and you might end

19   up in transfusing more frequently, which would also

20   decrease the availability of the blood supply.

21             DR. ALLEN:  Why don't we move on.  We have

22   got other speakers and then we will come back to

80

1   general discussion.  It sounds as though it is

2   going to be a fairly significant discussion.

3             Our next speaker is Dr. Gary Moroff,

4   Holland Laboratory, American Red Cross.

5                Presentation - Gary Moroff, Ph.D.

6             DR. MOROFF:  Thank you very much.

7             What I want to present today is our

8   historic data, because it was generated in the

9   1990s, early to mid-'90s, dealing with gamma

10   irradiation of whole blood derived ADSOL red cells.

11             [Slide.]

12             So, our study objective was to evaluate

13   the influence of gamma irradiation dose which we

14   deemed optimal, and I will talk about this in a few

15   seconds, for inactivating T cells in red cell units

16   on red cell properties with different scenarios for

17   time of irradiation and total storage time, and

18   this relates to the discussion over the last five

19   or 10 minutes.

20             [Slide.]

21             Basically, I am going to start with

22   talking about the 2,500 centigray issue because we

81

1   conducted studies before we did our red cell

2   studies to show that this was the optimal dose, and

3   then I will talk about the red cell property

4   studies.

5             [Slide.]

6             Just for review, these are the primary

7   types of instruments being used to irradiate blood.

8   There are free-standing irradiators with a cesium

9   or cobalt source, and also at hospitals, linear

10   accelerators are used, and the linear accelerators

11   are based on x-rays.

12             The irradiation effect is the same whether

13   it is the cesium source or the x-ray source when

14   you think in terms of total dose.  There is

15   currently now a free-standing x-ray machine that is

16   available also for blood units.

17             [Slide.]

18             Let me skip about the basics because that

19   was covered, but when we started thinking about

20   gamma irradiation in the early 1990s, we realized

21   that the optimal dose had not been identified in

22   appropriate studies with red cell units.

82

1             In collaboration with Dr. Luban and Dr.

2   Quinones at the Children's Hospital in Washington,

3   we identified 2,500 centigray as the optimal dose,

4   and I will talk about this for the next few

5   minutes.

6             [Slide.]

7             Before I do that, let me just say that the

8   dogma is that irradiation is needed to prevent GVHD

9   even when red cell units are leukocyte reduced.

10   The use of leukocyte reduction as an alternative

11   method has not been documented. Data on the log

12   reduction needed is not known.  All of our studies

13   were done on leukocyte-containing red cells.

14   Again, this is before the era of leukocyte

15   reduction.

16             [Slide.]

17             These are comments on the method that we

18   used in the assessment of the optimal dose of

19   irradiation.  We used a very sensitive limiting

20   dilution assay.  The assay was based on growth of

21   T-cells, and the assay measures up to approximately

22   5 logs of T-cell inactivation, and this work was

83

1   published in 1994 in Blood.

2             [Slide.]

3             This slide summarizes the results that we

4   found. This is a quantitative assay, but the

5   results are listed in qualitative fashion for this

6   presentation, and basically, what we found is that

7   there were still growth of T-cells at 2,000

8   centigrays, but at 2,500 centigrays, there was no

9   growth, and we did studies with 3,000 centigrays,

10   and there was also, of course, no growth at that

11   dose.

12             I should say that we also conducted

13   studies looking at two bags.  We looked at the

14   PL2209 bag, which is a citrate plasticized bag, and

15   the PL146 bag, which is a DHP plasticized bag for

16   red cells, and we found similar results.

17             We also conducted split unit studies where

18   we irradiated one-half with a linear accelerator

19   and one-half with the gamma irradiation source, the

20   free-standing cesium source, and we found no

21   difference.  We have very comparable results.

22             [Slide.]

84

1             I just wanted to mention that the current

2   nomenclature is centigrays.  For many years, the

3   nomenclature was rads - 2,500 centigrays equals

4   2,500 rads.

5             [Slide.]

6             I just briefly want to mention about how

7   we measured the dose that was delivered.  The

8   studies documenting 2,500 centigrays as the

9   appropriate dose measured the delivery at the

10   center of the simulated blood units.  We used water

11   as the simulated blood units, and we embedded

12   thermal luminescent dosimeter chips in the blood

13   bags containing water.

14             Currently, there are commercial systems

15   for dose mapping, which are based on dose mapping

16   of the canister, and I will show you an example of

17   the canister.

18             With 1-unit irradiators, there is

19   essentially no translation issue.  With

20   multiple-unit irradiators, there is a translation

21   issue.  With 2,500 centigrays delivered to the

22   canister centerpoint, some units will have a

85

1   greater dose delivered to their centerpoint.

2             [Slide.]

3             This is what we used in our studies.

4   These are these TLD dosimeter chips.

5             [Slide.]

6             This just is an example of a free-standing

7   irradiator.  This is by Nordians [ph], this gamma

8   cell 3000, and this is the canister that is used

9   for placement of the units of red cells or, for

10   that matter, platelets or plasma.

11             [Slide.]

12             This just shows one of the systems

13   available for dose mapping of the canister.

14             [Slide.]

15             Let's now turn to the evaluation studies

16   that we conducted to evaluate the influence of

17   gamma irradiation dose, 2,500, as I said, deemed

18   optimal on the red cell properties.  We used a

19   paired study approach to compare red cell

20   properties with radiation and without radiation, so

21   my control is without radiation.

22             The emphasis was on the evaluation of the

86

1   in vivo red cell viability properties.

2             [Slide.]

3             This is some study background.  Again, the

4   studies were conducted in the mid-1990s.  The

5   studies were sponsored and coordinated by the

6   American Red Cross Holland Laboratory.  The studies

7   were conducted at two sites.

8             The principal site investigators were Dr.

9   James AuBuchon at Dartmouth-Hitchcock Medical

10   Center in New Hampshire and Stein Holme, who at

11   that point was with the American Red Cross,

12   Mid-Atlantic Region, in Norfolk, Virginia.

13             This data was presented to the Food and

14   Drug Administration and helped establish the

15   guidelines in the early-mid 1990s for irradiation

16   of red cell units.

17             [Slide.]

18             In terms of methods, we used four

19   scenarios, and I will show you these scenarios in

20   detail in a few seconds.

21             Protocols 1 and 2 was the original study,

22   and these studies were done at two sites.  After we

87

1   saw the results from Protocols 1 and 2, we added

2   Protocols 3 and 4, and we did a small number of

3   studies at one site.

4             Each subject for all the protocols donated

5   two CBD  whole blood units at least 56 days apart.

6   The red cells were prepared with AS-1/ADSOL

7   preservative solution.  On one occasion, the AS red

8   cells were irradiated and stored, on the other

9   occasion, the AS red cells were stored with no

10   irradiation.  That is our control.

11             We used a PL2209 container, which was

12   being utilized at the time.  This container is not

13   utilized now, but as I said, there is no evidence

14   that the container influences the effects of

15   irradiation based on preliminary studies that we

16   conducted in the early 1990s.

17             [Slide.]

18             Again, the dose of irradiation was 2,500

19   centigrays delivered to the midsection of the blood

20   bag, as I explained before, and again the red cell

21   units were not leukocyte reduced.

22             [Slide.]

88

1             Our studies were stimulated by three

2   studies that predated our studies.  One study from

3   the NIH by Rick Davey and co-workers showed that

4   when AS-1 red cells were irradiated on day zero

5   with 3,000 centigrays, the irradiated red cells

6   gave lower results, lower 24-hour recoveries than

7   control red cells.

8             There was also a study published only in

9   abstract form from Ken Friedman at the University

10   of Mexico, and their means are listed here.  They

11   irradiated red cells on day 1, and they stored red

12   cells in one protocol for 21 days, and in the

13   second protocol for 28 days, and they found that

14   the irradiated red cells had lower recoveries.  The

15   N's were 6 in this study for both protocols, and N

16   in the NIH study.

17             [Slide.]

18             There was also a study by Paul Mintz.

19   This is an unpaired study.  The other two studies

20   were paired studies where control and irradiated

21   red cells were from the same donor, but there was a

22   study also done with AS-1 red cells from the

89

1   University of Virginia where they showed, in an

2   unpaired fashion, that irradiation after 35 days of

3   storage had a small effect compared to controls, so

4   again emphasizing that irradiation did influence

5   the 24-hour recovery, which is the kay parameter as

6   we just heard about.

7             [Slide.]

8             This slide shows the four protocols that

9   were used in the Red Cross study in the mid-1990s.

10             [Slide.]

11             I will go through each of these protocols

12   in more detail.

13             In Protocol 1, irradiation was at day 1,

14   storage for 28 days.  This was the conclusion of

15   storage.

16             [Slide.]

17             In Protocol 2, irradiation was on day 14,

18   because we thought that there is the need to show

19   the properties of the red cells when irradiation

20   takes place during the storage period.  So, in

21   Protocol 2, irradiation was 14 days, and here,

22   again we stored for 28 days through day 42, which

90

1   is the routine conclusion of storage for additive

2   solution of red cells in the United States.

3             It was based on some result that we found

4   in this study, which I will show you in a minute,

5   which showed that there were reduced recoveries,

6   24-hour recoveries, recoveries with irradiated red

7   cells of less than 75 percent, which led us to use

8   the next two protocols.

9             [Slide.]

10             This is Protocol 3.  Here, we irradiated

11   again at day 14, but we only stored for 28 days.

12   We wanted to see whether there was anything unique

13   about day 14 irradiation.

14             [Slide.]

15             This is Protocol 4.  Here, we wanted to

16   look at all the red cells, and that was just

17   discussed before.  We irradiated the red cells at

18   day 26, and we stored the red cells for two more

19   days when we did the red cells viability survival

20   studies.

21             I didn't emphasize that in the other

22   protocols, but when I say conclusion of storage,

91

1   that is when samples were taken, and the in vivo

2   viability studies were carried out.

3             [Slide.]

4             Just a few comments about the methods, and

5   I think I have covered this before basically, but

6   the autologous infusions were utilized, so the

7   individual received back a small portion of his

8   cells that were labeled with isotopes, and we

9   analyzed the red cell recoveries, 24-hour red cell

10   recoveries by the two standard methods, the single

11   label method which only involves the chromium

12   isotope, and by the double isotope method, which

13   involves chromium isotope and a 99 technetium

14   isotope.

15             [Slide.]

16             This is our data that we obtained from the

17   24-hour studies using the four protocols.  This is

18   date of irradiation and this is total storage time.

19             Let's talk about Protocol 2, because this

20   is where the results showed values less than 75

21   percent in the irradiated arm.

22             The control red cells--this is for the

92

1   single label, I will emphasize the single label

2   data because of time--the control red cells showed

3   a mean standard deviation of 76.3 plus or minus 7.0

4   percent, and the irradiated red cells had values

5   69.5 plus or minus 8.6, and there were comparable

6   values with the double label.

7             In the other protocols, the mean standard

8   deviation, the means were always above 75 percent.

9   I want to point out that the N for Protocols 1 and

10   2 were 16 paired studies, an N of 8 at both

11   studies.  At that time when we did these studies,

12   that was the dogma, to use a N of 8 at each of two

13   sites.

14             In Protocols 3 and 4, we carried out

15   studies at one site, as I said, and we had a

16   smaller number of subjects involved.

17             [Slide.]

18             This is just a graphic representation of

19   the means for the four different protocols.

20             [Slide.]

21             This is the individual data, and I will

22   show the individual data for Protocol 1 and

93

1   Protocol 2, and you can see that for Protocol 1, in

2   the control arm, no irradiation, there is one

3   value, less than 75 percent.  These values were

4   after storage for 28 days, and you can see, as was

5   mentioned, that there is a large variability in the

6   values in a set of donors.

7             This is the data with irradiation, and

8   again this is at day 28, and you can see that the

9   pattern is different than the control pattern, and

10   here there is an increased number of values less

11   than 75 percent.  This is at day 28 now with

12   irradiation on day 1.

13             [Slide.]

14             This is the data from Protocol 2, the

15   individual 24-hour recovery data, and here again

16   you see this large interdonor variation in 24-hour

17   recoveries in the control arms, no irradiation, and

18   you get the same pattern or same width with

19   irradiated cells or really a greater width, but it

20   is the same pattern.

21             You can see that there is a difference in

22   the patterns between irradiated and control cells

94

1   as indicated in the data that I showed previously.

2   We had one low value, and this is with the single

3   label method of 47 percent in Protocol 2.  This is

4   again irradiation on day 14 with storage for 28

5   days until day 42.

6             [Slide.]

7             This is the data from Protocol 1 and

8   Protocol 2 by site.  I just wanted to point out in

9   the control arm, for the studies in Protocol 2,

10   there is no irradiation, the mean from Site A was

11   above 75 percent, but as it turned out, with the

12   donors that were used at Site B, the mean was below

13   75 percent just slightly.  So, again, this

14   emphasizes the inner donor variability.

15             [Slide.]

16             This just combines all the data in a

17   summary slide, the data that I talked about that

18   was published in the three studies that predated

19   the American Red Cross study.  This is included in

20   the packet and I just wanted to show this in a

21   composite way.

22             This is when the irradiation was conducted

95

1   on day zero and day 1, and this is the data that I

2   have shown before, so I won't go over it now

3   because of time.

4             [Slide.]

5             This is a composite of the data where

6   there was mid-storage irradiation on day 14, the

7   University of Virginia study, and then the two

8   American Red Cross studies that I talked about

9   before.

10             [Slide.]

11             We also looked at long-term survival of

12   the red cells.  Long-term survival is not routinely

13   conducted when evaluating the viability properties

14   of stored red cells, but in view of the data from

15   the NIH study and from the University of Mexico

16   study, when we designed our studies, we decided to

17   also look at the long-term survival of the red

18   cells.

19             The long-term survival of red cells refers

20   to the survival of red cells that are circulating

21   24 hours after infusion, and we measured the

22   long-term survival of the red cells in terms of the

96

1   time that it took for the chromium to be reduced to

2   50 percent of initial values.

3             Without going into detail, we used a model

4   from the literature, and the circulating chromium

5   levels were determined in samples obtained from the

6   subjects.  The samples were obtained 7, 21, 28, 35

7   days after the labeled red cells were returned to

8   each subject, and we did not correct for elution of

9   chromium.

10             This has been an issue for many years

11   about elution of chromium, and there are still many

12   questions about how this elution occurs and what

13   influences the elution.

14             It is because of the questions, we decided

15   not to correct for elution.  Therefore, the

16   survival, and you will see the data in the next

17   slide, the survival time is in the range of 25 to

18   30 days, because we did not correct for elution,

19   and, of course, the normal red cell survival being

20   about 120 days total.

21             [Slide.]

22             This shows the data from the long-term

97

1   survival studies that we conducted.  We conducted

2   these studies with all four protocols, and

3   basically, there was no difference in the long-term

4   survival of the red cells for any of the protocols,

5   and also there was no difference between irradiated

6   red cells and control red cells, the values being

7   very similar.

8             This is the p value between irradiated and

9   control, and all of the p values were not

10   significant, p greater than 0.05, again, all the

11   values being around 25, 27 days in terms of the

12   mean values with limited SD values.

13             [Slide.]

14             Now, let me conclude by just briefly

15   presenting some of the in vitro results that we

16   obtained in these studies with the four protocols

17   that I have described.

18             ATP levels are used routinely to

19   characterize the quality of the red cells.  Again,

20   we did everything in a paired fashion.  These are

21   from the same units.  We took samples from the same

22   units.

98

1             With the irradiated cells, there was a

2   slight difference, a slight reduction in the ATP

3   levels in all four protocols.  There were some

4   statistical differences, but from a biological

5   point of view, it appeared to us that these

6   differences were very small.

7             This is known from the literature also

8   that irradiation of red cell units does affect ATP

9   levels.  ATP levels to some degree correlates with

10   red cell viability.  It is probably the best in

11   vitro test that correlates with viability, but is

12   not a perfect correlation test.

13             [Slide.]

14             Hemolysis.  We also measured hemolysis at

15   the conclusion of storage with all four protocols.

16   We took samples for the in vitro tests after the

17   samples were taken for the in vivo viability

18   procedures where the donors receive back their red

19   cells.

20             Here again the irradiated red cells had

21   slightly higher hemolysis, but nothing really

22   marked in terms of difference between control and

99

1   irradiated red cells.

2             [Slide.]

3             This is a composite slide of hemolysis

4   citing all the studies that I have talked about,

5   and overall, irradiated red cells in all the

6   studies have slightly higher hemolysis than the

7   controls.  Again, the controls are with no

8   radiation, but there is really no great increase in

9   hemolysis in my opinion.

10             [Slide.]

11             This slide talks about or addresses the

12   issue of potassium.  It has been known for many

13   years that potassium levels increase after

14   irradiation in stored red cells.  The potassium

15   levels increase during storage without irradiation,

16   but the irradiation enhances the leakage of the

17   potassium.

18             This slide shows that in all four

19   protocols with what we expected, the irradiated red

20   cells at the conclusion of storage had greater

21   supernatant potassium levels than the control red

22   cells.  Even in Protocol 4, where the cells were

100

1   only stored for two additional days after

2   irradiation, there was a statistically significant

3   difference in potassium level, irradiated red cells

4   versus control red cells.

5             [Slide.]

6             This is a composite of the potassium data,

7   and this is in the packet, and because the red

8   light just turned on, I won't go into any detail.

9   Again, this just shows what I said before, that

10   irradiated red cells have higher potassium levels

11   at the conclusion of storage.

12             [Slide.]

13             This is my last slide, and this is the

14   summary slide.  We concluded back in the mid-1990s,

15   as I said, these are historic studies with whole

16   blood derived red cell units, irradiation reduced

17   the retention of red cell properties during storage

18   including the 24-hour in vivo recovery.

19             The extent of change depended on storage

20   times post irradiation based on our studies.  For

21   the protocols utilized in the American Red Cross

22   study, the magnitude of the difference in the

101

1   24-hour red cell recovery between control and

2   irradiated units was limited.  The difference in

3   Protocol 2 was the greatest, and that was about 7

4   to 8 percent difference, albeit there were values

5   that were below 75 percent.

6             The last point is the long-term survival

7   parameter was comparable for control and irradiated

8   red cells under all the conditions that were

9   studied.

10             Thank you very much.

11             DR. ALLEN:  Thank you.  Very important

12   information for us.

13             A general question I have got that I

14   haven't heard answered so far, in terms of the

15   graft-versus-host potential from non-irradiated

16   cells versus irradiated, does it make a difference

17   when the cells are irradiated?  In other words, do

18   they do better if you irradiate them within 24 or

19   48 hours of collection versus waiting until day 14

20   or, as proposed, even waiting until day 26 or day

21   28, if they are infused within 48 hours after that?

22             DR. MOROFF:  I don't think there is data

102

1   to answer that question.  Susan Leitman would be a

2   better person than I on that.  Before I let Susan

3   respond to that, let me mention to you--and I

4   didn't mention this before--in terms of our initial

5   studies, which looked at the inactivation of the T

6   cells, most of our studies were done on day 1, but

7   we also did studies where we stored the red cells

8   for 7 and 21 days, and we found that the 2,500

9   centigrays inactivated the T cells from those

10   stored units also, just like they did at day zero.

11             Susan, do you want to comment on this?  I

12   think you are a better person to comment on this

13   than I am.

14             DR. LEITMAN:  There are older studies from

15   the '70s, it's a mean dose to inhibit mitotic

16   potential of circulation leukocytes is 200 rad, so

17   the reason we need to give 2,500 is simply that the

18   mechanism of irradiation is uneven and you want to

19   get every leukocyte in the bag.

20             But again if you get that dose in, you

21   inhibit the mitotic potential no matter when in the

22   life span of the cell it is given.

103

1             DR. MOROFF:  Dr. Allen, let me just be

2   clear.  To answer your question, the data does not

3   exist.  Susan shook her head yes as she agrees.

4             DR. ALLEN:  Thank you.

5             Other questions or comments?  Yes, Harvey.

6             DR. KLEIN:  Just one small technical

7   question, Gary.  These were paired studies and, as

8   I recall, you randomized the donations to

9   irradiation or non-irradiation.

10             DR. MOROFF:  Yes, they were randomized,

11   Harvey.  I did not mention that, but we randomized

12   the donations because we were bringing back the

13   donors two times with a difference of 56 days.

14             DR. ALLEN:  Yes.

15             DR. DiMICHELE:  I think your point about

16   the fact that these studies were done on

17   leukodepleted units is a very important one, and

18   you yourself said that you didn't know what the

19   impact would be.

20             In your opinion, are these studies that

21   need to be done, I mean given the fact that so many

22   of our units are now leukodepleted, do you think

104

1   this could reasonably change the standard for gamma

2   irradiation for red cell units?

3             DR. MOROFF:  I think the leukoreduction

4   systems that we now use remove 4, 5, 6 logs of

5   white cells, so I think potentially, you would not

6   need to irradiate gamma-irradiated red cells.

7             That is one point, but more direct to your

8   point, my assumption is that the white cells do not

9   influence the influence of the irradiation on the

10   red cell properties.  There is no data out there

11   suggesting that leukocyte-reduced red cells, in

12   terms of red cell properties, would behave

13   differently than the units that we used where there

14   were white cells in the units.  Susan might also

15   want to comment on that.

16             DR. ALLEN:  Dr. Leitman.

17             DR. LEITMAN:  We did that study.  At NIH,

18   we performed a paired donation control study where

19   we leukoreduced and irradiated, and the same donor

20   donated on another occasion, and irradiated without

21   leukoreduction.  We did that because synthesis told

22   us that non-nucleated cells are not significantly

105

1   damaged by irradiation, but cells are

2   non-nucleated, and the damage you see might be

3   secondary to the damage done to leukocytes in the

4   unit.

5             What we found in that paired control study

6   was if you leukoreduced prior to irradiation and

7   long-term storage to 42 days, there was no effect

8   on the 24-hour recovery of the red cell, so

9   leukoreduction prior to irradiation prevents the

10   majority of the irradiation injury we see in these

11   studies.

12             That is a very good point because many

13   centers are practicing universal leukoreduction on

14   the day of collection prior to irradiation, so this

15   whole discussion is not relevant to a huge

16   proportion of red cells collected today.

17             DR. MOROFF:  Susan, I forgot about your

18   study. That is a very important point.  So, you are

19   saying there was no effect on 24-hour recovery.

20             DR. LEITMAN:  It was not significant.  It

21   was the same, 70-something, I can't remember the

22   number.

106

1             DR. MOROFF:  Was the data from those

2   studies published?

3             DR. LEITMAN:  It's in abstract form.

4             DR. DiMICHELE:  I am sorry.  In those

5   studies, the irradiation was day 1?

6             DR. LEITMAN:  Day 1.

7             DR. DiMICHELE:  And the evaluation was

8   day?

9             DR. LEITMAN:  Forty-two.

10             DR. MOROFF:  Forty-two.

11             DR. LEITMAN:  Yes, full shelf life.

12             DR. MOROFF:  You are saying your 24-hour

13   survivals at day 42 were above 75 percent overall?

14             DR. LEITMAN:  They were very close to 75

15   percent, maybe 76, just above 75 percent in both

16   groups.

17             DR. ALLEN:  And that was the mean, so

18   there was a range around that mean, or was that--

19             DR. LEITMAN:  I can't remember the data.

20             DR. ALLEN:  Are those data likely to be

21   published in a more complete form than the

22   abstract?

107

1             DR. LEITMAN:  We can get that out.

2             DR. MOROFF:  I think that would be

3   important to publish, I agree, that would be

4   important to publish that data.

5             DR. ALLEN:  Again, I get the sense that

6   the storage solutions, that the bags may not have

7   much effect on what happens, but that the solutions

8   do, and that probably is something else that needs

9   to be commented on as we try to compare studies

10   over time when the other study parameters may have

11   changed.

12             DR. LEITMAN:  These studies were done from

13   the late '80s to the late '90s, during which the

14   practice of transfusion medicine changed, so

15   different storage vehicles were used, initially,

16   non-additive storage solutions were used, and later

17   the XL-1, XL-2, XL-3.

18             The effect of that, no one looked

19   prospectively at different anticoagulants and the

20   effect of the same dose of radiation, but if you

21   compare different studies using different

22   additives, they seem to give the same data.  It

108

1   does not seem to be a significant effective

2   additive.

3             DR. MOROFF:  That is my conclusion from

4   looking at the data that is out there.  We only

5   used one solution, but there were some studies with

6   AS-3 with neutrocil.

7             DR. LEITMAN:  They are done well, but even

8   the same study, then at two different sites gets

9   different results, so there is a lot of technique

10   related to doing these, so multi-site studies are

11   very helpful.

12             DR. MOROFF:  And there is a lot of donor

13   variability as we have been talking about.

14             DR. ALLEN:  Other comments or questions?

15   Dr. Strong.

16             DR. STRONG:  Just a couple of comments.

17   One, the issue of dosage, I think if you look back

18   into the '70s, as Susan commented, a lot of studies

19   were done in the transplantation era when we were

20   looking at, at that time, matching criteria for

21   using lymphocyte cultures that the dose of 2,500

22   was arrived at, which raises a concern about mixing

109

1   the data between 2,500 and 3,000, because one would

2   expect increasing effects with increasing doses of

3   irradiation as we see in lots of other biological

4   systems.

5             I had one question, Gary, about your data,

6   whether you have done any intergroup comparisons,

7   because of this wide variation between individuals

8   when you compare just looking at the data you have

9   presented, intergroup comparisons don't seem to be

10   very significant either.  In other words,

11   irradiation at day 1 versus day 14, if you compare

12   the 28-day storage of those two groups, it looks

13   like in some cases you have a better recovery in

14   the ones that were irradiated at day 14, so it just

15   seems to validate this problem on interdonor

16   variability, as well as interorganizational

17   differences.

18             DR. MOROFF:  I agree, Mike, it's the

19   interdonor variability which overpowers some of the

20   other comparisons, because there is such a wide

21   range in interdonor comparability or interdonor

22   values.

110

1             Let me mention about your first point.

2   Back in the '70s and '80s, a lot of the studies

3   which looked at dose were utilizing isolated

4   lymphocytes that were irradiated in the

5   transplantation model, and that is why we

6   irradiated in our dose studies red cell units,

7   because that data did not exist in the early '90s.

8             So, again what we did differently, we

9   irradiated the red cells, then, we isolated the

10   lymphocytes to do the T cell killing studies.  We

11   did not isolate the lymphocytes first.

12             DR. STRONG:  My only point here is I think

13   there is enough variability in the statistics that

14   are being presented that it is going to make it

15   difficult to make decisions about changing what we

16   have already established.

17             DR. ALLEN:  Other comments or questions?

18             DR. QUIROLO:  I just wondered if you could

19   comment on the feasibility of doing larger studies,

20   besides money, is that a possibility, or what would

21   be the impediments to that?

22             DR. MOROFF:  I think we are in an era

111

1   where we are doing larger studies.  As I said, back

2   in the early '90s, the dogma was to use 8 donors

3   per site, and now, as we heard before, the dogma is

4   to use 10 to 12 studies per site.  Yes, I think it

5   is feasible to do studies.

6             There are a lot of logistics involved for

7   getting the finances, the financial issue, but I

8   would say that we can do larger studies if needed.

9   I am not sure I understand your point.  To look

10   more at the donor variability issue?

11             DR. QUIROLO:  Yes.  There are so few

12   subjects in the studies, and the donor variability

13   is so large, how many subjects would you need to

14   negate that, so you could get good statistics?

15             DR. MOROFF:  I think we need some analysis

16   to show that.  More studies would help in the

17   statistics, I agree, and I think the studies are

18   feasible.  They would take longer, but I think they

19   are possible.

20             The question is you have to ask whether

21   larger studies are needed and on a case-by-case

22   basis.  We are going to be hearing a discussion of

112

1   statistics later which will enlighten your

2   question, which probably will address your point.

3             DR. ALLEN:  Dr. Leitman.

4             DR. LEITMAN:  This data on the factors in

5   the donor that may lower 24-hour recovery, Dr.

6   AuBuchon either presented that in abstract form or

7   published it, and we have similar data.  Low MCV,

8   low red cell size, and iron deficiency in the

9   donors is clearly associated with poorer in vivo

10   recovery in the autologous setting.

11             In addition, we have data that donors who

12   take high doses of antioxidants, vitamin E or

13   vitamin C, have better in vivo recoveries, so we

14   don't know all the factors, but there are specific

15   donor-related factors that impact.

16             DR. MOROFF:  We did not control for these

17   factors in these studies, you are right, Susan.

18   There may be some ways to understand the situation

19   in a more comprehensive way.

20             DR. GOLDSMITH:  I think in your studies on

21   Table 1, the recovery and control in irradiated red

22   cells, the studies that were done at one site are

113

1   not different statistically, and the studies done

2   at two sites are different in a statistical

3   fashion.

4             Again, this must speak to some kind of

5   innate variability site to site, as well as donor

6   to donor.  If you want to have really robust data,

7   you have to have multiple sites, I guess, to make a

8   conclusion.

9             So, I think your data support the use of

10   multiple sites very nicely.

11             DR. MOROFF:  I think multiple sites are

12   important.

13             DR. ALLEN:  Thank you.

14             DR. LACHENBRUCH:  Tony Lachenbruch,

15   Biostatistics at FDA.

16             One of the concerns that I would have

17   about larger studies is if we are looking at means,

18   we can always find a large enough sample size to

19   show that the lower confidence bound is greater

20   than 75 percent, but we still may have 30 percent

21   of the individuals are below the criterion of 75

22   percent recovery.

114

1             So, I think Dr. Kim will be addressing

2   this, but I think it is really important to say big

3   studies aren't going to solve that problem.

4             DR. ALLEN:  We will move on to our next

5   speaker at this point.  Thank you very much.  Larry

6   Dumont from Gambro BCT, Inc.

7                   Presentation - Larry Dumont

8             MR. DUMONT:  Mr. Chairman, members of the

9   committee, ladies and gentlemen, good morning.

10   Thanks for invitation to present the data.

11             [Slide.]

12             My mortgage gets paid by those people

13   right there. Mike is actually the guy that designed

14   the studies in collaboration with CBER, and he

15   couldn't be here, so I am just reporting the news.

16   I didn't make the news today.

17             [Slide.]

18             First of all, the groups that we worked

19   with, American Red Cross in Norfolk, Dr. Taylor and

20   Pam Whitley and that crew.  From

21   Dartmouth-Hitchcock, Dr. AuBuchon and his lab.

22   Blood Center of Southeastern Wisconsin, Dr. John

115

1   Adamson, Loni Kagan, and others.  At Gambro was

2   Mike and Marge.

3             [Slide.]

4             The objective of these studies was to

5   determine the in vitro and in vivo characteristics

6   of gamma irradiated, apheresis red cells compared

7   to concurrent controls prepared from whole blood.

8             [Slide.]

9             I am going to describe the methods here.

10   For past reference, essentially in Dr. He's

11   presentation, we are Company B, and in Gary's

12   presentation, we are going to be Protocol 2.

13             [Slide.]

14             In some studies, the Gambro Trima was used

15   to collect red cells.  Anticoagulant for that is

16   ACDA.  We get a packed red cell of about 250

17   milliliters at an 80 percent hematocrit.  Following

18   collection is added storage solution AS-3.  In some

19   studies that i will show you, the product was then

20   leukocyte reduced with a Pall filter, some studies

21   they are not.  I will make that clear.

22             Following that treatment, samples were

116

1   taken for testing.  The cells went into the cold

2   for 14 days.  Aliquots were taken out for testing

3   at that point.  The cells were irradiated at 25

4   grays, put back in the cold for an additional 28

5   days, and then testing was performed at the end of

6   that period.

7             Control arms, which came up earlier, was

8   whole blood collected in the standard fashion with

9   CPD as the anticoagulant where you get about 500

10   milliliters of anticoagulated whole blood.  This is

11   held for a couple hours, component processed into

12   two components, and storage solution is added, in

13   some cases filtered, some cases not, I will that

14   clear in a second.

15             What we tried to do at the different sites

16   is we wanted to use their standard procedures, so

17   we have got a couple different methods here.

18             One site actually used the Sepacell, which

19   is a whole blood filter, so that filtration, when

20   it occurred, happened right here prior to component

21   processing, and then AS-1 was added to the cells.

22             Another site utilized an AS-5, with the

117

1   Terumo bag, and they used the BPF-4 Pall filter.

2   For those products, this process actually happened

3   at 1 to 6 degrees.

4             Then, of course, after this processing,

5   testing was done and then things went into the

6   cold, they were irradiated, et cetera.

7             [Slide.]

8             The testing is the standard list of things

9   that you do with these studies - CBC, residual

10   white cells, pH, and gases, hemoglobin and the

11   plasma where we calculate hemolysis, ATP, sodium,

12   potassium, glucose, osmotic fragility, and

13   radiolabeled recoveries on the days indicated here.

14             [Slide.]

15             All the methods were standard.  The

16   important ones for this discussion, gamma

17   irradiation was at 2500 centigrays with a 1,500

18   centigray minimum using IBL 437C.

19             Radiolabeled red cell in vivo recoveries

20   were conducted at two sites.  The Red Cross in

21   Norfolk used the double label method with chromium

22   51 and technetium 99 where the technetium is used

118

1   to estimate the blood volume, as published by Dr.

2   Heaton and others.

3             Milwaukee used just the chromium label, as

4   Gary had published previously.

5             [Slide.]

6             We really have two parts to this study

7   that I am going to show you.  The first part was

8   actually evaluating non-leukocyte-reduced red

9   cells.  This was done at two centers,

10   Dartmouth-Hitchcock and Red Cross in Norfolk.

11             This is an unpaired, parallel group study

12   where there are 12 subjects per arm per center, so

13   there was a total of 48 subjects in this study.

14   These red cells were not radiolabeled.  We just

15   have in vitro data on that.

16             The second group is leukocyte-reduced red

17   cells, and that was done at two centers, Red Cross

18   in Norfolk and Milwaukee.  This was a randomized,

19   paired, crossover study with 12 subjects per

20   center.  These cells were labeled and reinfused.

21   So, you want to digest that for one second because

22   it's a little complex.

119

1             [Slide.]

2             This is the unpaired non-leukocyte-reduced

3   summary, and I am showing mean and standard

4   deviation.  At each site, there were 12 Trima

5   collections from 12 different donors and 12 whole

6   blood collections from 12 different donors.

7             The whole blood was collected in both

8   sites with Terumo system with an AS-5 additive, and

9   gamma irradiation happened after day 14 sample was

10   taken.

11             So, this shows supernatant potassium, ATP,

12   hemolysis for whole blood and Trima,

13   Dartmouth-Hitchcock, Red Cross.  So, day 14 and

14   again irradiation happened right after these

15   samples were taken.

16             To try to simplify it a little bit, the

17   blocks that are highlighted in white are those that

18   have a small p value, and the p's are actually a

19   comparison between this group and this group, in

20   this case an unpaired T test.

21             So, we can see that the potassium is

22   increasing over storage, continues to increase.  It

120

1   is slightly less in the Trima.  Same type of

2   picture you get at Red Cross.  ATPs decrease over

3   storage and after radiation, seems to be no

4   difference between the two test arms.  Hemolysis

5   also increases.  There was no difference between

6   the test arms at Dartmouth-Hitchcock, and there was

7   some difference at Red Cross.

8             [Slide.]

9             This is the second group summary.  So,

10   this is a paired leukocyte reduced red cell, again

11   mean and standard deviation, 12 pairs at each

12   center.  So, this was randomized and there was at

13   least 56 days in between donations.

14             Again, potassium, ATP, hemolysis.  In this

15   case, we had 24-hour recovery, whole blood Trima.

16   Red Cross, Norfolk, their whole blood was again the

17   Terumo system with an AS-5 additive solution.

18   BPF-4 filter was used for leukoreduction.

19   Milwaukee again used the RS-2000 filter with an

20   AS-1 additive.

21             So, we see the same picture where we have

22   got increasing potassium.  There are some

121

1   differences there. Clinically, those probably

2   aren't significant.  No difference there.  ATP, the

3   same type of picture where it decreases over

4   storage.  Hemolysis increases, and there is a

5   consistent difference between the test arms.

6             Twenty-four recoveries, mean, and standard

7   deviation again.  A significant difference was

8   found at Blood Center of Southeastern Wisconsin.

9             To get away from these giant tables, next

10   slide, please.

11             [Slide.]

12             I wanted to give you a broader picture of

13   these data in perspective with some previous

14   studies that we have done.  Hemolysis at 42 days

15   for the red cells, and I have shown here a study

16   that we conducted with Trima at Hoxworth in 1999,

17   where we have Trima red cells, and these were not

18   leukoreduced, and they are not irradiated.

19             In 2000, the Red Cross in Norfolk, and

20   here are the publication references.  These are

21   Trima red cells, and these are historical controls

22   for that center.  They were not leukocyte reduced

122

1   and they were not gamma irradiated. Shown is the

2   mean in one standard deviation unit.

3             This is the current study where we have

4   the paired study, Trima and the paired controls,

5   and where we leukocyte reduced and gamma irradiated

6   at Red Cross, Blood Center of Southeastern

7   Wisconsin.

8             Here is the unpaired parallel group study

9   where they were gamma irradiated and not

10   leukoreduced at Red Cross and Dartmouth-Hitchcock.

11   I have put the 1 percent hemolysis specification

12   for a reference.

13             You might tend to say, well, here is the

14   same center and can we compare these two irradiated

15   versus non-irradiated, and I would caution you that

16   it might tell you something, but the studies

17   weren't designed to answer that question, so be

18   very careful about jumping to conclusions even

19   though we think that is the direction things go.

20             So, what we can see is that actually, the

21   apheresis cells seem to be superior to all the

22   control cells in all these studies.

123

1             [Slide.]

2             This is the 24-hour recovery for the same

3   studies again, Hoxworth Red Cross 2000, and in the

4   current study. Again, these are means and standard

5   deviations, and I put this 75 percent in here for a

6   reference.

7             There are a couple of points that I want

8   to make with this slide.  First of all, we have a

9   hard time with what is really the specification,

10   and we can't actually find that written down.

11             We got a letter in response to an IDE

12   submission where it said that the FDA acceptance

13   criteria are the following, that for red cells, the

14   mean red cell recovery should be greater than 75

15   percent with a standard deviation less than 9

16   percent in more than 20 volunteers, and this will

17   assure that the 95 percent lower confidence limit

18   is greater than 70 percent.

19             This kind of a complex statement in my

20   view.  I think what that means is that they want to

21   have an adequate sample size, so that you can

22   demonstrate with 95 percent confidence that your

124

1   mean recovery is greater than 70.  That is the way

2   I read that.

3             So, it would be very helpful to have that

4   clarified in a guideline for us.

5             The other point I would like to make about

6   the thoughts concerning are we trying to evaluate

7   differences in the mean versus trying to approach

8   it from a tolerance perspective, allowing a certain

9   percentage that might be below some maybe arbitrary

10   number.

11             Given the fact that there is a large

12   variation between centers, which we have already

13   seen, and we have already seen the large variation

14   between subjects, and these are normal subjects,

15   these aren't patients, I am very concerned about

16   having a tolerance specification on that.  I think

17   that would be very difficult to interpret.

18             I can tell you right now we would like to

19   do all our studies over there if that would be the

20   case.  They seem to have better donors.  That was a

21   little silly statement in a way, but I am not sure

22   that that makes a lot of sense from my perspective.

125

1             I am also a little concerned with the

2   proposition that Dr. He presented earlier, defining

3   75 percent as failure, because if you look at her

4   slides, and we are Company B, that we have got so

5   many percent failures in these, and I am sure that

6   that is well defined what failure really means and

7   is 75 percent really the right number.

8             [Slide.]

9             In conclusion, from these data, the 42-day

10   apheresis red cells collected by Trima, stored in

11   ACD-A and AS-3, gamma irradiated on day 14 of

12   storage, and taken out to day 42, are at least not

13   inferior to control red cells from whole blood that

14   were stored and irradiated in a similar manner.

15             The average 24 hour in vivo recoveries for

16   Trima red cells are greater than 75 percent.

17             Thank you very much.

18             DR. ALLEN:  Thank you.

19             Questions or comments pertinent to the

20   presentation?

21             DR. KLEIN:  Larry, I just have one

22   question for you, and that is, we are going to

126

1   discuss this issue in a different context later,

2   but how do you feel about comparing processed red

3   cells from two different components without a

4   non-irradiated, in this case, control?  I mean

5   isn't it the case that if you see small changes

6   that may show that they are statistically

7   different, that they are different, that, in fact,

8   eventually, you are going to have a large

9   difference from a non-irradiated or

10   non-leukoreduced or non-processed control?

11             MR. DUMONT:  Well, that is a very good

12   point, and there is a lot of ways to design these

13   studies.  Of course, today, I am just reporting the

14   news.

15             Apparently, Mike and FDA felt that this

16   design would answer the questions that were on the

17   table.  If we wanted to ask different questions,

18   that might dictate a different study design.

19             I do want to make a comment about, you

20   know, what do we need to do irrespective of the

21   money.  I don't think we can ignore the money.

22   These are really expensive studies.  I mean the

127

1   data that I just showed you, it doesn't look like a

2   lot of data, but, you know, they probably cost us

3   $200,000, and that is just out-of-pocket expenses.

4   That is not talking about all the people that focus

5   on it.  So, that is not a small issue.

6             DR. EPSTEIN:  Harvey, you always have a

7   good question or a good comment, and really, your

8   point is going to be the subject of the next topic,

9   which is whether we can use previously licensed

10   product as the control for evaluating or comparing

11   new products or whether we need some reference

12   preparation as a way to do that.

13             I think in this case, when you speak about

14   the non-irradiated control, you are saying that

15   should be the reference population, not the

16   performance of previously approved product.

17             We agreed, and that is really how this

18   whole issue arose, is that when we looked at these

19   data, we discovered that they weren't meeting our

20   current standard for a non-irradiated product.

21             I have a question for Mr. Dumont.  Larry,

22   could you just clarify, my understanding is that

128

1   the studies you are presenting are the same ones as

2   were Manufacturer B in the FDA presentation, and

3   you focused on the mean recovery greater than 75

4   percent, but I just want to clarify, and correct me

5   if I am wrong, but these are, in fact, the same

6   studies where the test product and the control

7   product had high absolute rates of individual units

8   failing the 75 percent recovery, in other words, 5

9   out of 24 of test article and 11 out of 24 control

10   article.

11             But the point that Dr. He was making was

12   that if you looked at those distributions, they

13   were, in fact, unexpectedly high frequencies of

14   failed individual units, and that is really where

15   all the head scratching is about.

16             MR. DUMONT:  I believe that is correct,

17   that this is the Study B data.  The point that I

18   would suggest is why a 75 percent failure.

19             DR. EPSTEIN:  That is a fair point which

20   we need to discuss.

21             DR. LEITMAN:  A question on the design.

22   When we think of these studies, we come up with the

129

1   fact we would like to do four studies in a donor in

2   a year, two paired, unirradiated, two paired

3   irradiated, we have across this, and our IRB tells

4   us, as yours does, that that approaches the limit

5   of acceptance of giving a radionuclide to a normal

6   healthy volunteer.

7             So, that study may simply not be doable.

8   In addition, you would have to have the donor

9   donate a unit of blood four times in one year, and

10   at the end of that, on the fourth donation, their

11   iron status is not as good as the first time.

12             Now, you would randomize the donations to

13   make things comparable, not to introduce bias, but

14   then there is already movement in the statistics,

15   in the means, and that introduces a new variable.

16   So, you may not be able to do whatever someone just

17   suggested, unirradiated and irradiated test and

18   control in the same paired donor four times.

19             DR. LACHENBRUCH:  Excuse me, but if you

20   are splitting the samples and doing one irradiated

21   and one non-irradiated each time, isn't that

22   controlling for the iron status each time?

130

1             MR. DUMONT:  You don't split the samples.

2             DR. LEITMAN:  There is four individual

3   donations.

4             DR. ALLEN:  Thank you very much.

5             It is right at the time for our break.  We

6   were supposed to have one more the break.  I think

7   let's go ahead and take our break right now.  I

8   would like to have people back in the room and

9   settling down in 15 minutes, and we will restart

10   the session in 20 minutes instead of 30.

11             [Break.]

12             DR. ALLEN:  Our next speaker who will

13   speak on the dating of irradiated red blood cells

14   is Dr. Dean Elfath, Vice President of Global

15   Medical and Scientific Affairs at Baxter

16   Healthcare.

17                 Presentation - Dean Elfath, M.D.

18             DR. ELFATH:  Good morning.

19             I am Dean Elfath, Vice President of Global

20   Medical and Scientific Affairs at Baxter Healthcare

21   Corporation.  I am happy to be with you this

22   morning.  I represent Company A that was in Dr.

131

1   Ping He's presentation.

2             [Slide.]

3             When we debated how to go about doing this

4   study at Baxter, we actually were influenced in our

5   decision with the ongoing debate about creeping

6   inferiority by comparing one product to another

7   approved product that was approved by being

8   compared to another product.

9             So, we actually made the decision that we

10   will not include a control and our target would be

11   to evaluate the red cells that we irradiated using

12   the 75 percent in vivo recovery, as well as

13   hemolysis.

14             So, we were evaluating the in vivo and in

15   vitro characteristics of irradiated ACD-A/ADSOL red

16   blood cells, and the ACD-A/ADSOL combination is

17   actually used in two of Baxter's devices, the

18   Amicus, as well as the ALYX cell separators.

19             [Slide.]

20             We again did also our study in three

21   different sites - Yale University where Dr. Ed

22   Snyder was our principal investigator, Dr.

132

1   Greenwald was our principal investigator at

2   Hoxworth Blood Center, and Dr. John Adams was

3   actually the principal investigator at Blood Center

4   of Southeast Wisconsin.

5             [Slide.]

6             Our units were collected using the Amicus

7   because our evaluation was focused on the

8   ACD-A/ADSOL preparation, so the collection process

9   was actually done using ACD-A.  After collection,

10   ADSOL was added to these units collected, and were

11   leukoreduced by filtration.

12             Then, we actually divided the units into

13   two groups.  Group A was irradiated on day 1 and

14   evaluated on day 28, and Group B was irradiated on

15   day 14 and evaluated on day 42.

16             [Slide.]

17             The irradiation, again, we struggled with

18   the issue of irradiation because we know that some

19   blood centers and blood banks use 30 gray for

20   irradiation, not everybody used 25, so we tried to

21   actually include a group of red cells that were

22   treated that way to capture all the practice out

133

1   there in real life.

2             So, one-third of our groups that were

3   actually performed or tested at Hoxworth, who by

4   the way, used 30 gray to irradiate the red cell

5   products, were chosen to do the same, to use 40

6   gray to irradiate the units that were tested there.

7             [Slide.]

8             This is a slide I borrowed from my friend,

9   Larry Dumont, so competing companies do talk to

10   each other.  It shows the testing that we did.

11   Actually, the testing that was performed on day 28

12   included in vivo recovery for those units that were

13   irradiated on day 1, and on day 42, for those that

14   were irradiated on day 14.

15             [Slide.]

16             These are the results that we had and the

17   pertinent values.  You can see that when you use

18   two-thirds of our study, of our units that we

19   tested, when you used 25 grays for irradiation, you

20   meet the standards for hemolysis, as well as for in

21   vivo recovery, which the mean in vivo recovery

22   actually was 81.7 percent.

134

1             [Slide.]

2             For the units that were irradiated on day

3   14, and tested on day 42, and using 25 grays,

4   again, the hemolysis met the standard, and we had

5   in vivo recovery of 80.1 percent.

6             [Slide.]

7             It is different when you look at the data

8   using the 30 grays.  If you irradiate on day 1 and

9   test on day 28, we got these results that are

10   listed here, and again, even if you look at the

11   hemolysis, it was an acceptable hemolysis, low

12   hemolysis level below the acceptable limit, and the

13   in vivo recovery was 78.6.

14             So, you see this actually in vivo recovery

15   is lower than when you use the 25 grays for similar

16   treatment, irradiation on day 1 and evaluation on

17   day 28.

18             [Slide.]

19             The group that we had trouble with

20   actually was the units that were irradiated using

21   30 grays on day 14 and evaluated on day 42.  You

22   can see that the in vivo recovery here failed the

135

1   standard of 75 percent.  The mean in vivo recovery

2   was 71.5 percent.

3             [Slide.]

4             So, our conclusion was that the

5   leukoreduced apheresis of red blood cells toward an

6   ACD-A/ADSOL combination, and irradiated using 25

7   grays, met the FDA in vivo recovery requirements of

8   75 percent, and also the percent hemolysis was

9   lower than the 1 percent, which is the standard or

10   the acceptable limit, 28 days or irradiation

11   regardless of when you irradiate them.

12             So, the current FDA recommendation applies

13   and is met by this group of units when you use

14   irradiation of 25 grays of irradiation, but when

15   you use 30 grays, if you irradiate on day 1, you

16   have acceptable results on day 28, but if you

17   irradiate on day 14, day 42, data does not meet the

18   standards.

19             So, that is basically the conclusion from

20   the data that I presented to you.  This is also my

21   last slide.

22             I just wanted to make a couple of comments

136

1   actually, that in the U.S., there is no limit,

2   above limit for irradiation dose.  The AABB lists

3   25 grays and also the FDA recommendation lists 25

4   grays, and the Council of Europe actually use a

5   range of 25 to 50, and maybe that is the reason

6   behind the more stringent criteria for irradiation,

7   because we are only allowed to irradiate in the

8   first two weeks after collection, and units expired

9   on day 28 from the date of collection.

10             Also, you saw in Dr. Ping He's

11   presentation, actually the day required for

12   neonatal transfusion and intrauterine transfusion,

13   that these units will be used within two or three

14   days or irradiation.

15             This is the end of my presentation, and I

16   think we will get into some discussion after.

17             DR. ALLEN:  Thank you.

18             Comments or questions?

19             DR. KLEIN:  Dean, since you brought it up,

20   do you know the European data?  I mean is this just

21   a guess as to why they are doing it?

22             DR. ELFATH:  No, I have searched for them,

137

1   but I have not been able to actually identify the

2   actual data that was used for studying the European

3   standard.

4             DR. ALLEN:  Other comments or questions?

5             [No response.]

6             DR. ALLEN:  Thank you very much.

7             Our last formal presentation on this topic

8   is by Dr. Jessica Kim from FDA.

9                Presentation - Jessica Kim, Ph.D.

10             DR. KIM:  This presentation is about a

11   statistical point of view in evaluating quality of

12   red blood cell products.  It especially focuses on

13   in vivo red blood cell recovery.

14             [Slide.]

15             Current red blood cell in vivo recovery

16   criteria, as several times through this session

17   have mentioned, sample mean recovery is at least 75

18   percent, and sample standard deviation is no more

19   than 9 percent.

20             A one-sided 95 percent lower confidence

21   limit for the population mean is at least 70

22   percent.

138

1             The studies need to be done with at least

2   20 units and at least two sites.  In general, 20

3   units are equally divided into the two sites.

4             [Slide.]

5             The next couple of slides, we show the

6   data from the in vivo recovery studies, and this

7   table, you have seen several times before from the

8   previous speakers.

9             There were two manufacturers in this data,

10   in the in vivo recovery studies, Manufacturer A and

11   B, with the irradiation and the reinfusion

12   schedule.  Manufacturer A has two schedules,

13   Manufacturer B has the one schedule, and the

14   Manufacturer A has three sites with two different

15   irradiation dose levels, and the Manufacturer B has

16   one dose level with the test group with a new

17   product and control with old product, which is

18   already on the market.

19             [Slide.]

20             Now, this scatterplot shows the combined

21   data, and the X axis on this scatterplot shows the

22   various samples, and as you may see, the first

139

1   letter refers to the Manufacturer A and B, and

2   second code numbers represent the site, and the

3   third represent the condition or the group.

4             The Y axis of this scatterplot is red

5   blood cell recovery in percent.  Overall, this old

6   combined data shows the sample mean is greater than

7   75 percent, but I drew the 75 percent recovery

8   line, and as you may see, some of the data, some of

9   the samples shows many individuals who failed the

10   75 percent criteria.

11             That rate is our concern and we want to

12   have a further investigation for the cases.

13             [Slide.]

14             This slide shows the combined data look

15   fairly normal, so that we can use it to make

16   inferences you can use our normal curve.  That is

17   our purpose, to check the normality from this

18   combined data.

19             [Slide.]

20             For further investigation, data from test

21   units of Manufacturer B was considered, and the

22   sample mean of the particular data shows 80.78

140

1   percent and standard deviation was 9.29 percent.

2   One sided 95 percent lower confidence limit was the

3   77.53 percent, and it was conducted at two sites,

4   and the total site sample size was 24.

5             So, this data pretty satisfies all the

6   items of the current criteria except that the

7   latter allows you standard deviation 9.29, which is

8   greater than the 9 percent condition.

9             The more serious concern in this case is

10   for individual values, 5 out of 24, which is about

11   21 percent, the recovery were below 75 percent, the

12   cutting value.

13             [Slide.]

14             Now, one of the current criteria include

15   the 95 percent confidence lower limit is at least

16   70 percent, and that statement is the equivalent to

17   null hypothesis that the population mean is less or

18   equal to 70 percent, and the alternative

19   hypothesis, the population mean is greater than 70

20   percent.

21             As statisticians would think about this

22   hypothesis, what we want to show goes to the

141

1   alternative hypothesis, and everything else goes to

2   the null hypothesis.

3             [Slide.]

4             Using normality assumption, two normal

5   distributions with different means were drawn in a

6   common plane, and you see the two red normal curve

7   and with a color blue normal curve.

8             I located it at 70 percent and the 75

9   percent red blood cell recovery, and as you may

10   see, that both cases, the lower 95 percent

11   confidence limit meets the 70 percent criteria, but

12   what we have to pay attention to in here is the 75

13   percent cutting value.

14             The lefthand side, area under the curve of

15   the lefthand side of the 75 percent, that will be

16   the proportion of individuals who receive red blood

17   cell products less than 75 percent cutting value,

18   so I am going to call that proportion of failures,

19   and the righthand side of the 75 percent cutting

20   value is the proportion of successes, proportion of

21   individuals who would have received the 75 percent

22   red blood cell recovery.

142

1             [Slide.]

2             Our thoughts on the individual values are

3   as follows.  It may be clinically relevant to

4   ensure that each patient, each recipient has a high

5   viability as measured by in vivo red blood cell

6   recovery evaluation.

7             So, test samples determine percent of in

8   vivo red blood cell recovery for each subject, and

9   then defines success as in vivo red blood cell

10   recovery at least 75 percent.  That has been used

11   as a criteria a long time.

12             Then, what we want to have is the

13   proportion of successes should be large.

14             [Slide.]

15             In relation to the individual units,

16   thinking about the individual unit, the modified

17   criteria would be hypothesis about individual

18   units, and then I want to look at the alternative

19   hypothesis, that is what we want to have, that is

20   what we want to show.

21             Population proportion of successes should

22   be greater than minimum acceptable proportion of

143

1   successes. Now, I define another new word of the

2   minimum and acceptable proportion of successes.

3   Proportion of successes, as we defined it

4   previously, that is the proportion of individuals

5   who would receive high-quality red blood cell

6   product, but what low can we allow to make sure

7   each patient will have high quality of red blood

8   cell product.

9             So, I define that as a minimum acceptable

10   proportion of successes.

11             [Slide.]

12             Let me go again to the two items that we

13   need to agree upon.  The first one is the

14   definition of success.  An individual unit whose in

15   vivo red blood cell percent recovery is at least 75

16   percent.  That is pretty much I thought it is

17   agreed, but again we need further discussion if you

18   want.

19             The next item is that we need to agree on

20   a value for the minimum acceptable proportion of

21   successes.

22             [Slide.]

144

1             The pictorial explanation about this

2   hypothesis and related to individual units.  Using

3   the same normal curve, and using the cutting value

4   of 75, which categorize individual units as a

5   success or failure, and the one we want to have to

6   assure the future recipients receive the high

7   quality red blood cell products, we want to have

8   this proportion of successes, which is defined by

9   area under the curve of the distribution of the

10   individuals cut by the 75 successes should be

11   large.

12             [Slide.]

13             The following two slides, I tried to

14   calculate the 95 percent lower confidence limit for

15   the population proportion of successes where the

16   sample size is 24, and then the next slide will

17   show where the sample size is 20.

18             The first column is the number of

19   successes when sample size is 24.  None of the

20   samples failed, meaning the 24 success, when you

21   have the 24 successes, of course, in the

22   parenthesis, that number is observed proportion of

145

1   successes, meaning the 24 out of 2,400 percent

2   yield success, but when you compute the population

3   proportion of successes--this result is from, by

4   the way, the statistical program Sted Exec [?]--the

5   lower limit will be 88.27 percent, and then one

6   failure out of 24, the observed proportion of

7   successes will be 23 out of 24, which is 95.8

8   percent, and the confidence interval, 95 confidence

9   limit of the population proportion will be 81.71.

10             The same procedure was applied, and then

11   these are the numbers that relate to the number of

12   successes out of the sample size 24.

13             [Slide.]

14             This table, I applied exactly the same

15   procedure with the sample size of 20.  For example,

16   when you look at this 18 row, the third row, the 18

17   successes out of 20 would give you lower, 95

18   percent of lower confidence limit for the

19   population proportion of 71.74 percent, and the 18

20   out of 20, which is 90 percent observed proportion

21   of successes.

22             [Slide.]

146

1             Summarizing the previous two tables, let

2   me use one of the examples that I just pointed out

3   from the previous slide.  When sample size is 20,

4   if no more than two failures is acceptable, failure

5   means again the less than 75 percent recovery,

6   then, minimum acceptable proportion of successes

7   will be 71.74 percent.

8             In such a case, the sample proportion of

9   successes will be 90 percent.

10             [Slide.]

11             In conclusion, to assure the future

12   population of patients, in addition to or instead

13   of the current criteria, the mean approach, we may

14   specify the minimum proportion of the successes

15   that assures that the patients or each patient

16   receive a high quality of red blood cell products,

17   which means again higher than the 75 percent red

18   blood cell recovery.

19             Thank you.

20             DR. ALLEN:  Thank you.

21             Comments or questions with regard to the

22   statistical methodology?  Yes.

147

1             DR. GOLDSMITH:  Have you had a chance to

2   apply this methodology to previously licensed

3   product trials to determine what the outcomes were?

4   Would they meet this criteria?

5             DR. KIM:  Actually, one of the data that I

6   presented in the slides, the scatterplot, I believe

7   that one of the samples is already--let me make

8   sure, is it already--yes, it is already approved

9   one, and the results, there are many failures.  I

10   think Dr. He showed the failure rate of the control

11   and the test of the Manufacturer B.

12             DR. VOSTAL:  Yes, that is correct.  The

13   Manufacturer B had a control arm that was an

14   approved product, and it was I think a 45, 46

15   percent rate in that product when it was stored out

16   to 42 days.

17             DR. GOLDSMITH:  Is that typical for all

18   prior licensed products?

19             DR. VOSTAL:  No.

20             DR. GOLDSMITH:  That is what I am trying

21   to get a feel for.

22             DR. VOSTAL:  No, I think this is a special

148

1   case.  We are trying to point out that if you store

2   irradiated cells out to 42 days, there is a high

3   failure rate even in products that are currently

4   approved.

5             DR. ALLEN:  Dr. Davis.

6             DR. DAVIS:  How is the number 75 percent

7   recovery derived, what is the significance there?

8   What happens if somebody gets blood that is on the

9   failure side of the curve, I mean does it matter

10   clinically?

11             DR. VOSTAL:  Maybe I can try to answer

12   that.  The 75 percent is really a regulatory line

13   in the sand in terms of evaluating red cell

14   products.  I don't think it really means that much

15   in terms of transfusing individual patients. Unless

16   you have a really damaged red cell unit that is

17   going to cause a lot of hemolysis, that could cause

18   renal failure or other adverse events, having a

19   unit that would be at 60 percent, 65 percent

20   probably wouldn't have much of an effect on an

21   individual patient.

22             But when you extrapolate that criteria to

149

1   the population across the nation of all the cells

2   that are transfused, we were trying to keep the

3   quality standard of those cells relatively high, so

4   you wouldn't run into problems where you have more

5   frequent transfusions which would be associated

6   with more blood use and more exposure to possible

7   pathogens in those red cell units.

8             So, for individual patients, it probably

9   doesn't mean that much.

10             DR. ELFATH:  May I make a comment also on

11   that?  Actually, during my early training, I did

12   studies with Dr. Irma Shemanski, and we looked at

13   the nonviable cells that are present in a unit, of

14   red cells, and we found out that these nonviable

15   cells gets cleared within 24 hours by the

16   reticuloendothelial system.

17             I am not sure that hemolysis is the only

18   thing that can cause harm to the patient.  If you

19   have a sick patient that received multiple units,

20   and one-quarter of these red cells that we are

21   transfusing are engaging the immune system, that

22   this patient will fare as well as we would like

150

1   them to be.  I am not sure that any studies have

2   been actually done to look at the effect of the

3   clearance of these nonviable cells, so although the

4   units do not hemolyze when you transfuse them into

5   the patient, and they are not hemolyzed when they

6   leave the blood bank, we know that 25 percent of

7   them, on the average, are nonviable cells that gets

8   cleared within 24 hours of transfusion by the

9   reticuloendothelial system.

10             So, again, I am actually surprised at the

11   number of 75 because I would not go to the

12   supermarket and buy oranges knowing that 25 percent

13   of them are bad and may give me stomachache for 24

14   hours.

15             DR. DAVIS:  I am not sure the viable cells

16   are not bad for the immune system either.

17             DR. ELFATH:  I think they probably are

18   bad.  Actually, the immune systems gets engaged in

19   clearing these cells when they are transfused, and

20   I think it influences the patient's ability to

21   fight infections and other immune functions.

22             So, no studies have been done and I think

151

1   maybe we need some leadership in that regard

2   because I have not been able to get anybody to get

3   on this idea of improving the quality of red cells

4   during storage, because the debate is not exciting

5   and nobody is going to put money or do a lot of

6   work, so maybe the FDA would provide some

7   leadership in that regard.

8             DR. ALLEN:  We have had a presentation by

9   Dr. Kim on statistical methods and I would like to

10   keep the discussion at this point pertinent to

11   that.  We will get into the open discussion after

12   we have had the open hearing.

13             Dr. Epstein, do you want to comment on

14   that?

15             DR. EPSTEIN:  Just that I think that we

16   will get quickly mired the more we argue the

17   standard as opposed to the findings that certain

18   products currently under review don't appear to

19   meet the standard, the current standard, because I

20   don't think anyone would dispute the need for a

21   standard, and no one has pretended that it was not

22   derived arbitrarily, but does anyone really think

152

1   that there could be an approvals process or a

2   quality control process without a standard?

3             Maybe changing the standard is an

4   important issue, but it is in a certain way an

5   issue for another day.

6             DR. ALLEN:  You are absolutely right.

7   This is not the question before the committee at

8   the present time.  I think it is important for the

9   committee, however, to understand how the standard

10   was set and what changes with regard to irradiated

11   products we may want to make with regard to where

12   the standard is set.

13             But let's address the statistical paper,

14   first of all, and then we will move on with our

15   process.

16             Dr. Leitman.

17             DR. LEITMAN:  I have a question for the

18   FDA.  We are being asked to consider this in the

19   context of irradiation as something one does to red

20   cells, but you can't just apply the standard to

21   irradiation.  It has to be, I said this earlier,

22   applied across the board to everything.

153

1             We all know from the older data that an

2   additive solution, anticoagulant solutions still

3   use CPDA-1, wouldn't meet the standard at day 35 of

4   storage.  This kind of standard wasn't applied.

5   That is not used very widely, we use additive

6   solution, but it is still a licensed, approved

7   anticoagulant.

8             So, there would be big repercussions, I

9   think, of this lower limit.  The data wouldn't be

10   available for you to evaluate, because those

11   studies won't be redone.

12             The use of the word "success" bothers me.

13   You can have a regulatory guidance, you can have a

14   target, but for a clinician, if I give a unit of

15   blood, and there is a 1 gram increase in the

16   hemoglobin and no adverse effect, that is a

17   successful transfusion.  If 70 percent of the cells

18   circulated 24 hours later, it is okay with me.

19             So, what is a clinical success as opposed

20   to a statistical success?

21             DR. LACHENBRUCH:  Probably the phrase

22   "success" comes from what statisticians think of

154

1   when they are talking about the binomial

2   distribution.  If you aha, we reached the criteria,

3   that is a success, but reached target, whatever,

4   calling it a success is words, we can figure out

5   the right thing to do to make people happier.

6             DR. ALLEN:  I think the point from a

7   clinician's perspective is you don't want the

8   lawyers to get ahold of it, you transfused a failed

9   unit of cells.

10             Go ahead.

11             DR. SCHREIBER:  How do you define or how

12   do you propose that we define what the minimum is?

13   In one of the alternatives, it looked like there

14   was a minimum acceptance of 60 percent for all

15   samples, and when I eyeballed your graph, it looked

16   like that was chosen as the grandfather in all of

17   the existing products that were on your chart.

18             How would you propose a process if we look

19   at that last alternative?

20             DR. KIM:  The third alternative that Dr.

21   He proposed, it is pretty similar to the second

22   item that she proposed.  The only thing, the third

155

1   one, because we dropped the current criteria about

2   the sample mean, we want to make sure that the

3   very, very bad item, like at 40 percent, I don't

4   know whether that is feasible or not, the 40

5   percent or 50 percent may occur if we only have

6   restriction on the proportion.

7             So, that is why we want to add all the

8   sample data should be at least 60 percent if we do

9   not include the current criteria, which is about

10   the sample mean.

11             But in addition to the sample mean, if we

12   used the 95 percent confidence limit, we don't need

13   to use the restriction on all the individual unit.

14             About the 70 percent, the criteria, it's

15   data driven, it is not the statistical method that

16   we come up with this number.  When we look at the

17   previous chart, the table shows when we have no

18   more than 2 failures out of 20, then, the 95

19   percent limit will be at least 70 percent, 71.74

20   percent.  It is data driven, it is not from other

21   methods, the statistical method.

22             Our statisticians want to go from the

156

1   reverse way, but here, to weigh out the 71.74, at

2   least the 70 percent to meet the 95 percent lower

3   limit, is from to allowing no more than 2 failures

4   out of 20.

5             The third item that Dr. He proposed, that

6   all data samples should be at least 60, we don't

7   want to have bad cases if we only have a

8   restriction on the proportion.

9             DR. ALLEN:  Any other questions or

10   comments?

11             [No response.]

12             DR. ALLEN:  This will conclude our formal

13   presentations.

14             Did you want to make one quick comment

15   with regard to the statistical?

16             DR. SNYDER:  Yes, Ed Snyder from Yale.

17   When I was at Yale, Joseph Bovey [ph], who was the

18   director, gave me a copy of the Journal of Clinical

19   Investigation from 1947, which contained a

20   compilation of all of the data on blood transfusion

21   that was done for World War II.

22             In that entire issue, which was devoted to

157

1   blood transfusion, there are innumerable graphs and

2   charts and data which form the basis for much of

3   what we are talking about today.

4             In that volume, for example, is the number

5   70 percent derived from data showing the best

6   viability for various storage solutions.  It also

7   lists the 10 degree storage conditions for blood

8   transport and explains how that numbers were

9   derived.

10             So, the number wasn't just pulled out of

11   the air. At least the initial number of 70 percent,

12   and many of the other things we do, are based on

13   work that was presented in that volume, and that

14   can be made available to the committee obviously if

15   they wish to look at that further.

16             DR. ALLEN:  Thank you.

17             We will now move to the open public

18   hearing.  I have got two speakers that wish to make

19   a short presentation.  Again, I need to read the

20   open public hearing announcement for general

21   matters meetings.

22             Both the Food and Drug Administration and

158

1   the public believe in a transparent process for

2   information gathering and decisionmaking.  To

3   ensure such transparency at the open public hearing

4   session of the Advisory Committee meeting, FDA

5   believes that it is important to understand the

6   context of an individual's presentation.

7             For this reason, FDA encourages you, the

8   open public hearing speaker, at the beginning of

9   your written or oral statement to advise the

10   committee of any financial relationship that you

11   may have with any company or any group that is

12   likely to be impacted by the topic of this meeting.

13   For example, the financial information may include

14   a company's or a group's payment of your travel,

15   lodging, or other expenses in connection with your

16   attendance at the meeting.

17             Likewise, FDA encourages you at the

18   beginning of your statement to advise the committee

19   if you do not have any such financial

20   relationships.  If you choose not to address this

21   issue of financial relationships at the beginning

22   of your statement, it will not preclude you from

159

1   speaking.

2             Our first speaker is Allene Carr-Greer

3   from the American Association of Blood Banks.

4                       Open Public Hearing

5             MS. CARR-GREER:  Good morning.  My name is

6   Allene Carr-Greer, and I am an employee of the

7   American Association of Blood Banks.  While that

8   may make me conflicted, I personally don't feel

9   conflicted.  We have previously introduced

10   ourselves this morning.

11             BPAC previously discussed the issue of

12   irradiation of blood products collected in novel

13   anticoagulants and additive solutions at the March

14   2003 advisory committee meeting.

15             These novel solutions are used with

16   automated collection systems and have not been

17   specifically approved for use in red blood cells

18   that are to be irradiated or, for that matter,

19   frozen and deglycerolized.

20             At that time, FDA explained the rationale

21   used in deciding to permit the affected blood

22   components to continue to be irradiated and

160

1   distributed for transfusion.

2             During the open public hearing at that

3   time, it was brought to the attention of the

4   committee that once the product has been irradiated

5   and stored for a period of time, there may be

6   issues of concern related to the 24-hour percent

7   recovery of red blood cells.  This was a potential

8   concern for irradiated products collected both

9   manually and by automated methods.

10             AABB appreciates that FDA permitted the

11   continued use of blood products collected in these

12   novel solutions pending review of the sponsor

13   studies on RBCs collected and stored in

14   anticoagulant and preservative solutions that have

15   been approved for their devices.

16             Beginning with the 9th edition in 1978, of

17   the "Standards for Blood Banks and Transfusion

18   Services," AABB has included language regarding

19   irradiation of blood or blood components, and

20   anticoagulants and additive solutions are not

21   discussed.  Memoranda and guidance documents issued

22   by FDA have never mentioned restrictions on

161

1   irradiation with regard to anticoagulants or

2   additive solutions.

3             Given the questions posed to the committee

4   by FDA, AABB wishes to modify the statement

5   provided earlier with the following:

6             Gamma irradiation of transfusion products

7   is used to prevent graft-versus-host disease.

8   These transfusion patients are severely

9   immunocompromised, and the benefit to the

10   irradiated blood product grossly outweighs any

11   decreased percentage of recovery at 24 hours that

12   may be measurable.

13             The current requirement for percent

14   recovery is an arbitrary number.  In addition, the

15   several layers of evaluation that FDA is proposing

16   with regard to the age of the blood product and the

17   expiration date assigned to the product post

18   irradiation is a cumbersome and not insignificant

19   process that will be difficult to implement in

20   blood establishments.

21             We suggest that practical discussions are

22   necessary with the blood community before making

162

1   any change to the current processes for providing

2   irradiated blood products to transfusion patients,

3   and, as always, AABB is ready to work with FDA on

4   this important issue.

5             DR. ALLEN:  Thank you.

6             Comments or questions with regard to that

7   statement?

8             [No response.]

9             DR. ALLEN:  The next speaker is Dr.

10   Michael Fitzpatrick, America's Blood Centers.

11             DR. FITZPATRICK:  Mike Fitzpatrick.  I

12   have confirmed that I am still employed by ABC.

13             I will not read the first paragraph.

14             Our members appreciated the position, just

15   as AABB mentioned, that was taken in March of 2003,

16   when this subject was first brought up, and that

17   FDA's awareness of the clinical need for these

18   products in support of patient care required that

19   FDA allow them to continue to be produced.

20             I will start with the third paragraph here

21   more pertinent to this morning.

22             The latest study cited by FDA this morning

163

1   concludes that average 24 hour in vivo recoveries

2   for irradiated TRIMA RBC are greater than 75

3   percent, but FDA contests that conclusion in their

4   statement.  While FDA rightfully points out that

5   the donors at one site had a mean survival less

6   than 75 percent in the test units and another site

7   had mean survival in the control units of less than

8   75 percent, the conclusion that one site's results

9   cannot be included in the analysis and the other

10   site's results should be used to change the dating

11   period seems to ignore t he third site that had

12   greater than 75 percent survival in both the test

13   and the controls.

14             We don't really understand the parsing of

15   the data in this manner.  It would appear that

16   parsing out the three sites data results in a tie

17   with the tiebreaker's data tainted by poor control

18   data.

19             The FDA analysis makes no comments about

20   the fact that the mean survival in the whole blood

21   (control) irradiated units is less than the mean

22   survival of the units collected by an automated

164

1   instrument and in novel anticoagulant at all sites.

2             A review of this limited data indicates to

3   us that more analysis is required before drawing

4   conclusions.  What was the impact of the use of

5   3,000 centigray at one test site, why are the

6   irradiated red cells from the control units

7   exhibiting lower survival rates than the test

8   cells, was there a single individual with an

9   extremely low survival rate in the study conducted

10   at site 3 that skewed the results?

11             We believe FDA should have answered these

12   and other questions before bringing that matter to

13   the BPAC.  In FDA's defense, they are bringing the

14   larger question this morning, which is how should

15   these studies be evaluated, and that is really the

16   larger question as opposed to the specific one

17   here.

18             Blood products are irradiated for the

19   needs of specific patients and are usually

20   transfused within hours of being irradiated.  There

21   appears to be no assessment of the current

22   practices of hospitals and blood banks regarding

165

1   the dating and transfusion practices.

2             We sent out an urgent survey two days ago

3   in response to seeing the agenda and the

4   presentations.  I immediately had 7 responses.  All

5   7 responses indicated that the major impact would

6   be an increase in errors from having to use a

7   manual system for dating irradiated products until

8   software used to track and implement expiration

9   dates could be changed.

10             All of them were concerned about

11   implementing such a change manually and inserting a

12   process that could create more clerical errors

13   when most of the irradiated units they transfuse

14   are transfused shortly after irradiation no matter

15   what the age of the unit is at the time or

16   irradiation.

17             In addition, two centers expressed

18   concerns about hospitals that may irradiate and

19   store products for inventory.

20             Also, this morning I received some

21   comments concerning directed donations.  Directed

22   donations from family members are often irradiated,

166

1   usually very soon after collection, but may be kept

2   longer than 28 days depending on what the patient

3   is actually scheduled for.

4             Our members were also concerned that this

5   change is being promulgated in the absence of

6   reports of any potential harm being prevented.  We

7   would like to see an analysis of blood product

8   deviation reports and adverse event reporting prior

9   to make a change of this magnitude.

10             In closing, I would like to say that ABC,

11   as always, supports the use of good clinical and

12   practical laboratory data in the formulation of

13   regulation, policy, and guidance.  We feel that the

14   data presented today is very important to review,

15   but does not begin to establish the evidence needed

16   to change the current guidance and information

17   published by FDA regarding the dating period of

18   irradiated products.

19             We ask that the questions presented in

20   this statement be asked and answered - first, to

21   establish the need to change current requirements

22   and, second, to base the changes on a body of

167

1   evidence - not one study of 24 donors that does not

2   meet the number required by FDA to provide

3   significant data.

4             Thank you for the opportunity to provide

5   this statement, and if I could just make a couple

6   of extemporaneous comments, I would like to point

7   out that the EU data, the EU guidance that was

8   quoted, referred to intrauterine and fetal

9   transfusion and potassium leakage. It did not

10   appear to express a concern on red cell survival.

11             The statistical model that was presented,

12   I would like to personally say that I think that as

13   far as our experience with FDA and my previous

14   experience, that the attempt to develop a model

15   that allows you to incorporate a way of looking at

16   individual variation in donors who may have a

17   survival less than the desired 75 or 70 percent is

18   wonderful.  I think we need to move forward in that

19   direction.

20             My concern from just seeing it is that

21   there is a huge amount of variation in donors, and

22   that is the population we collect blood from, so

168

1   that is the population that is being transfused.

2             You compared single labels versus double

3   labels, leukoreduced versus non-leukoreduced,

4   apples and oranges and figs, and truly, and I am

5   speaking because I did my dissertation, a large

6   amount of it on red cell survival, and I looked at

7   a lot of that old literature that you just heard

8   about from 1947 and beyond, in my estimation of

9   that literature, the best method of assessing the

10   impact of a change in an additive solution, a

11   change in a manipulation of a red cell, whether

12   it's irradiation or leukoreduction or whatever you

13   are doing, is looking at the change in the

14   individual donor, knowing what the cell survival,

15   using the method you use whether it's single label

16   or double label, in an individual donor with

17   untreated cells, and then using the test method,

18   whatever that is, in that same donor, and using the

19   same labeling method, and then looking at the

20   change in survival in that donor.

21             Did you increase survival in the donor,

22   did you decrease survival of those cells in that

169

1   donor, allows you to assess what you are doing to

2   red cells and the impact of what you are doing to

3   them.

4             A model that sets a limit of 75 percent,

5   while your model tries to incorporate that

6   individual variation, it won't encompass all of it,

7   and you will still be stuck with a conundrum later

8   on that says I have a study where 8 of the donors

9   are below the desired 75 percent, and 15 are above,

10   and what do I do with the results.

11             Unless you know the history of that donor,

12   I don't think you will be able to assess that.

13             Thank you.

14             DR. ALLEN:  Thank you very much.

15             Comments or questions?  Yes.

16             DR. DiMICHELE:  Actually, my question is

17   for both ABC and AABB.  One of the reasons for

18   these standards that we have heard is, I mean I

19   guess there is two things we need to consider.  One

20   is patient safety, the other is resource

21   availability.

22             One of the issues for issuing a standard,

170

1   and a rather stringent standard, is to manage

2   resources, because certainly decreased red cell

3   viability per unit might increase the units needed

4   to be transfused, but the supply issue is on both

5   ends, because if you have to defer donors because

6   of this problem with individual donor red cell

7   survival, and basically, issues of guidelines that

8   are so stringent that we might not be able to

9   collect enough blood, I guess what is the impact on

10   either side.

11             I am wondering if I can get both

12   organizations to comment.

13             MS. CARR-GREER:  Currently, the standards

14   we are working with don't address that level of

15   detail, the variability with the donor.  I think

16   that is coming into play in the sponsor study

17   protocol.

18             For availability, as I believe Mike and

19   maybe someone else mentioned, practically, units

20   are transfused shortly after having been

21   irradiated, primarily, shorter dated products are

22   irradiated.  However, there are situations when

171

1   that cannot occur because a particular unit is

2   necessary.  There is the directed family units that

3   were spoken of.

4             But availability of the product is not an

5   issue that I believe blood centers are looking at.

6   It is the process of providing quickly what a

7   patient needs if it falls outside these guidelines,

8   and what I see is a cumbersome process being

9   proposed for selecting what is an appropriate unit

10   that can be irradiated and still be compliant with

11   the regulations, with the recommendations.

12             But individual donor variability is not

13   something that blood centers are dealing with.

14             DR. FITZPATRICK:  That sums most of it up.

15   I mean you wouldn't defer donors.  We don't look at

16   individual survivals, and it is a process question.

17   One is the greater question, how do you evaluate

18   these new product for licensure, and how should

19   they be evaluated reliably and comparably, so that

20   you are providing good products to patients.

21             That is the broader question that FDA is

22   asking, which we all need answers to.  Just as you

172

1   heard the manufacturers, sometimes it is hard until

2   you get into pre-IND discussions with FDA to know

3   exactly what parameters you should be evaluating to

4   make sure you provide the right data for licensure

5   of the product.

6             The other is not all hospitals have

7   irradiators, so they have to go to blood centers

8   and ask for irradiated products and have inventory

9   on hand for patients.  A 48-hour window might make

10   the difficult.  We don't have data on that, but it

11   might, the directed donation question, and then

12   also not all collection centers have irradiators

13   and sometimes have to use other sites to irradiate

14   blood to provide it to hospitals.

15             So, there is a practical aspect to it that

16   would need to be worked out.

17             DR. STRONG:  Just a practical comment.  I

18   posed this same sort of question to my staff when

19   we saw the questions to be addressed here, and it

20   caused near panic in our transfusion service

21   laboratory.

22             First of all, software systems that don't

173

1   allow for different kinds of dating applications

2   depending on how long the unit was started before

3   you irradiated, how we would manage units that were

4   irradiated 28 days after storage.

5             We would probably just have to eliminate

6   those, so that would have an impact on inventory

7   for sure, and because we would not have a means of

8   managing the dating process with the current

9   software systems, it might force us to essentially

10   make the hospitals buy the unit, and not allow them

11   to return it.  That could, in fact, result in

12   higher outdating.

13             But I think the main issue here is that

14   there is no data.  I mean we have not really

15   evaluated what the impact would be.  So, to make

16   this kind of a change, I think we ought to have

17   some data on what those impacts would be.

18             DR. ALLEN:  Can I get a clarification of

19   the statement you just made?  You were talking

20   about you might force the hospital to purchase the

21   unit and not be able to return it.  Is that if they

22   choose to irradiate it on site in the hospital?

174

1             DR. STRONG:  In our setting, I mean we are

2   a little bit different in Seattle, because we are

3   the transfusion service for the hospitals in

4   Seattle, so when they order a unit, they either

5   transfuse it or return it.

6             DR. ALLEN:  Dr. Klein.

7             DR. KLEIN:  In terms of supply, I think

8   the impact of a more stringent standard or a

9   different statistical evaluation of the standard

10   would be primarily to shorten the period of time at

11   which you can store different components, and that

12   might have some impact on supply, but it certainly

13   doesn't mean that you reject individual donors,

14   because you don't know anything about the

15   individual donor's red cell survival.

16             DR. ALLEN:  Do you want to make a

17   statement in the open--

18             DR. DAVEY:  Yes, I do, if I could, Mr.

19   Chairman.

20             I am Richard Davey.  I am the chief

21   medical officer at the New York Blood Center.  The

22   New York Blood Center is the largest blood center

175

1   in the United States.  We are members of ABC.

2             I would just like to make a couple brief

3   comments.  We certainly support the comments of the

4   AABB and the ABC, and support their conclusions.

5   Just a couple of comment for the committee.

6             I think it is important for the committee

7   to remember that the data we have seen, both

8   historical and the data today, are representing the

9   most extreme situations under which red cells are

10   irradiated and stored.

11             The data that Susan and I did a number of

12   years ago, irradiating with 3,000 centigray on day

13   zero until day 42 resulted in I think 68 percent

14   survival.  The data from Dr. Elfath today, just one

15   bit of data, irradiating on day 14 with 3,000

16   centigray at day 42 with 71.5 percent survival.

17             So, these are the most extreme situations.

18   It is very unusual, I think Mike would agree, that

19   in normal blood banking situations, that patients

20   would receive units that have been irradiated and

21   stored under those conditions, and it is highly

22   unlikely that they would ever receive two units

176

1   that would be stored and irradiated under such

2   extreme conditions.

3             So, I think the clinical sequelae of these

4   stringent conditions are very minimal at best.  As

5   a matter of fact, we don't know what the clinical

6   sequelae are, there is no clinical data that

7   transfusing cells stored and irradiated in this

8   manner have any adverse clinical effects.

9             However, and again to echo what Mike has

10   said, I polled my staff yesterday at the New York

11   Blood Center about the implications, and they were

12   likewise very concerned about what this would mean

13   with our computer system, with managing our

14   inventory, with dealing with our member hospitals,

15   with managing inventory at hospitals whether they

16   have irradiators or not, and we feel that there are

17   problems or potential problems with increasing

18   error.

19             So, I think, in summary, I am not sure we

20   have a problem here that we really need to solve,

21   but I think we may be creating problems, and I

22   think, too, to maybe use a phrase that is a little

177

1   bit too worn, but might be applicable here, I think

2   we have to be very careful of making perfect here

3   really the enemy of the good.

4             Thank you.

5             DR. ALLEN:  Any other comments or

6   questions in this open public session?

7             [No response.]

8             DR. ALLEN:  The open public hearing is now

9   closed. We will move on to the open committee

10   discussion with the presentation of FDA's current

11   thinking and questions for the committee.  I will

12   just point out that the questions, actually, there

13   is two sets of questions that need to be considered

14   independently.

15             Dr. Vostal from the FDA will present FDA's

16   current thinking.

17                FDA Current Thinking and Questions

18           for the Committee - Jaro Vostal, M.D., Ph.D.

19             DR. VOSTAL:  Thank you.

20             [Slide.]

21             The first question to the committee is:

22   Do the committee members agree that the current

178

1   recommendations regarding the dating period of

2   gamma-irradiated red blood cells should be

3   modified?

4             Just to remind you what the current

5   recommendations are, as Dr. He presented earlier,

6   the dating period for red cells should not be more

7   than 28 days from the date of irradiation, but

8   should not exceed the dating period of the original

9   products.

10             So, basically, that gives you, you are

11   allowed to irradiate the cells for the duration of

12   42 days.  We saw some of the data that was

13   presented today, red cells seem to have difficulty

14   maintaining the storage lesion, as well as the

15   irradiation lesions when they are stored out to

16   long periods of time.

17             Would you like me to run through all of

18   them?  Certainly.

19             [Slide.]

20             If you agree to No. 1, please comment

21   whether the available scientific data support the

22   following candidate modifications to FDA's current

179

1   guidance on irradiated red blood cells.

2             a.  For red blood cell products that are

3   irradiated within the first 26 days after the date

4   of collection, the products should not be stored

5   more than 28 days from the date of collection.

6             b.  For red blood cell products that are

7   irradiated on or after 26 days from the date of

8   collection, the post-irradiated products should be

9   stored no longer than 48 hours after irradiation.

10             [Slide.]

11             Does the committee have any alternative

12   modifications to FDA's current guidance regarding

13   the dating period for gamma-irradiated red blood

14   cells that should be considered?

15             [Slide.]

16             These are questions on red cell in vivo

17   recovery acceptance criteria.  Please comment on

18   the following options:

19             a.  Keep the current criteria, which is

20   the sample mean is greater or equal to 75 percent,

21   the standard deviation is less than 9 percent, and

22   a 95 percent lower confidence limit for the

180

1   population mean above 70 percent.

2             b.  Proposed new criteria, criteria 1:

3   Sample mean will be 75 percent, standard deviation

4   will be less than 9 percent, and a 95 percent

5   one-sided lower confidence limit for the population

6   proportion of successes to be greater than 70

7   percent.  Here, success is defined as meaning

8   greater or equal to 75 percent of red blood cell

9   recovery in individual donors.

10             c.  Proposed new criteria 2:  A 95 percent

11   one-sided lower confidence limit for the population

12   proportion of successes to be greater than 70

13   percent.  Again, success is defined as greater than

14   75 percent, or equal or greater than 75 percent red

15   cell recovery in individual donors, and a minimum

16   individual recovery of all samples in all donors to

17   be greater than 60 percent.

18             DR. ALLEN:  May I ask for a clarification

19   on that last proposed new criteria 2?  That, I

20   assume is a statistically derived minimum

21   individual recovery of greater than 60 percent.

22             DR. VOSTAL:  Sixty percent.  That actually

181

1   is open to discussion.  From the current dataset

2   that we have, that was the appropriate value,

3   however, we could decide that we may want to set

4   that higher or lower based on clinical opinion.

5             DR. KUEHNERT:  I just have a clarification

6   on that last proposed criteria.  That says that no

7   matter how big a study you have, if you have one

8   sample that is below 60 percent, then, it flunks,

9   is that right?

10             DR. VOSTAL:  That's right.  I mean the

11   study doesn't flunk, but there is a failure that

12   will be counted as the proportion against the

13   total, so when you derive your proportion, that

14   will be one of the failures.

15             DR. ALLEN:  Except as I interpret that,

16   when we say that minimum individual recovery for

17   all units is greater or equal--if you had one that

18   came in at 57 or 58, essentially, the method could

19   not ever be adopted because it failed under that

20   criteria.

21             DR. KUEHNERT:  Is that true?

22             DR. VOSTAL:  Actually, I misspoke, you are

182

1   correct.  What we want to do is fix the proportion

2   of successes, and a success is greater than 75

3   percent.  We want to make sure that the failures

4   that we see are not down in the 20 percent range,

5   so we want to limit the failures to--

6             DR. KUEHNERT:  All it takes is one.

7             DR. VOSTAL:  Right.

8             DR. LEITMAN:  Could I ask a question?  I

9   am not a statistician, but Dr. Schreiber, are you a

10   statistician, because in most studies, a member of

11   the panel who has more experience in this, if there

12   is one piece of data, one outcome, one result that

13   is extraordinarily different from the group of the

14   others, it is okay statistically to call that an

15   outlier of unknown significance and eliminate that,

16   state that it is eliminated.

17             So, in most studies, one doesn't place

18   much weight at all on a single outlier, yet, in

19   this proposal, a single outlier can derail a

20   proposed new solution, new method of storing blood

21   components, but this method proposed is distinctly

22   different from what we apply to studies in general.

183

1             DR. SCHREIBER:  I think your comment is

2   well taken.  I think when you have studies and you

3   have outliers, we pull our hair out trying to deal

4   with how do you deal with outliers.  You know, you

5   Windsorize [?] them, do you take a normal, do you

6   set a limit, but I think when you look at a process

7   like this, what you are doing is you are just

8   trying to decide how much lower you are willing to

9   go with a product, that it is safe or usable.

10             I think it is okay to take your standard,

11   because that is what your population is, and you

12   know that wherever you take your cut, that there

13   are going to be some people that walk through the

14   door, that they contribute, and you are going to

15   have products that are much lower than that.

16             So, I don't have any problem with setting

17   a standard or a cut, I just don't know where that

18   cut should be, should it be 60 or 50 or 40.  The

19   question you originally raised I think is the key

20   one.  If you knew that a product was only X percent

21   viable, would you transfuse it. If it was the only

22   thing available, I am sure you would.

184

1             DR. ALLEN:  We sort of started our

2   discussion on the second set of questions.  I think

3   what we will do, can we go back and see the first

4   committee question, and let's take that in

5   sequence.

6             Dr. Klein, do you want to open the

7   discussion?

8             Committee Discussion and Recommendations

9             DR. KLEIN:  First, I want to make sure

10   that I understand the current status, which I think

11   I do, and that is that the dating period for

12   irradiated red cells, is your guidance, should not

13   be more than 28 days from the date of irradiation,

14   and should not exceed the dating period of the

15   original product.

16             So, it is not allowing something to be

17   irradiated on day 1 and kept for 42 days.

18             DR. VOSTAL:  That is right.

19             DR. KLEIN:  I may have misunderstood what

20   you said, because I thought you said the latter,

21   and not the former.  So, what we are dealing with

22   is 28 days from the time of irradiation, not to

185

1   exceed the total time allotted for collection.

2             DR. VOSTAL:  That is true.  What I was

3   pointing out, that you can actually irradiate at

4   day 14, and store it out to 42 days, or you can

5   irradiate at day 41 and transfuse.

6             DR. ALLEN:  Other questions or discussion

7   on this? Yes.

8             DR. LEITMAN:  Let me just throw the

9   question on its head.  There are members of the

10   committee that studied this, that have thought that

11   the current guidance is too restricted, that the

12   most extreme circumstance that Dr. Davey quoted in

13   this study, and Moroff's study, irradiation on day

14   1 and storage to the full extent of storage life of

15   the unit, 42 days, was chosen for those studies to

16   choose the worst case scenario, because everything

17   shorter than that is better.

18             One gets a 68, 69 percent recovery, in

19   vivo recovery at 24 hours, and that is not that bad

20   for an extreme situation, when perhaps that unit is

21   the only one available, or it's a directed donor

22   unit, or it went to the OR intended to be given,

186

1   but wasn't given, and is needed further on, all

2   sorts of scenarios one can think of.

3             Again, that is an unusual circumstance, a

4   very small percent of all such units that would be

5   transfused, but it is interdicted now.  So, there

6   is a contrary opinion to what is already in effect.

7             So, my response to should the current

8   recommendations be changed, I have no data

9   presented today that would make me think I am

10   adversely affecting the outcome of patients using

11   the current standard.

12             DR. GOLDSMITH:  If I understood the

13   presentations and our argument is that the mean is

14   one aspect of the data analysis, and hidden in the

15   mean sometimes, because of variations, there are

16   some failures which are quite serious failures.

17             In Dr. Ping's presentation, what she

18   showed that in the one study, if I remember the

19   numbers correctly, there was close to 46

20   percent--maybe not as high as that--was 26 percent

21   that had actual failures, in other words, it was

22   below 75 percent.  Some of those failures, I don't

187

1   have the numbers, were much less than 75 percent.

2             So, hidden in your study, and maybe you

3   know the data, you say the mean was 68 percent, but

4   I don't know what the N was, and I don't know, of

5   that N, how many failures there were, in other

6   words, less than 75 percent, and how serious were

7   those failures, were some of them 50 percent and 40

8   percent, and if that is the case, doesn't that

9   raise a problem.

10             I mean the problem that we are raising is

11   not that the mean is a problem.  When we look at

12   the data, we see that individuals fail, and those

13   individuals can fail in a serious manner.

14             DR. LEITMAN:  If you want a specific

15   question, I think there were two, 57 and 58

16   percents of an N of 8.  Dr. Davey is the first

17   author on that study.  When we looked at those,

18   those were the two individuals with the lowest MCV,

19   and that was an older study using perhaps standards

20   that aren't as good as they are right now.  Those

21   are the lowest two numbers.

22             Is that what you remember?

188

1             DR. ALLEN:  Dr. Klein.

2             DR. KLEIN:  I guess my question is what is

3   the compelling need to change what we currently

4   have, and I guess what I am hearing is because in

5   unpublished studies that you have shown us this

6   morning, and the details of which we have seen from

7   the presenters, you have a number of what would be

8   considered failures statistically.

9             I would think that, based on this

10   relatively small number, I would be reluctant

11   personally to change the current guidance.  When we

12   get to Issue No. 2, we may want to redefine what

13   our standard is, and I don't know whether we will

14   or not, but we may.

15             My suspicion is we won't want to redefine

16   it in a say that will be so restrictive that we

17   would have to go back and relicense all of the

18   various products we currently have that are being

19   irradiated.

20             So, I am not sure that I see the pressing

21   need to change the current guidance.  As best I can

22   tell, there is no clinical problem, we don't have

189

1   any data, and what we are seeing is maybe a

2   heads-up that the issue needs to be addressed, not

3   just for irradiation, but for everything we do to

4   red cells and for the length that we store them,

5   but I think it might be precipitous to make a

6   change that would make life more difficult without

7   a pressing need to do so.

8             DR. ALLEN:  Dr. Kuehnert.

9             DR. KUEHNERT:  I would agree with that

10   sentiment. It seem like--and this sort of goes back

11   to actually considering the last question

12   first--you want to have acceptance criteria that

13   everyone is comfortable with before trying to

14   change things on specific issues, and taking a

15   cohesive approach after there is agreed-upon

16   recovery acceptance criteria would be a better

17   approach than trying to change things for

18   irradiation, and then changing the acceptance

19   criteria and then going back and changing things

20   all over again.

21             DR. ALLEN:  Yes.  I don't disagree with

22   you, nonetheless, I think we will continue in the

190

1   order that the FDA has asked to consider.

2             DR. KUEHNERT:  Yes, that's fine, just

3   consider that.

4             DR. ALLEN:  Dr. Leitman, could I ask you

5   to restate what you said with regard to irradiation

6   of leukoreduced units?

7             DR. LEITMAN:  I am not sure what I said,

8   but this is the actual data.  We did two sets of

9   paired studies on a different group of N equal 8

10   each time.  In the first set, quoted correctly by

11   Dr. Moroff, the control arm was non-leukoreduced,

12   non-irradiated, with a mean recovery at 24 hours,

13   autologous single-labeled chromium 51 recovery of

14   78 percent, and then the test arm was

15   non-leukoreduced, irradiated, 68 percent.

16             Then, we presented that data at a BPAC

17   meeting and a physicist, a radiation oncologist I

18   think, got up and commented wasn't that interesting

19   that red cells are non-nucleated, and isn't the

20   damage more likely to be secondary, indirect damage

21   to red cells, do damage to granulocytes, which

22   degranulate, perhaps lysosomal enzymes might be

191

1   doing the red cell damage.

2             That was the impetus for the second study,

3   which I think I didn't quote correctly.  In the

4   second study, both arms were irradiated.  The

5   control arm was irradiated, non-leukoreduced, and

6   thus was the same as the test arm from the previous

7   study.

8             The test arm was irradiated, but

9   leukoreduced prior to irradiation.  So, the

10   question asked in that study was what is the effect

11   or pre-storage, pre-irradiation, leukoreduction, if

12   you are going to irradiate, and that data was

13   control arm 72 percent, leukoreduced, irradiated 78

14   percent.

15             So, in the same institution, with the same

16   technologists, with the same methodology, the mean

17   for irradiated, non-leukoreduced, went up 4

18   percent, 68 percent, 72 percent, only within two

19   years perhaps of performing those studies.

20             The irradiated, leukoreduced survival,

21   then, of in vivo recovery of 78 percent was the

22   same as non-irradiated, non-leukoreduced, but not

192

1   in a cohort study.

2             DR. DiMICHELE:  And what day was the

3   irradiation, and what day was the survival done?

4             DR. LEITMAN:  We always performed a worst

5   case scenario, so leukoreduction was within several

6   hours of collecting the unit, which is the industry

7   standard, and irradiation was within 24 hours of

8   collection, so day 1, and storage was 42 days in

9   additive storage solution.

10             DR. DiMICHELE:  Do you have the ranges on

11   those numbers?  No?

12             DR. LEITMAN:  No.

13             DR. DiMICHELE:  That's okay.  I am glad

14   that you asked that question because, you know,

15   that is what I was going to say.  In terms of

16   changing the guidelines, I am more concerned about

17   this issue of--that is one of the issues I was

18   getting at in terms of effect on supply, because if

19   you can only store all these units that we

20   irradiating more and more, and we can only store

21   those units for 28 days versus 42 days, I was just

22   thinking there must be some impact on supply.

193

1             From what we have seen here today, with

2   the data that has been supplied by some of industry

3   with respect to using the old guidelines for their

4   new apheresis products, but actually leukoreducing,

5   you know, using leukoreduction is part of the

6   process, the data at 42 days with the current

7   guidelines--I am not saying irradiation at day zero

8   and then evaluation at 42 days--but certainly

9   irradiation at day 14 and evaluation at 42 did not

10   look any worse than the 28-day data.

11             I am certainly not a blood banker or a

12   transfusion medicine person, but it would seem to

13   me that if we are going more and more toward

14   leukoreduction, that the same standards have to be

15   looked at with respect to leukoreduction, and that

16   we may not need to decrease the storage time of

17   irradiated cells.

18             So, I am not sure--we have good data, but

19   given the standard that has changed, I am not sure

20   we have the data we need for right now in order to

21   change the standards.

22             DR. ALLEN:  Dr. Epstein.

194

1             DR. EPSTEIN:  It is probably apparent to

2   the committee members, but we have put two

3   different issues on the table.  One is whether we

4   should be changing the approval standard for new

5   products and processes, and the other is whether we

6   should be changing the guidance on products that

7   are appropriate for use, namely, how long stored

8   after irradiation.

9             It is possible to dissociate those two

10   questions and answers.  It was suggested a little

11   bit earlier that maybe we ought to take the second

12   set of questions first, and I am starting to think

13   that that is the wiser course.  I think it will be

14   easier for us to resolve the question of whether we

15   need to change the guidance once we have talked

16   about the question of the approval standard.

17             The core issue with the approval standard

18   is whether we should be looking at the range of the

19   distribution or only the mean, and it is a simple

20   enough question.  What has been said many, many

21   times is that there is interdonor variability, that

22   if you look at a group of tests, a group of donors,

195

1   a group of sites, you are going to see a range.

2             What the FDA has said to you is that you

3   can have a very good mean, which obscures the fact

4   that some processes are allowing through a large

5   number of outliers, not 1, not 2, but 40 percent,

6   and should we care about that or are those bad

7   processes, are those processes that are introducing

8   too much variation to be acceptable from an

9   approvals process point of view.

10             That question can be dissociated from

11   whether the data are worrisome enough to change the

12   guidance.  What is at stake here is that we know

13   that irradiation is reducing survival and it also

14   seems to be broadening that range, and so you are

15   going to have some of these lower recoveries.

16             I am hearing it said that, well, it is a

17   rare event because we don't generally irradiate

18   early and store for a long time, and that in any

19   case, we don't have an apparent clinical problem.

20             If that is the prevailing point of view,

21   the argument would be don't change the guidance,

22   but that doesn't mean don't refine the approval

196

1   standard.  So, I think it might be helpful to

2   follow the suggestion--I forget, it might have been

3   Harvey who said it--and look first at the approval

4   standard, deal with the issue of product approvals,

5   and then come to the guidance issue.

6             DR. ALLEN:  Thank you for the permission

7   to switch the order.  I would like to just ask one

8   question before we do that, however.

9             Can somebody tell me about what percentage

10   of red cell units transfused have been irradiated

11   at the present time and what the trend is on that?

12             Peter.

13             MR. PAGE:  Peter Page, American Red Cross.

14             We reviewed all the red cell units that we

15   irradiated in the 12-month period ending June 30,

16   2004.  We irradiated 121,825.  That is out of just

17   over 6 million units of whole blood collected, so

18   it is about 2 percent were irradiated.

19             I have slides, I am not sure it is worth

20   the time to show, but 80 percent of those were

21   irradiated by day 9, and by 28 days, 97 percent

22   were irradiated.  So, the bulk of irradiation

197

1   occurs really in the first week.

2             Now, that is when it is irradiated, we

3   don't know when it was transfused.

4             MR. WAGNER:  Steve Wagner, American Red

5   Cross.

6             I recently looked over in 2002, a

7   marketing study done by a commercial company that

8   looks at the blood industry, and in it was

9   categorized the percentages of different products

10   that were provided.

11             I believe that it's about 10 percent of

12   the red cells of the entire country are gamma

13   irradiated.  The blood centers irradiate a minority

14   of them.  Most of them are irradiated at the

15   hospitals.

16             DR. ALLEN:  And presumably, then, would be

17   transfused fairly shortly after--well, I guess we

18   can't say that because the presumption would be

19   that they would be transfused, but then for

20   whatever reason, they might not be and they would

21   go back into storage if they are still within date.

22             MR. WAGNER:  There was no data on that.

198

1             DR. ALLEN:  Thank you.

2             Yes, Susan.

3             DR. LEITMAN:  Dr. Klein and I are probably

4   going to say the same thing.  It is not captured

5   now, but centers that have a vast proportion of

6   patients requiring irradiated blood products,

7   cancer treatment centers, such as M.D. Anderson,

8   Sloan-Kettering, Stanford might do this, the NIH,

9   especially, if such a center has seen a case of

10   lethal graft-versus-host disease, practice

11   universal irradiation, and that is clearly an

12   increasing trend, but the data hasn't been

13   captured.

14             I would put it--about more than 2 dozen

15   centers you think?

16             DR. KLEIN:  I would add that it is a

17   moving target.  All of the blood in Japan is

18   irradiated, all of the blood, and increasingly, you

19   are going to see this in the United States, because

20   if you think TRALI is an issue, so is

21   graft-versus-host disease.

22             DR. ALLEN:  Do you know, do they also

199

1   leukoreduce, do they filter first and then

2   irradiate?  I mean are these going to become

3   mutually exclusive that we will move to one process

4   or the other?

5             DR. KLEIN:  Japan doesn't leukoreduce, but

6   I would like to correct an earlier statement.

7   Current leukoreduction will not prevent

8   graft-versus-host disease, and should not be used

9   for that purpose.

10             DR. ALLEN:  I assumed that the term is

11   chosen carefully is reduction, it is not

12   elimination.

13             DR. LEITMAN:  They are done for two very

14   different procedures, are applied for several very

15   different purposes, so leukoreduction is applied to

16   eliminate the most common transfusion reaction,

17   which is afebrile nonhemolytic transfusion

18   reaction.  Irradiation won't touch that.

19             It is also done to avoid HLA

20   alloimmunization, and gamma irradiation won't touch

21   that.

22             So, many centers, again, the centers we

200

1   just spoke about, that deal with a high proportion

2   of long-term transfusion dependent patients with

3   hematologic disorders and post-transplant, do both.

4             DR. ALLEN:  Thank you.

5             Dr. Strong.

6             DR. STRONG:  Just to make it more

7   confusing, I think there is a broad range.  For

8   example, in Seattle, which is a mixture of cancer

9   centers, transplant units, and general hospitals,

10   our irradiation practice is about 30 percent of our

11   blood is irradiated, so it is somewhere in that

12   range.

13             DR. ALLEN:  Okay, very helpful to have all

14   of this information out.

15             Shall we now move to the second set of

16   questions?  We will entertain discussion on this

17   set of questions and vote on this first before

18   going back to the other set.

19             Questions and comments?  Yes, Dr.

20   Goldsmith.

21             DR. GOLDSMITH:  Just to confuse this a

22   little, I would like to suggest a fourth

201

1   alternative here to be considered by the committee,

2   which would be Item A, which is sample mean greater

3   than plus the standard deviation, which does

4   include variability obviously, and the 95 percent

5   lower confidence for the population mean above 70

6   percent, and put a rider on it, plus a minimum

7   individual recovery of all samples, and I just

8   picked greater than or equal to 50 percent, which

9   would make this a stricter kind of criterion than

10   the ones that exist today, but might not deal with

11   these other population proportion of success

12   issues, that I don't have a good feel about for

13   prior products that have been licensed.

14             I just don't have a good feeling, would

15   they fall into this or would they not fall into

16   this, and this would get rid of studies that had

17   very low outliers.

18             DR. ALLEN:  Could you read that again,

19   please, and if you have that written out, if you

20   could bring it to Dr. Smallwood, that would be

21   helpful.

22             Are you proposing this as a motion to add

202

1   an option D?

2             DR. GOLDSMITH:  Yes.

3             DR. ALLEN:  The answer was yes.  Dr.

4   Goldsmith, this is a modification of (c)?  It is a

5   modification of (a), okay.

6             DR. VOSTAL:  The way I understand it, it

7   is adding a lower limit to the failures over here.

8             DR. ALLEN:  So, it is keeping the same

9   basic statement, but adding a lower limit failure.

10             DR. VOSTAL:  Right, and the proposal has

11   been 50 percent would be the cutoff for failures,

12   and we can discuss that.  We thought 60 percent

13   would be appropriate, but I think that is open for

14   discussion.

15             DR. ALLEN:  Dr. Leitman.

16             DR. LEITMAN:  I would like Dr. Moroff, who

17   has reviewed this data very thoroughly, to help me

18   if I am wrong here, but on b, every study of an N

19   of 8, and of 6, and of 8, and of 10, and of 20--an

20   N of 16, there haven't been N of 20 studies--in the

21   control arm, there has always been at least one

22   subject whose unit was not irradiated, one arm

203

1   there was no irradiation, where the recovery was

2   less than 75 percent.  So, 1 out of 8 is always

3   greater than 1 out of 10, which is the 20 arm.  It

4   is not doable, nothing would have been acceptable.

5             I am sure that in our control arm, we had

6   1 out of 8, that was a 68 percent recovery at 24

7   hours, at 42 days.

8             Gary, do you want to comment?

9             DR. MOROFF:  Susan, I agree, that in all

10   the studies that I have reviewed, there are always

11   at least one, in control studies, at least one

12   recovery that has less than 75 percent.  It varies,

13   sometimes it is 72 percent, sometimes it is 70,

14   sometimes it is even less, 65 percent, it varies.

15   In some cases, depending on the N, it is two or

16   three studies that are less than 75.

17             DR. LEITMAN:  In the non-irradiated

18   control.

19             DR. MOROFF:  Yes, I am talking in the

20   controls now, non-irradiated.

21             DR. ALLEN:  It is a good thing that humans

22   aren't subject to quality assurance.

204

1             So, this is the proposed (d) option for

2   consideration.  The sample mean of greater than or

3   equal to 75 percent, standard deviation greater

4   than or equal to 9 percent, and a 95 percent lower

5   confidence limit for the population mean above 70

6   percent.  That is the same as (a) and then adding

7   "and a minimum individual recovery of all samples

8   greater than or equal to 50 percent."

9             DR. LACHENBRUCH:  Isn't that giving that

10   one outlier, which Dr. Leitman said, a tremendous

11   amount of weight?

12             DR. GOLDSMITH:  I was allowing an outlier

13   to remain in as long as it was over 50 percent.  If

14   it was an extreme outlier, then, it would

15   disqualify the study the way that this was written.

16             DR. LACHENBRUCH:  So, what that means is

17   the study has to be redone completely, is that

18   right?

19             DR. GOLDSMITH:  That's correct.

20             DR. LEITMAN:  We are going to pick our

21   subjects very carefully.

22             I want to comment, in our studies, we

205

1   didn't have an outlier that low, but the test group

2   with the lowest percent recovery were the same

3   people as the lowest percent recovery in the

4   control.  So, the person who was 67, 68 percent in

5   the non-irradiated arm of the study was the 57

6   percent in the irradiated arm.  So, the percent

7   change was the same.

8             DR. ALLEN:  Which suggests that maybe what

9   you want to consider is--I mean you are doing it as

10   a matched pair, that what you want to look at is

11   the maximum percentage decrease for each matched

12   pair rather than for the mean.

13             You know, in somebody in the control arm

14   has a very low recovery, it is their ownselves

15   transfused back into themselves, maybe what you

16   ought to look at is a maximum percentage change for

17   each matched pair rather than for the mean.

18             DR. GOLDSMITH:  But each person still has

19   to have a certain level of recovery, otherwise, it

20   is not valid.

21             DR. ALLEN:  You could still put that on.

22             DR. GOLDSMITH:  The reason I think 50

206

1   percent here is you think about it clinically, even

2   though this is just a chromium survival, which is

3   not really a red cell survival, it is a survival of

4   label that is modestly associated with the red

5   cells, as a clinician, I don't want to give a bag

6   of blood of which half of it disappears somewhere,

7   and doesn't circulate.

8             To me, that is kind of a cutoff just from

9   a clinician's perspective.  You want half of it at

10   least to survive.

11             DR. ALLEN:  I would certainly at least

12   agree with you on that.  I am less concerned if the

13   person in the control arm goes from, let's say, 62

14   percent down to 52 percent, than if the person goes

15   from 82 percent to 52 percent, which is maybe why

16   looking at individual controls ought to be a

17   consideration.

18             Dr. Bianco.

19             DR. BIANCO:  Celso Bianco, America's Blood

20   Centers.  I want in a certain way to support what

21   you are saying, and more, I think that we are stuck

22   to the 75 percent, and if we could select subjects

207

1   for those studies that were all above 75 percent,

2   then, we wouldn't have a problem, but we can't.  We

3   just enlist these individuals and then we discover

4   that they are 68 percent or something like that.

5             So, it would be fair to consider reduction

6   instead of the 75 percent cutoff, and assume that

7   you cannot have more than, I don't know, 10 percent

8   or 20 percent loss, or whatever it is, by

9   re-analyzing the data and not be so fixed on the 75

10   percent, or consider everybody that starts at below

11   75 percent as an outlier and only consider the ones

12   that are within what you would expect as the mean

13   of the population.

14             DR. ALLEN:  Thank you.

15             Dr. Strong.

16             DR. STRONG:  I am always concerned about

17   having a 100 percent requirement for anything in a

18   biological study from a research perspective.

19   Having a study completely invalidated on the basis

20   of one sample seems inappropriate, and we certainly

21   wouldn't do that in any other setting, you know,

22   you would apply outlier statistics when you have

208

1   something that is that far out, or you would

2   eliminate that particular sample or donor from

3   future studies or other studies.

4             But if you look at the ultimate conclusion

5   of this approach, you could run 100 samples and

6   have every working, and all of a sudden you have

7   one donor that drops down below 50 percent, and

8   your entire study is gone, I mean that doesn't seem

9   like a very good option.

10             DR. VOSTAL:  Maybe if I could suggest

11   something, usually, when we see the companies, the

12   companies approach us with a study design, we give

13   them an option of either meeting the absolute

14   standard and the donors of 75 percent, or doing a

15   control study alongside, so either a single arm or

16   a double arm study.

17             Most companies opt to do the single arm,

18   and the risk there is if you do get somebody who is

19   a normal person, but has lower recovery, you

20   wouldn't know that because you could always blame

21   it on the procedure, but if you include or require

22   a control arm in every study, you will be able to

209

1   eliminate the persons who don't have good recovery.

2             So, one option would be to always request

3   a control arm in these studies.

4             DR. ALLEN:  Yes, Dr. Schreiber.

5             DR. SCHREIBER:  I am not quite sure on

6   your last point.  Are you saying that there is

7   another standard for evaluation of the studies if

8   someone comes to you and say I want to do a matched

9   analysis, that the 75 percent or whatever the magic

10   number doesn't count, and what you are really

11   looking at is the type of difference that Dr.

12   Leitman and Dr. Strong suggested, we are looking at

13   changes in the individual?

14             DR. VOSTAL:  No, in that type of design,

15   the individuals that have poor recoveries with

16   standard products, you know, with their own red

17   cells, if they fall below the 75 percent, they

18   would not be included in the final analysis.

19             DR. LEITMAN:  There are so many reasons

20   why that low value can occur, so it can be

21   increased fragility of that subject's cells related

22   to something you can't even measure, some

210

1   hemoglobinopathy, something that is real that

2   happens in normal donors, and we know that every

3   now and then we see a unit transfused, and there is

4   no bump in the hemoglobin, and we say let's try

5   another unit, and the bump is fine.

6             So, we know that that happens, and they

7   are not irradiated necessarily, or there could be

8   something in processing, maybe someone irradiated

9   to 3,500 or 5,000 rad although it said 2,500.  Lots

10   of things can go wrong in processing.

11             So, we don't know why.  I certainly

12   believe that number, but there are so many reasons

13   why that could happen as a single outlier.  So,

14   again, just to agree with everything that has just

15   been said, to hold the whole study invalidate, not

16   invalidate, the study is valid, but negate the

17   potential for moving on with that type of

18   processing or type of storage solution because of

19   one data point, it seems excessive.

20             DR. ALLEN:  I want to focus on these four

21   options. So, you are arguing that we ought to

22   consider preferably option (a) or option (b) rather

211

1   than (c) or (d), which kind of puts a limit on the

2   outlier.

3             DR. LEITMAN:  I don't think that we can

4   use (b) because the data don't support its use.

5             DR. ALLEN:  I am sorry?

6             DR. LEITMAN:  The one I am supporting is

7   (a).

8             DR. KUEHNERT:  What about (a) plus

9   requiring a control arm?

10             DR. LEITMAN:  You always need a control

11   arm.

12             DR. VOSTAL:  Actually, we don't, because

13   we have an absolute standard of 75 percent, so you

14   could run the study with a single arm, and if the

15   mean comes in at 75 percent and all the other

16   criteria is met, then you would get an approval.

17             If you are unlucky and one of those donors

18   has a very low recovery, you cannot eliminate them

19   from the single arm study, but you could eliminate

20   them if you had a control arm.

21             DR. LEITMAN:  Could you repeat that?  What

22   is the proposal?  The proposal is not to have a

212

1   control arm at all?

2             DR. ALLEN:  No, he is just saying you

3   wouldn't have to have a control arm given an

4   absolute standard.

5             DR. VOSTAL:  That is what we have

6   currently.

7             DR. KUEHNERT:  I think some of the

8   discussions have led to the conclusion, I mean

9   there should be a control arm, so should that be

10   another option?  I am not sure how it would be

11   phrased, though.

12             DR. ALLEN:  I think the committee could

13   certainly make the point to the FDA without taking

14   a formal vote on it that, you know, control studies

15   have utility beyond just meeting a certain

16   standard.

17             Donna.

18             DR. DiMICHELE:  What I was going to add

19   was if you actually look--I was just trying to look

20   back very quickly at some of the control data that

21   was presented to us today, and it is true, none of

22   them do meet the uniform criteria of greater than

213

1   60 percent, but based on what we have seen and what

2   most people are presenting, given the N of 8 or the

3   N of 10, oftentimes we are dealing with either 1 or

4   2 outliers, so that if there really was to be a

5   greater stringency criteria and we wanted to have

6   that 75 percent, just say we have that 75 percent.

7             Then, we want to make it tighter by making

8   sure that no fewer than 30 percent of the samples

9   actually fall out of that 75 percent range, but

10   there needs to be a minimum, but maybe that minimum

11   can apply to 85 percent of the sample, for

12   instance, as opposed to 100 percent of the sample,

13   which is what is being proposed.

14             Then, I think it would almost fit a little

15   bit better with the data that was presented today,

16   so that if you wanted to use that kind of criteria,

17   you could say, for instance, okay, you know, 75

18   percent is the mean, we want 70 percent of the

19   sample to be above 75 percent, but our minimum,

20   let's say, is 60 percent, and at least 85 percent

21   of the cohort has to come in above 60 percent.

22             I don't know if that is the kind of

214

1   quality control that you are kind of looking for.

2   It might fit better with the data that we have

3   seen.

4             DR. KIM:  I would like to make it clear

5   about the criteria that you are talking about.  The

6   (b) options, you see the 95 percent one-sided lower

7   confidence limit for the population proportion of

8   successes is greater than 70 percent.

9             If that confuses you to understand what is

10   going on, I want you to look at the table.  The

11   table, when you have a sample size of 24, allowing

12   3 failures will meet the criteria.  So, the outlier

13   that we are talking about, the 1 or 2, can be

14   considered outlier, but not 30 percent or 40

15   percent.  So, we have a room for making failure.

16             DR. DiMICHELE:  I think we were actually

17   addressing the minimum criteria that you have in

18   (c), as well, that there shouldn't be any outliers

19   under 60 percent. Isn't that what you are

20   implicating in (c)?

21             DR. KIM:  The (c) part is the minimum of

22   all the individuals must meet the 60 percent.

215

1             DR. DiMICHELE:  That is what I was

2   suggesting.

3             DR. KIM:  The other option, the (b) says

4   you do have room for the making of failures,

5   allowing the outliers.

6             DR. ALLEN:  I think what concerns me in

7   the tables is that you have study sizes of 20 and

8   24, and, in fact, we didn't have any of the

9   presented studies today that went that high.

10             DR. KIM:  This is the combined data.  I

11   think all the previous presenters, they separated

12   sites.  Site, 1, they reported, like, a sample size

13   of 6 or 7 or 8.

14             DR. ALLEN:  I think we got up to 16.  I

15   didn't see anything above 16.

16             DR. KIM:  But there was 24.

17             DR. VOSTAL:  Our current standard is at

18   least 20 donors.

19             DR. ALLEN:  Okay.

20             DR. KUEHNERT:  If I could make just one

21   more suggestion about this minimum number because

22   it seems to be hanging everyone up on it.  It

216

1   definitely sets a concerning precedent about making

2   outliers basically shoot down a study.

3             What about something like a 99 percent

4   confidence interval and setting some rate for that?

5   I am searching for something more statistical that

6   addresses this minimum number of all samples part

7   of this.  I am just wondering if that might be

8   useful.

9             DR. FITZPATRICK:  Before the committee

10   votes on this or makes a decision, I want to go

11   back to something Jay said earlier, which I think

12   is the fundamental question, should FDA approve a

13   product that has a 40 percent or greater failure

14   rate when we are talking failure or success.

15             I would like you to turn around that

16   thinking.  The fundamental question is should the

17   FDA approve a process that either improves or what

18   is the impact on the cell that we are discussing.

19             The failure rates, I think we are

20   misleading ourselves into thinking that if you set

21   the standard, then, 75 percent of the transfusions

22   are going to have recoveries greater than what you

217

1   say, because you can see in just the small numbers

2   of people that are being used for these samples,

3   the diversity of, as Dr. Goldsmith said, the impact

4   of assessing chromium survival rates of cells.

5             We are not really doing red cell

6   survivals, we are doing chromium labeled survivals,

7   and that has an impact on cells, and it impacts

8   different individual cells differently.  As Celso

9   said, if you want to go to the fundamental

10   question, what is the impact of what you are doing

11   to those cells, and I think the only way you can

12   assess that is to say, okay, here is chromium

13   labeled survival of this individual cells with

14   nothing done to them, or with something done to

15   them in a standard solution that we have been using

16   for years, so we are obviously not going to

17   delicense and change.

18             Here is the change in that individual

19   cells when we did something to them, we irradiated

20   them, we changed the additive solution, did it make

21   it worse or did it make it better?  That is the

22   fundamental question.  If it made it worse, how

218

1   much worse are you willing to tolerate in order to

2   prevent graft-versus-host disease, if it made it

3   better, well, you probably did a good thing.

4             That, I think is what should be being

5   modeled, and I don't think any of those proposals

6   modeled it.

7             DR. ALLEN:  It is 12:36.  We actually do

8   have a deadline for getting into the lunchroom

9   here, and we have got a second set of questions

10   which we can always come back after lunch and take

11   up if we need to.

12             I would like to move this forward, and I

13   am impressed that we don't seem to be arriving or

14   coming together in a consensus on the options.  We

15   are still quite far apart.  Can we entertain

16   discussion around trying to come towards a

17   consensus, and if these four don't meet it, I am

18   hesitant to try and have us craft language today

19   that will do that.

20             Maybe as an escape option, we could craft

21   text language that doesn't come up with specific

22   parameters, that provides guidance for the FDA, but

219

1   I sense, Dr. Epstein, that you really would like to

2   have us come up with a more formal recommendation

3   that includes numbers, is that correct?

4             DR. EPSTEIN:  I think that what we have

5   heard from the discussion is that--I mean the

6   committee can vote on these propositions if they

7   wish--but that the sense of the discussion is that

8   there should be a preference or a bias toward

9   control studies in which the subject is its own

10   control, that that would enable us to get away from

11   the issue of outliers in the studies, and that in

12   doing so, we may wish to consider a criterion where

13   we look at the relative percent recovery of a

14   process versus control in the same subject.

15             So, I think that there is a general sense

16   that that is a better pathway.  It leaves open the

17   question of if we still allow companies to submit

18   one-arm studies.  The company runs a greater risk

19   because if you have the outliers, you can't tell

20   whether they are caused by a bad process or they

21   are caused by donor variability.

22             So, the companies will need to understand

220

1   that upfront, but the FDA shouldn't be put in the

2   position of hand waving when we get data that shows

3   a lot of outliers, because we can't answer the

4   question whether it was donor variability without

5   the subject being his or her own control.

6             So, I would suggest that the advice that I

7   am hearing from the committee, although you have

8   not explicitly stated it, is that you do endorse a

9   strategy where we look at the percent failures in

10   the trial, but that it should be flexible enough to

11   allow for some level of outliers.

12             Now, to my way of thinking, that is really

13   option (b), but I don't know that you have to vote

14   it to send us that message.  We can go back and

15   think about it.

16             Those are my takes on the discussion.

17   Now, whether we need votes on these specific

18   proposals, you know, with their numbers, I am not

19   sure that is productive anymore at this stage.

20             It may be better for us to take these more

21   general messages and come back at a later date with

22   a concrete proposal.

221

1             DR. ALLEN:  Without taking a formal vote,

2   may I see a show of hands of people who are in

3   agreement with Dr. Epstein's summary statement, and

4   then we will ask for those people who are proposed.

5             All those who favor Dr. Epstein's

6   statement?

7             [Show of hands.]

8             DR. ALLEN:  All who would like further

9   discussion who do not agree with Dr. Epstein's

10   statement?

11             DR. LEITMAN:  I agree with everything Dr.

12   Epstein stated except his choice of option (b),

13   because I stated before that the control group

14   won't meet that standard, and never has in any

15   study.

16             DR. ALLEN:  Dr. Smallwood.

17             DR. SMALLWOOD:  Excuse me.  I don't want

18   to prolong it, but the procedure is that I do have

19   to take a roll call vote, so if you are voting on

20   anything, I will have to take a roll call vote.

21             DR. ALLEN:  I was asking just for a sense

22   of the committee.  As I said, it is not a formal

222

1   vote.

2             DR. DAVIS:  Are we voting whether to vote?

3             DR. ALLEN:  Dr. Davis, you may state a

4   preference if you have one.

5             Dr. Epstein, are you satisfied with that

6   as a conclusion without a formal vote?

7             DR. EPSTEIN:  Yes, I am prepared to accept

8   discussion without voting on this issue.  I do

9   think it is important to have some general sense of

10   the committee whether those summary statements

11   reflect the discussion, because it's my summary,

12   not your summary.

13             DR. ALLEN:  That was unanimous with the

14   except that Dr. Leitman very specifically stated

15   that she is not in favor of option (b), which you

16   sort of lumped in with that.

17             DR. EPSTEIN:  The statement is that

18   control studies would not pass option (b) as a

19   criterion, but I think it is not actually true,

20   Susan.  I think we have lots of control data that

21   pass option (b), because the issue is whether the

22   lower confidence bound for the mean is above 70

223

1   percent, and that is true in studies that can have

2   significant numbers of outliers.

3             In other words, what we are saying is that

4   the proportion of failures can't exceed 30 percent

5   as the lower confidence range for proportion of

6   failures, and that construct does allow significant

7   number of outliers.

8             In the data sets that you have seen, the

9   controls, where they have been done, have passed,

10   so I am not sure I fully can accept the underlying

11   premise for your comment.

12             I am not saying we should simply adopt

13   option (b) by fiat.  I am just saying that a

14   concept along the lines of allowing outliers, but

15   having some constraint over the proportion of

16   outliers, as for example expressed on construct

17   (b).  It doesn't have to be construct (b), but to

18   allow some proportion of outliers.

19             DR. LEITMAN:  That is acceptable to me.

20             DR. ALLEN:  Dr. Schreiber.

21             DR. SCHREIBER:  I was just going to

22   suggest in the control study, really, what we want

224

1   to do is have our limit based on the treatment

2   recovery, and that would fulfill your argument,

3   because then you have your control and your

4   treatment as long as you have a 50 percent

5   difference, whatever the difference is, you are

6   okay, and that would eliminate the outlier problem

7   completely.

8             DR. ALLEN:  Any further comments or

9   questions on this set of questions?  If not, we

10   have resolved without a formal vote the questions

11   on RBC, red blood cell in vivo recovery acceptance

12   criteria, and we will go back to the questions on

13   the dating period for gamma irradiated RBCs and try

14   to conclude that within 15 minutes or so, so we can

15   break for lunch in a timely fashion.

16             The first question was:  Do the committee

17   members agree that the current recommendations

18   regarding the dating period of gamma irradiated red

19   blood cells should be modified?

20             Dr. Klein.

21             DR. KLEIN:  I am just going to repeat my

22   earlier comment, that I see no compelling need at

225

1   this point in time to make any modification based

2   on the data that I have seen today.

3             DR. ALLEN:  Other comments or questions?

4   Are we ready to call the question?

5             Dr. Smallwood.

6             DR. SMALLWOOD:  I will repeat the

7   question, so that it can be entered into the

8   record.

9             Do the committee members agree that the

10   current recommendations regarding the dating period

11   of gamma irradiated red blood cells should be

12   modified?

13             Dr. Allen.

14             DR. ALLEN:  No, they do not need to be

15   modified.

16             DR. SMALLWOOD:  Dr. Davis.

17             DR. DAVIS:  Yes.

18             DR. SMALLWOOD:  Dr. DiMichele.

19             DR. DiMICHELE:  No.

20             DR. SMALLWOOD:  Dr. Doppelt.

21             DR. DOPPELT:  No.

22             DR. SMALLWOOD:  Dr. Goldsmith.

226

1             DR. GOLDSMITH:  No.

2             DR. SMALLWOOD:  Dr. Klein.

3             DR. KLEIN:  No.

4             DR. SMALLWOOD:  Dr. Laal.

5             DR. LAAL:  No.

6             DR. SMALLWOOD:  Dr. Harvath.

7             DR. HARVATH:  No.

8             DR. SMALLWOOD:  Dr. Kuehnert.

9             DR. KUEHNERT:  No.

10             DR. SMALLWOOD:  Dr. Leitman.

11             DR. LEITMAN:  No.

12             DR. SMALLWOOD:  Dr. Quirolo.

13             DR. QUIROLO:  No.

14             DR. SMALLWOOD:  Dr. Schreiber.

15             DR. SCHREIBER:  No.

16             DR. SMALLWOOD:  Dr. Whittaker.

17             DR. WHITTAKER:  No.

18             DR. SMALLWOOD:  Ms. Knowles.

19             MS. KNOWLES:  No.

20             DR. SMALLWOOD:  And, Dr. Strong, how would

21   you vote if you could?

22             DR. STRONG:  I will try.  No.

227

1             DR. SMALLWOOD:  The results of voting,

2   there are 12 No votes, 1 Yes vote, and the industry

3   representative agreed with the No vote.

4             DR. ALLEN:  Any further comments or

5   questions?

6             [No response.]

7             DR. ALLEN:  If not, we will adjourn for

8   lunch.  I would like to have people here by quarter

9   of 2:00, so that we can come back into session at

10   1:45 p.m.

11             [Whereupon, at 12:45 p.m., the proceedings

12   were recessed, to be resumed at 1:45 p.m.]

228

1             A F T E R N O O N  P R O C E E DI N G S

2                                                    [1:45 p.m.]

3             DR. SMALLWOOD:  I would like, before we

4   begin, to make a correction to the vote taken

5   before lunch.  The vote on the previous topic, I

6   gave the vote as being 12 No votes and 1 Yes vote.

7   The correction is that there were 13 No votes and 1

8   Yes vote.  The vote of the consumer rep was omitted

9   when I was counting, but she is eligible to vote

10   and to be counted.

11             Also, as I mentioned, there would be a

12   brief announcement regarding the conflict of

13   interest statement for this particular topic, so at

14   this time, I will read it for your hearing

15   pleasure.

16             This brief announcement is an addition to

17   the conflict of interest statement read at the

18   beginning of the meeting, and it is part of the

19   public record for the Blood Products Advisory

20   Committee meeting on July 22nd.

21             This announcement addresses conflicts of

22   interest for Topics II and III.  Drs. Liana

229

1   Harvath, Matthew Kuehnert, Susan Leitman, Keith

2   Quirolo, George Schreiber, Donna Whittaker, and Ms.

3   Katherine Knowles have been appointed as temporary

4   voting members for this meeting.

5             Dr. Michael Strong is participating in

6   this meeting as a non-voting industry

7   representative acting on behalf of regulated

8   industry.

9             The Food and Drug Administration has

10   prepared general matter waivers for the special

11   government employees participating in this meeting

12   who required a waiver under Title 18, United States

13   Code 208.

14             A speaker for Topic II, Dr. Edward Snyder

15   is employed by the Yale-New Haven Hospital Blood

16   Bank.  He also has associations with clinical

17   trials that involve red blood cells.

18             Dr. James AuBuchon has grants and/or

19   contracts with firms that could be affected by the

20   committee discussions.  He is also a scientific

21   advisor for several affected firms.

22             A speaker for Topic III, Dr. Steven

230

1   Kleinman receives consulting fees from two firms

2   that could be affected by the committee

3   discussions.

4             Dr. Jerry Holmberg has a financial and

5   professional interest in several firms that could

6   be affected by the committee discussions.

7             In addition, there may be regulated

8   industry and other outside organization speakers

9   making presentations. These speakers have financial

10   interests associated with their employer and with

11   other regulated firms.  They were not screened for

12   these conflicts of interest.

13             Again, if there is any update to any of

14   these statements that I have read, please be

15   advised to announce them before you speak.

16             Thank you.

17             At this time, I will turn the proceedings

18   of the meeting again to the Acting Chairman, Dr.

19   James Allen.

20             DR. ALLEN:  Thank you, Dr. Smallwood.  I

21   will note for the record that all of our federal

22   employees over here are busy with their

231

1   blackberries.  It is just interesting to see, the

2   whole left side.

3             Our second topic of discussion is a New

4   Standard for Platelet Evaluation.  The background

5   statement will be made by Dr. Salim Haddad of the

6   Food and Drug Administration.

7             II. New Standard for Platelet Evaluation

8                   Introduction and Background

9                        Salim Haddad, M.D.

10             DR. HADDAD:  Thank you.  Good afternoon.

11   My topic in this session is the proposed standard

12   for platelet evaluation.  More precisely, it is the

13   standard on the use of radiolabeled platelets for

14   the assessment of in vivo viability of platelet

15   products.

16             Before delving into the substance of the

17   topic, I would like to place it in its proper

18   context, which is that of the FDA evaluation of the

19   efficacy of platelet products.

20             [Slide.]

21             We can start be defining platelet

22   efficacy.  It is the ability of the transfused

232

1   platelets to circulate for the expected life span

2   after transfusion and for their ability to

3   participate in hemostatic processes to prevent or

4   stop bleeding.

5             [Slide.]

6             What kind of testing does FDA look for

7   when we are evaluating a new platelet product?

8   Well, that depends on our level of concern.  This

9   what we refer to as our pyramid of concerns.

10             At the bottom, we have the minimal

11   concerns and at the top, the major concerns.  At

12   the bottom, for example, for current storage

13   conditions, we like to see in vitro studies on

14   platelet biochemistry and physiology, and I will

15   have a few words to say about in vitro testing.

16             As we move up to more serious concerns,

17   such as for a new five- or 7-day storage container

18   or for a new apheresis collection device, then, we

19   will require radiolabeling studies in human

20   volunteers.

21             [Slide.]

22             However, for pooled platelet products,

233

1   instead of radiolabeling studies, then, we would

2   accept post-transfusion assessment in

3   thrombocytopenic patients rather than radiolabeling

4   studies in human volunteers.

5             The reason is that it would be unethical

6   to subject the human volunteer to products for a

7   number of donors since the issue is pooling, and

8   the assessment can be done using the CCI, the

9   corrected count increment.

10             At the top of the pyramid, usually, we are

11   dealing with new methodologies that may alter

12   platelet physiology with unexpected consequences on

13   platelet performance, and fitting into that

14   category would be the platelet substitutes and

15   chemically modified platelets like in pathogen

16   reduction, and for those kind of products, we do

17   like to see hemostatic clinical trials.

18             [Slide.]

19             Obviously, we run into some gray areas

20   that do not fit neatly into prespecified

21   categories.  For minor modifications to current

22   storage conditions, that would probably fit between

234

1   radiolabeling studies and in vitro studies,

2   whereas, new storage media or extension beyond 7

3   days, that would probably fall between

4   radiolabeling studies and hemostatic clinical

5   trials.

6             [Slide.]

7             As I mentioned earlier, the first line of

8   evaluation would be the in vitro testing, and this

9   is a list of some of the basic in vitro tests that

10   we like to see. Obviously, the platelet count of

11   the product would be of value.  Under the

12   morphologic category, the microscopic examination

13   of the platelets would give us an indication as to

14   the shape of the platelets.

15             Normal platelets circulate as discs,

16   whereas, spheroidal platelets are usually

17   associated with some level of damage that may

18   affect their circulation and maybe their function.

19             The MPV is the mean plated volume, can be

20   helpful in certain instances.  In terms of

21   metabolism of the platelets, platelets consume

22   glucose and oxygen, and will produce carbon dioxide

235

1   and lactic acid.  The lactic acid will decrease the

2   pH.  LDH can be an indication of the platelet

3   lysis.

4             Platelets, once activated, they will have

5   in increase in surface expression of the CD 62,

6   which is also known as P-selectin, and the

7   platelets can display specific physiologic

8   responses when they are exposed to certain

9   conditions, for example, for the hypotonic shock

10   response, the platelets are placed in a hypotonic

11   solution, the platelets would swell, and the

12   healthy platelets that have maintained its membrane

13   integrity and its energy metabolism will be able to

14   extrude water and regain its origin and shape.

15             The extent of shape change would happen

16   when the platelets are put in contact with an

17   agonist, such as ADP or thrombin, and the

18   disc-shaped platelets would rapidly turn into

19   discoid, into spherical shape.

20             Of all those in vitro tests, the most

21   reliable would be extent of shape change, hypotonic

22   response, the pH, and the morphology score.

236

1             [Slide.]

2             The problem with in vitro testing is that

3   they did not correlate well with in vivo viability,

4   and over time, there isn't a single individual test

5   that has stood out as a surrogate marker for

6   platelet efficacy.

7             However, those tests, they are still

8   performed and FDA still requires them because they

9   constitute a screening process to eliminate those

10   procedures that would clearly result in suboptimal

11   products.  Additionally, you would avoid subjecting

12   a volunteer to radioactivity from an unwarranted

13   radiolabeling study.

14             [Slide.]

15             So, the next step for the evaluation is to

16   look at the in vivo radiolabeling studies.  This is

17   really our main topic here.

18             These studies are conducted to evaluate

19   the platelet survival in the circulation, and the

20   assumption is that viable circulating platelets

21   that have shown no defect in the individual phase

22   of the testing would be able to participate in the

237

1   physiological responses that constitute clinical

2   platelet efficacy.

3             As such, those studies act as surrogate

4   markers for hemostatic efficacy.  The way it works

5   is that if you are evaluating a new platelet

6   product, which would be a test product, that

7   product would be labeled, it would be reinfused

8   back into the healthy autologous volunteer from

9   whom it was drawn, and then with any level of

10   damage, the new platelet product would be cleared

11   at the faster rate than a control platelet product

12   in the paired comparison testing.

13             [Slide.]

14             So, this is our current approach for

15   evaluating radiolabeling studies.  This is the

16   approach that we are seeking to replace with a new

17   approach.  The main difference between the two

18   approaches is the choice of the control.

19             In the current approach, the control is an

20   old, established licensed product.  For example, if

21   we are evaluating a 7-day apheresis platelet, the

22   control would be a 5-day platelet product.

238

1             Here is how it would work.  The donor

2   comes in on day zero.  You have a unit of apheresis

3   platelets that is withdrawn, stored out to 5 days.

4   On day 5, a sample is taken from that apheresis

5   unit.  It is radiolabeled with either chromium or

6   indium, and reinjected back into the donor.

7             Two days later, on day 7, another sample

8   is taken from that same unit, labeled with the

9   alternate isotope, and then infused back into the

10   donor.

11             [Slide.]

12             Then, we would trace the radiolabeling in

13   that same subject of the two products.  So, here we

14   have the survival curves of the two products where

15   they link the Y axis, that would be the recovery,

16   and where they intersect with the X axis would be

17   the survival, and here I am oversimplifying.

18             You would notice that 7-day platelet

19   product would obviously have a lower recovery and a

20   lower survival than the 5-day product.  The way you

21   would compare those two products in the current

22   approach is that you would look at the difference

239

1   in recovery and the difference in survival, and you

2   would allow for about 10 to 20 percent difference,

3   because we think that such a difference would not

4   be clinically significant and also allow for some

5   variability in experimentation results.

6             [Slide.]

7             However, there are problems with this

8   current approach, which is that we have no minimum

9   standard set for platelet quality.  All we are

10   doing is testing a new product, comparing it to an

11   established product, and allowing 10 to 20 percent

12   difference.

13             The repeated application of this approach

14   to successive products can lead to what has been

15   described as the "inferiority creep" in the quality

16   of the product, so this observation, this concern

17   has prompted a call.

18             [Slide.]

19             A call for a new standard for

20   radiolabeling studies.  This standard has been

21   first voiced by Dr. Scott Murphy in August 2002 at

22   an FDA Workshop.  Dr. Murphy is a recognized

240

1   authority in platelet research and platelet

2   radiolabeling study.

3             [Slide.]

4             He published his proposal for a standard

5   in the January 2004 Transfusion journal.

6             [Slide.]

7             The main tenets of this new standard is

8   that now the control is no longer an established,

9   licensed product, but rather it is a fresh platelet

10   specimen, and I will qualify further that fresh

11   specimen in the next slide.

12             So that fresh platelet specimen would be a

13   reference specimen, would be a standard of quality

14   for the testing.  Any new product would be compared

15   to that fresh platelet, and the outcome measures,

16   the outcome parameters, recover and survival, would

17   be expressed as a percentage of the control, which

18   again is the fresh platelets.

19             [Slide.]

20             So, that is how it would work in the new

21   standard. We have the donor.  On day zero, he

22   donates an apheresis platelet unit.  This time it

241

1   is stored out until day 7, and again, here, we are

2   evaluating a day 7 product, so it is stored out to

3   day 7.  We no longer have a day 5 product.

4             On day 7, the donor comes back.  This

5   time, a fresh blood specimen is collected, about

6   maybe 50 to 100 ml, and Dr. Snyder will elaborate

7   further on the proper protocol for the collection

8   of the specimen.

9             A platelet concentrate is prepared from

10   that blood specimen and then it is radiolabeled and

11   reinfused back into the donor.  Then, on that same

12   day 7, a specimen is withdrawn from the platelet

13   bag, radiolabeled with the other tag, and then

14   infused back into the donor.

15             [Slide.]

16             Again, we would follow the survival curves

17   of those two products.  Again, the fresh, you would

18   expect a better recovery and survival than with the

19   stored platelets. However, this time, to compare

20   the two products, we are no longer looking at the

21   difference like we did before, but we are looking

22   at the ratio.

242

1             So, the new parameter, that is the percent

2   recovery of the test would be the recovery of the

3   test product divided by the fresh x 100, and for

4   the survival, it would the survival of the test

5   product divided by the survival of the fresh x 100.

6             [Slide.]

7             A workshop on this new standard was

8   convened in Bethesda in early May with the stated

9   goals of orienting the transfusion community

10   towards this new approach and to obtain the

11   opinions of the experts in the field.

12             [Slide.]

13             At that workshop, there were discussions

14   on the rationale and merits of the new approach.

15   These was presentation on appropriate study

16   protocols for comparing platelet products to the

17   standard along with the statistical approaches, and

18   also we had preliminary data presented using the

19   new approach.

20             At the end of the day, an expert panel was

21   convened and there was initial agreement between

22   FDA and the panelists to adopt this new approach.

243

1   I say "initial" because we did indicate during that

2   workshop that you would be taking up the issue to

3   this committee.

4             There were also discussions on

5   establishing minimal performance levels for

6   platelet recovery and survival.

7             [Slide.]

8             Now, Dr. Murphy, in his original proposal,

9   he suggested setting the bar for the recovery at 66

10   percent, so a new product having a percent recovery

11   of 66 percent or higher would be acceptable, and

12   again, that would be the recovery of the test or

13   fresh.

14             There was common agreement again between

15   FDA and the panelists that 66 percent sounds like a

16   reasonable bar to set.

17             [Slide.]

18             For the survival, Dr. Murphy, in his

19   proposal, suggested setting the ratio at 50

20   percent, and the expert panel concurred with him.

21   His rationale for being more lenient with the

22   survival than with the recovery is based on this

244

1   article where the authors, Hanson and Slichter,

2   where they demonstrated that there is a shortened

3   platelet survival in thrombocytopenic patients

4   irrespective of the cause of the thrombocytopenia.

5             The reason is that in any individual, you

6   have a fixed daily loss of about 7,000 platelets

7   per microliter that leave the circulation early to

8   support the endothelium. So, in thrombocytopenic

9   patients, that fixed loss would constitute a high

10   percentage of the actual platelet count in the

11   patient, hence, the shortened survival.

12             The second argument for being more lenient

13   on the survival than on the recovery is that in

14   actual prophylactic transfusion, the time to next

15   transfusion is usually 2 to 3 days.  It is rarely

16   more frequent or rarely more than that, and the

17   platelets actually never live up to their full

18   potential of 7 to 9 days.

19             So, due to those two arguments, the

20   survival seems to be less than important than the

21   recovery.

22             As I mentioned, the expert panel concurred

245

1   with Dr. Murphy, however, our position, FDA

2   position is that we prefer to have a 66 percent as

3   the ratio.

4             Our reasoning, so it would be 66 percent,

5   the same level as for the recovery, and our

6   argument is that those radiolabeling studies are,

7   after all, being conducted in healthy volunteers,

8   not in thrombocytopenic patients, so that increase

9   the platelet clearance mechanism that I just

10   detailed that occurs in the thrombocytopenic

11   patients, does not occur in the healthy volunteers,

12   so there would be no reason to make correction or

13   adjustment for it.

14             Also, we are trying to set a quality of

15   standard against which all platelet performance

16   would be measured.

17             Our second argument is that the platelet

18   product performance, as obtained via those

19   radiolabeling studies, would be a reflection of the

20   product processing conditions from collections all

21   the way through storage without any extraneous

22   factors, divorced from any external factors.

246

1             We think that the standard should be set

2   accordingly.

3             Our third concern is that additive effects

4   of damage from the storage and also an increased

5   clearance in the thrombocytopenic patients by the

6   mechanisms I detailed earlier, so since the product

7   is going to take two hits, one might as well start

8   with a higher standard.

9             [Slide.]

10             These are examples of some of the products

11   that would be subjected to this new standard - new

12   storage containers, extended platelet shelf life,

13   additive solutions for platelet storage, low

14   temperature storage conditions, and

15   pathogen-reduced platelets.

16             [Slide.]

17             Now, what if a product fails to meet the

18   criteria? It wouldn't necessarily be the end of the

19   road for that product because those with

20   alternative merits, such as pathogen reduction or

21   extended shelf life, they could be licensed if the

22   benefits outweigh the shortcomings.

247

1             However, they would have to be labeled

2   accordingly and each alternative product will have

3   to be considered on a case-by-case basis.

4             [Slide.]

5             So, in conclusion, FDA has committed to a

6   gold standard for platelet product performance.

7   This would be a fixed standard to maintain platelet

8   product quality over time.  This would allow for a

9   uniform and less subjective regulatory review

10   process with common research protocols to minimize

11   differences in methodology and improve

12   inter-laboratory compatibility, and a standard of

13   this nature would facilitate product development in

14   a competitive but fair environment.

15             So, there are too many issues here that we

16   are seeking advice from the committee are whether

17   the new approach would constitute scientific

18   advancement compared to the old approach, and we

19   would like that comment on what to set the criteria

20   for recovery and survival.

21             DR. ALLEN:  Thank you.  Would you just

22   refresh my memory, I believe Dr. Murphy's 66

248

1   percent was selected based on some physiologic

2   data, but it was just an arbitrary selection point

3   that he chose.

4             DR. HADDAD:  Right, essentially, it is

5   arbitrary. I think it was basically set to kind of

6   meter [?] the recovery when you transfuse a

7   patient, you know, a third goes to the spleen, and

8   you have a recovery of two-thirds, but it is not

9   etched in stone.

10             DR. ALLEN:  Thank you.

11             Comments or questions?

12             DR. KUEHNERT:  I just had a quick

13   question, maybe I missed this, but you talked about

14   recovery and survival, and this is survival after

15   what time period?

16             DR. HADDAD:  You would do the

17   radiolabeling, you would reinfuse into the donor,

18   and then you start taking samples, you know, from

19   time zero all the way to 10 days, and then you

20   would trace the radioactivity, which would go down.

21             DR. KUEHNERT:  So, if you are seeing

22   survival greater than 66 percent or 50 percent,

249

1   that is in comparison to after how many days?

2             DR. ALLEN:  You would compare the survival

3   curve in the control platelet population, the

4   radiolabeled population versus the sample.

5             DR. KUEHNERT:  So, it is just

6   comparatively over however many days you are

7   studying, there is no limit.

8             DR. HADDAD:  Usually, survival, I said it

9   is the section of the curve with the x axis, but,

10   in fact, you have computer programs that work out

11   the survival and recovery. You have different

12   theories, multiple hit theory, you have the

13   weighted mean theory, and these can be defined by

14   computer programs.

15             DR. ALLEN:  Dr. Harvath.

16             DR. HARVATH:  I wanted to ask you a

17   question about your criterion for platelet survival

18   and how you came up with your position of equal to

19   or greater than 66 percent.

20             Did you have data that you worked with in

21   normals to come up with that?

22             DR. HADDAD:  This is a new standard, and

250

1   the data, we don't have much data using this new

2   approach.  At the workshop, there were preliminary

3   data, and I think today we will have more data.

4             But the reason I presented those three

5   points, what we think why the standard should be

6   set at the same level as the recovery, and we think

7   if you set the standard, we are willing, you know,

8   based on the data that will come up later on, to

9   maybe make adjustments if, let's say, the product

10   don't meet 66 percent.

11             However, we would like to set the standard

12   and rather than make adjustments from the get-go,

13   and dilute the standard, we prefer to set the

14   standard and then make any necessary adjustments

15   later on.

16             DR. HARVATH:  So, in one of the clinical

17   trials that a sponsor would come in with, instead

18   of using a platelet product in an autologous

19   setting, in a normal volunteer donor population,

20   they chose to go directly into a patient

21   population, let's say, because of a critical need,

22   would you still hold them, if the patients were

251

1   thrombocytopenic, to the 66 percent, or would you

2   be willing to work with the recommendation of Dr.

3   Murphy and colleagues at 50 percent?

4             DR. HADDAD:  You mean if the study would

5   be conducted in a thrombocytopenic patient?

6             DR. HARVATH:  In a thrombocytopenic

7   patient population.

8             DR. HADDAD:  Looking at CCI, for example.

9             DR. HARVATH:  Exactly.

10             DR. HADDAD:  Again, we would weigh the

11   benefits and the shortcomings, and if the product

12   has benefits, but does not meet all our criteria,

13   let's say, we might label it with altered

14   transfusion kinetics, this way the transfusing

15   physician would know that this product does not

16   have the same kinetics as the regular product,

17   therefore, maybe he needs to increase the frequency

18   of transfusion.

19             DR. LEITMAN:  If I could state something

20   that may be obvious, but in the morning, we were

21   talking about 75 percent as an absolute recovery of

22   cells infused.  The 66 and 50 percent here is not

252

1   an absolute recovery, it is not an absolute

2   survival, it's a proportion of fresh.

3             So, we are not talking about the recovery

4   of cells.  The standard is, because we don't have

5   any other, you use fresh platelets, and that

6   defines your 100 percent, if you will, that is the

7   best we can do, so the rest is as a percent of that

8   fresh.

9             Does that answer the earlier question?

10             DR. KUEHNERT:  That helps.  Thanks.

11             DR. ALLEN:  Other comments or questions?

12             DR. QUIROLO:  This is really an

13   enumeration of platelets and doesn't mean that they

14   work.  Has the FDA ever thought about PFA testing

15   for platelet function, has that ever been

16   considered as a standard?

17             DR. HADDAD:  The common wisdom and the

18   assumption for years now is that if the platelets

19   circulate, you assume that they are functioning.

20             Regarding the PFA, these have not been

21   validated. They need to be validated against

22   radiolabeling studies to be used as a replacement

253

1   for labeling studies.

2             DR. QUIROLO:  There is literature about

3   PFAs used in donors and recipients that indicate

4   that enumeration is probably not the gold standard

5   for function, at least that is my reading of the

6   literature.

7             DR. HADDAD:  The enumeration that I

8   mentioned, that was part of in vitro testing, so

9   obviously, you would like to know how much

10   platelets you have in the bag.

11             DR. ALLEN:  Thank you very much.

12             The second presentation on this topic is

13   by Dr. James AuBuchon, Dartmouth-Hitchcock Medical

14   Center.

15               Presentation - James AuBuchon, M.D.

16             DR. AuBUCHON:  Good afternoon, ladies and

17   gentlemen, and thank you to the FDA for asking me

18   to provide an update and a follow-up from the

19   conference held in May on the use of radiolabeled

20   platelets for the assessment of viability of

21   platelet components.

22             [Slide.]

254

1             The concept, as you heard Dr. Haddad

2   discuss, is that the normal subject would serve as

3   his or her own control using fresh platelets from

4   the subject establishing essentially the baseline

5   against which a platelet product that has been

6   collected, treated, or stored in a new manner could

7   be compared on reinfusion after radiolabeling.

8             [Slide.]

9             The conference looked at a number of

10   different issues, and I would like to walk you

11   through some of these in a little more detail than

12   previously you had an opportunity to see.

13             [Slide.]

14             There are a number of issues that had to

15   be addressed, and the first was although it was

16   understood that some type of standard would be

17   efficacious, should the standard come from the same

18   subject or should it be an absolute standard, as

19   you have heard discussed for red cells.  For red

20   cells, there is an absolute standard that has to be

21   met, but should we have a similar absolute standard

22   for platelets.

255

1             After some discussion and consideration,

2   it was the conclusion of the group that the

3   standard should come from the same subject, should

4   be based on fresh platelets in order to decrease

5   problems associated with intersubject variability

6   and procedural variability between laboratories.

7             [Slide.]

8             Indeed, this kind of approach will work.

9   We had the opportunity to work with Dr. Murphy and

10   develop an initial validation of this fresh

11   criterion using apheresis platelets that we

12   reinfused on day 1 and then, after further storage,

13   on day 5.

14             This was a licensed collection system or

15   licensed storage system.  We would expect the day 5

16   outcome here to be successful outcome, so this was

17   really intended as a validation of the criterion

18   proposed by Dr. Murphy rather than an attempt to

19   actually verify whether or not this particular

20   apheresis collection system was worthy of

21   licensure, which it had already been.

22             [Slide.]

256

1             The absolute recovery and the absolute

2   survival with the fresh platelets reinfused within

3   20 hours of collection, shown here, in comparison

4   to the absolute recovery and absolute survival as

5   measured on day 5, and then the ratio between the

6   two.

7             For both recovery and survival, as you can

8   see, the ratio between the stored platelets and the

9   fresh platelets met the criteria as suggested by

10   Dr. Murphy.

11             [Slide.]

12             Another issue that the panel had to

13   address is whether the fresh platelets should be an

14   aliquot of the unit taken shortly after collection

15   and then reinfused, as we did in the initial

16   validation, or should it be collected separately,

17   should the fresh platelets be collected on the day

18   of reinfusion of the stored or treated product.

19             If one used an aliquot from the unit that

20   was being stored, one would have the assurance that

21   the platelets being used as the fresh standard were

22   representative of what was in the bag.

257

1             However, if those platelets had been

2   collected by a technique that ultimately was

3   actually injurious to the platelets, then, the bar,

4   the standard established by the fresh platelets

5   might be inadvertently too low.

6             One could avoid that by collecting the

7   fresh platelets separately by a standard, and by

8   that we mean manual technique, on the day of

9   reinfusion of the stored platelets.  This would

10   avoid any potential damage to the fresh platelets

11   that would be unexpected because the new collection

12   technique was an unknown technique.

13             However, a problem arises there because of

14   the potential for actually having a different

15   population of platelets represented in the fresh

16   sample than the stored sample.  This is not a

17   concern probably when the platelets are being

18   produced from a whole blood collection, but it is a

19   potential concern when they are being produced via

20   apheresis, and apheresis collection could lower a

21   normal subject's platelet count by 30 or 40

22   percent.

258

1             As a result, the thrombopoietin level

2   could increase and the mean age of platelets, say,

3   a week or 10 days later, at the time at the end of

4   the storage when the study was going to be

5   performed, the mean age of platelets could actually

6   be younger.

7             So, that would inadvertently set the fresh

8   bar, the standard, too high against which the

9   stored platelets would then be compared.

10             After some discussion back and forth, and

11   some consideration of the pros and cons of the

12   different approach, it was ultimately decided by

13   the expert panel that the better way to go would be

14   to collect the fresh platelets separately on the

15   day of reinfusion of the test platelet using a

16   defined manual technique.

17             This would then allow the standard-setting

18   platelets to be collected outside of any commercial

19   system, be collected in the standard way 10 years

20   or 20 years from now, the same way that we are

21   doing today, even after some apheresis systems

22   might have left the market.

259

1             Another issue that was addressed is when

2   should these studies be performed, should the study

3   be performed on the last day of storage that the

4   manufacturer is looking for, that is, if it is

5   meant to be a day 5 storage system, should the

6   platelets be reinfused on day 5, or should they be

7   reinfused on the day after the end of the storage

8   period, day 6 for a day 5 platelet, providing

9   assurance that the platelets were indeed able to be

10   recovered and able to survive all the way to

11   midnight on the last day of storage.

12             The group felt that it was better to

13   perform these studies on the last day of storage

14   for comparative purposes.  This is the way that

15   these studies have always been done in the past.

16   They noted that no dramatic changes occur with

17   platelet storage over time, the platelets don't all

18   keel over and die at midnight on the last day of

19   storage.   It's a gradual change over time.

20             [Slide.]

21             There are many technical details of

22   platelet radiolabeling that were begun to be

260

1   considered during this conference, but the

2   conference only lasted one day.  We couldn't

3   possibly hope to deal with all of them.

4             [Slide.]

5             As a result, issues have continued for

6   discussion, and the forum for this has been the

7   Biomedical Excellence for Safer Transfusion, or

8   BEST, collaborative.  The BEST collaborative has

9   essentially convened a small ad hoc group of

10   individuals with expertise and interest in platelet

11   radiolabeling that have traded documents back and

12   forth.

13             They met right after the May conference of

14   the FDA, and the BEST group recently had an

15   international meeting, as well, where some of these

16   points were considered.

17             The first thing that BEST was able to do

18   was to pull together thoughts and submit them to

19   the Agency regarding a new platelet guidance

20   document.  This submission by BEST included

21   consideration of not only in vivo issues, but also

22   in vitro testing parameters that could be applied

261

1   prior to a new technique being taken to in vivo

2   testing, and there were comments in there about the

3   issues that were just raised a few minutes ago

4   about the functional characteristic of platelets

5   and how these should be evaluated in the in vitro

6   testing before moving to in vivo analysis.

7             Currently, BEST is working on a standard

8   radiolabeling protocol.  I appreciate the

9   leadership of Ed Snyder on this and a number of

10   others through BEST who have been cooperating,

11   sharing ideas, sharing data, sharing protocols.

12             The approach that we are using is the one

13   that Drs. Heaton and Holme developed over 15 years

14   ago and has stood the test of time.  There are a

15   few minor changes which are occurring in this, but

16   we hope soon to be able to have a standardized

17   technique that all will be able to use.

18             [Slide.]

19             The fresh platelet approach that Dr.

20   Snyder will be talking about shortly has been

21   adapted from previous experience.  The one that we

22   have been using, for example, uses 43 ml of blood

262

1   anticoagulated with 7 ml of ACD, followed by a

2   short holding time before centrifugation and

3   production of PRP platelets.

4             Again, this is a manual-based system,

5   doesn't require any particular manufacturer's

6   collection device in order to produce the fresh

7   platelets in a standard way.  We are attempting to

8   refine this technique and Dr. Snyder has some data

9   that shows, for example, that if one increases the

10   initial ACD amount from 7 ml to 9 ml, one can do

11   away with the holding time and speed up the entire

12   process.

13             After the PRP is made, it is acidified

14   with ACDA, any other cellular elements are spun

15   out, and then the supernatant PRP is spun again to

16   produce a platelet pellet that can be used for

17   labeling, as well as supernatant platelet-poor

18   plasma that is involved in the labeling procedure.

19             [Slide.]

20             The labeling technique itself, as I said,

21   is one that Drs. Holme and Heaton have worked with

22   for a number of years, and can be applied to stored

263

1   platelets, as well.  The stored platelets have an

2   aliquot taken that is acidified, and again cellular

3   elements other than platelets are spun out, the

4   pellet is acidified, brought back into solution,

5   and then used for radiolabeling in the same way

6   that fresh platelets would be used, with either

7   sodium chromate or indium oxine.

8             This is a brief, 20-minute, room

9   temperature incubation that allows for the easy

10   production with relatively good uptake efficiency.

11             [Slide.]

12             So, the BEST group is working on a

13   standard protocol that we hope will be able to be

14   published in detailed from within the next year,

15   but certainly will be available and circulated

16   within the scientific community before then.

17             We hope this will provide a standard means

18   for all laboratories to perform this testing and

19   reduce another form of variability between

20   different clinical trials.

21             [Slide.]

22             A question of how to actually calculate

264

1   the result, as Dr. Leitman mentioned, is something

2   that bears further consideration, because

3   initially, as proposed by Dr. Murphy, he was

4   anticipating that one would look at fresh

5   platelets, one would look at stored platelets, and

6   simply divide the result of the stored platelets by

7   the fresh platelets, ending up with a ratio.

8             Larry Dumont spoke at the May meeting and

9   noted that statisticians do not like ratios because

10   they are difficult to work with, and instead,

11   another approach would be better, one in which one

12   could document the lack of inferiority of the

13   proposed system as opposed to the standard.

14             To do this, rather than just simply taking

15   the mean across the fresh and the mean across the

16   standard, one would calculate the difference for

17   each subject involved in the study between the test

18   infusion and the control infusion, determine the

19   mean, and then the upper confidence internal for

20   that difference.

21             In this case, from our first validation of

22   Dr. Murphy's idea, the mean difference was 16.5

265

1   percent when looking at recovery, and the upper

2   confidence limit was 20.7 percent.

3             This would be compared to what the target

4   is.  The observed fresh recovery was 75 percent as

5   it turned out in that study.  The multiplier that

6   would be used would be two-thirds or 67 percent, or

7   66 percent depending on how you want to express

8   two-thirds, multiplying those two values together,

9   the target then actually became 50 percent.

10             The 50 percent subtracted from 74 percent,

11   to give the maximum acceptable difference between

12   the test and the stored of 24.7, and the upper

13   confidence interval limit was 20.7, in this case

14   allowing us using a more robust non-inferiority

15   approach to the statistics, to actually accept the

16   proposal.

17             The same sort of approach would be used in

18   the comparison of survivals, and we saw that was

19   also acceptable.

20             [Slide.]

21             We moved on to a 7-day validation of Dr.

22   Murphy's idea using the fresh collection of

266

1   platelets on day 7 and reinfused simultaneously

2   with the stored apheresis component.  You can see

3   the actual absolute recoveries here, and the simple

4   ratio approach would seem to match the requirement

5   of Dr. Murphy, but these have been verified as not

6   being inferior to the standard according to the

7   approach that Larry Dumont had outlined back in

8   May.

9             [Slide.]

10             Another issue is actually how the survival

11   is calculated.  Most laboratories use a program

12   that is called COST, developed by a South African

13   researcher probably 20 years ago.  It has been

14   published, but it had been handed about from

15   laboratory to laboratory, and has required some

16   modifications between laboratories in order to

17   allow it to run on the current PCs and Window

18   environment.

19             He created a sample set of data and sent

20   them around to a number of different laboratories,

21   asked them to determine the recovery and survival

22   using their version of the COST program.  In

267

1   addition, one participant also created a similar

2   program in a standard statistical package, SAS

3   software, in order to see whether they would get

4   the same thing using that approach.

5             I am not going to go through all the data,

6   but the bottom line is that all of the laboratories

7   get exactly the same numbers, so that gave us the

8   confidence that this was really a program which was

9   turning out numbers that could be relied upon

10   across different laboratories.

11             [Slide.]

12             Another question was whether or not

13   chromium and indium labeling gave the same results

14   after extended storage.  Some data had recently

15   been presented that called that conclusion into

16   question, although that at five days, it appeared,

17   everyone agreed, that chromium gave the same result

18   as an indium label, that at longer time periods,

19   that may not be the case.

20             We attempted to address this, knowing that

21   it was important as we moved into an era of

22   hopefully, longer storage time for platelets.  We

268

1   collected an apheresis unit from a dozen

2   individuals, stored that for 8 days, and then took

3   two separate aliquots, labeled those aliquots with

4   either chromium or indium, reinfused them

5   simultaneously using now the soon to be

6   standardized method for platelet labeling in order

7   to see whether the two different radiolabels gave

8   the same result after the extended storage time.

9             [Slide.]

10             I am happy to report, as you can see here,

11   that there is no difference between either the

12   recovery or the survival after 8 days of storage of

13   apheresis platelets.  There was, in my opinion,

14   sufficient power in this study in order to pick up

15   what would have been a clinically important

16   difference.

17             [Slide.]

18             In conclusion, the meeting on May 3rd, the

19   expert panel's conclusion was that they accepted

20   the concept of Murphy's law.  This is the good

21   Murphy's law rather than the bad Murphy's law.  It

22   seems to be very widely supported internationally

269

1   based on discussions at the recent BEST meeting,

2   and I am happy to report that we are progressing

3   toward standardization and hopefully will have soon

4   a standard protocol that all laboratories will be

5   able to follow step by step.  I appreciate the help

6   of all those who are participating in that effort.

7             Thank you very much.

8             DR. ALLEN:  Thank you.

9             Comments or questions for Dr. AuBuchon?

10             DR. STRONG:  Could you comment on the

11   66/66 versus 66/50 differential here?

12             DR. AuBUCHON:  I agree with Scott Murphy.

13   In my experience, most patients who require

14   platelets for thrombocytopenia, require platelets

15   frequently, every 2 to 3 days.  Whether we supply

16   platelets that have a projected life span in

17   normals of, let's say, 4 days versus 8 days, will

18   make no difference to those patients, because of

19   their relatively low platelet count, the platelets

20   will be used up and won't ever have the opportunity

21   to reach senescence anyway.

22             The other group of patients who would use

270

1   platelets would be those undergoing surgery,

2   usually cardiovascular surgery or after trauma, for

3   example, and in those patients also, a prolonged

4   life span is not of clinical importance.  You need

5   to get them through the next hour or through the

6   next day, not through the next week, with the

7   platelets.

8             Therefore, I agree with Scott that if we

9   are able to show that the stored or treated

10   platelets have at least half of the survival of

11   fresh platelets, that will be more than adequate

12   for clinical use.

13             It appears that we are able to meet that

14   criterion by a wide margin with at least the

15   current apheresis technology.

16             DR. STRONG:  Was there consensus on that,

17   did your working group all agree with that?

18             DR. AuBUCHON:  Yes, there was, and

19   actually, when I left the meeting, I thought there

20   was also agreement with that consensus from the

21   FDA, but perhaps they have changed their mind since

22   then.

271

1             DR. DiMICHELE:  There was a question about

2   whether recovery or survival might relate to

3   function, platelet function, which is certainly

4   much more difficult to get to the bottom of.

5             Is there any potential that if we set our

6   sights too low on survival, we may actually be

7   looking at platelets that don't function as well?

8   It is not just how long they last in the

9   circulation, but really are they functioning, and

10   could survival, in particular lower survival, be a

11   surrogate for lower function or decreased platelet

12   function?

13             DR. AuBUCHON:  I agree with Salim's

14   statement, which is the general mantra in this

15   field, that a platelet that survives and circulates

16   is a platelet that functions.  Because of the

17   complex metabolic machinery that is involved, if a

18   platelet can function such that it can maintain its

19   shape and continue in circulation, it is likely to

20   be a functional platelet.

21             All the platelets that would reach the

22   stage of an in vivo study would have already gone

272

1   through in vitro analyses, which includes challenge

2   with various agonists in order to document that the

3   platelet can indeed respond to physiologic

4   stimulus, change its shape, release its contents,

5   and do what it is supposed to do.

6             That is a standard part of in vitro

7   testing and one that the BEST group certainly

8   supported continuation of using in future studies.

9   I don't have a concern that a platelet product that

10   would pass through all of the different studies,

11   including the in vivo radiolabeling study,

12   successfully, would not be functioning properly for

13   a patient to achieve hemostasis.

14             DR. DiMICHELE:  So, basically, what you

15   are saying is that there would be the in vitro

16   testing and this in vivo radiolabeling, but both

17   have to be done.

18             DR. AuBUCHON:  As Salim showed the

19   pyramid, the first step is always the in vitro

20   testing before any subject is exposed to

21   radioactivity.

22             DR. ALLEN:  Dr. Epstein.

273

1             DR. EPSTEIN:  Thank you, Jim, for that

2   very clear and comprehensive summary.

3             On the question of 66 percent versus 50

4   percent, my understanding is that all of the

5   currently licensed products can meet the 66 percent

6   standard for survival, so the question is why

7   should we lower what is currently an achievable

8   state of the art.

9             I think, as Salim pointed out, if we have

10   novel processes that produce less survivable

11   platelets, we could then have a reasonable debate

12   over whether the benefits outweigh the risks.

13             One of our goals in setting standards is

14   to try to identify the state of the art, and the

15   state of the art is that the platelets we make can

16   have 66 percent survival compared to fresh

17   preparation.  So why lower it?

18             DR. AuBUCHON:  I think this continues in

19   time-honored tradition in blood banking, when, as I

20   understand the original requirement for red cell

21   recovery at 24 hours, which was set at 70 percent,

22   actually the discussion occurred in a pub, so often

274

1   we are forced to make decisions without adequate

2   data to know exactly how we should set this up.  It

3   is a point that can be argued, I understand that.

4             DR. LEITMAN:  Jim, were you here this

5   morning for the discussion on red cells?

6             DR. AuBUCHON:  No, United Air Lines still

7   had me flying around.

8             DR. LEITMAN:  We were very concerned then

9   about the range around the mean.  It is not

10   tremendously relevant maybe to this discussion, but

11   we were concerned about outliers who have low

12   recoveries.

13             You gave some very nice mean data, could

14   you tell us about the interdonor variability in

15   platelet recovery?

16             DR. AuBUCHON:  There certainly is

17   intersubject variability, we see that all the time.

18   There would be some attempt to address some of the

19   major features, major sources of that variability

20   in the proposed protocol because both the fresh and

21   the stored platelets would be infused

22   simultaneously, so there would be no day-to-day

275

1   variability on the status of the subject.

2             One important point, and one point of

3   apparent variability relates to the estimation of

4   the subject's blood volume, which in platelet

5   studies is done based on a height and weight

6   formula rather than using a separate means of

7   actually determining their blood volume, and the

8   same blood volume would be used for both studies,

9   so you eliminate that variable, as well.

10             It certainly is true that some subject's

11   red cells don't store well, possibly through some

12   altered metabolic pathway that they may have, and

13   that may be true for platelets, as well, it has

14   never been thoroughly studied.

15             That would come up in this approach as

16   having a test platelet that did not survive well.

17   Whether that was due to the subject and his or her

18   own genetics, or the test article would not be

19   known, and one would have to then enroll that

20   subject in another study using a licensed approach

21   to see whether or not the platelets did not do

22   well.

276

1             I think the usual approach would be to

2   just have a large enough number of subjects,

3   usually at least 24 to 30, involved in a study, so

4   that a single donor that didn't do well, the

5   platelets didn't do well, would not cause the whole

6   study to fail.

7             DR. KLEIN:  Jim, since your name was taken

8   in vain several times this morning, and you weren't

9   here, I can't resist the temptation to ask you

10   whether this strategy for platelets might also be

11   applied to red cells.

12             DR. AuBUCHON:  It could be, however, we

13   have years, decades of experience with stored red

14   cells without trying to compare them to fresh red

15   cells.  In addition, we would have a problem in

16   that we do not have another radiolabel that we

17   would be able to use over a long enough time period

18   to assess fresh red cell recovery and survival in

19   comparison to stored.

20             Indium has been tried as a label of red

21   cells, and unfortunately, it elutes at too great a

22   rate.  Technetium labels red cells very nicely, but

277

1   decays at too rapid a rate, and also has some

2   elution problems over time.

3             So, I think that we are stuck with only

4   looking at stored red cells.  The issue is should

5   then the stored red cells in a new system be

6   compared to a standard system, in other words, have

7   a control arm.  That is something that we generally

8   like to do in studies in our laboratory, because it

9   allows us to see whether a subject has some

10   biochemical abnormalities in their red cells that

11   just prevent their red cells from storing properly

12   in any system, let alone a test system.

13             But that part of the study really isn't

14   helpful unless you need to explain an outlier, a

15   particularly poor recovery in a test system red

16   cell unit.  So, it is not exactly an analogous

17   situation.

18             DR. KLEIN:  Mr. Chairman, with your

19   forbearance, since the 24-hour survival of the red

20   cells is thought to be equivalent to looking at

21   full survival, couldn't you simply use chromium and

22   technetium, for example, look at 24-hour without

278

1   doing a full survival?

2             DR. AuBUCHON:  No, there is really too

3   much elution of technetium even in 24 hours to be

4   able to do that.

5             DR. ALLEN:  I have got a couple of

6   questions also. You mentioned that for the control

7   collection, that there would be a separate

8   collection using a manual method, which you then

9   went on to describe.

10             Has that been compared with a standard

11   platelet collection or an apheresis collection to

12   look at any differences?

13             DR. AuBUCHON:  Through several of our

14   studies in the last year looking at Murphy's law,

15   it does appear that at least some apheresis devices

16   are less injurious to platelets than this manual

17   PRP method, and that they do not cause pelleting of

18   the platelet at any time during the collection

19   process.

20             In our hands anyway, some of these

21   apheresis products can give 70 or even slightly

22   above 70 percent recovery as a fresh platelet.  So,

279

1   therefore, one might say, well, we should use the

2   best standard available.  However, if we were to

3   use an apheresis device today, that apheresis

4   device may not be available tomorrow.

5             Fresh platelets handled in this manual

6   method seem to have a slightly lower recovery, in

7   our hands, and you will see Dr. Snyder's data,

8   somewhere between 60 and 65 percent, so not hugely

9   different, but a little bit lower, and I guess that

10   is not surprising when you look at the two

11   different ways that the platelets are separated

12   out.

13             The expert panel was not truly concerned

14   about that and they felt that it was still better

15   to use a standard manual method that could be

16   reproduced over time, and just accept that the bar

17   is being set a little bit lower.

18             DR. KUEHNERT:  I just had a question about

19   the 66 percent number again, and I realize that is

20   sort of a soft number, but I am just wondering how

21   much variability there may be between donors.  I

22   mean does it vary with racial ethnic groups, how

280

1   much variability are we talking about here?

2             I see in Scott Murphy's paper, 62 percent

3   range, 50 to 80 percent, so it sounds like there is

4   a pretty wide range.  I wonder if you could give

5   some more background on that.

6             DR. AuBUCHON:  There certainly can be.

7   There certainly can be.  I would not be surprised

8   if a mean were, let's say, 60 percent recovery, to

9   see included in that group of data some individuals

10   who had 45 percent recovery and some individuals

11   who had close to 90 percent recovery.

12             That is why it is important to then use

13   the statistical approach that Larry Dumont

14   suggested, so that we are able to say with

15   statistical confidence that the test system that is

16   being evaluated is not inferior to the proposal

17   that Scott Murphy has made.

18             One can be certain that the variation that

19   is seen between donors, for example, does not

20   obliterate at true difference that really exists

21   between the test system and the standard that has

22   been proposed.

281

1             DR. ALLEN:  I had a second question.  You

2   mentioned the discussion about whether to infuse

3   the test platelets on the last day or the day after

4   the last day to assure that you had gone through

5   that full time period.

6             Would there be any advantage, given the

7   decision to infuse them on the last day, to look at

8   the storage time just for the test, not obviously

9   in an actual working situation, but for the test

10   period, to consider the storage time in hours, and

11   therefore, you could adjust for whatever collection

12   time was on day 1 and infusion time on day 5 or day

13   7, and look at it in hours versus just in whole

14   days?

15             DR. AuBUCHON:  I suppose that could be

16   done.  The laboratories that I am familiar with,

17   who do this approach, generally collect first thing

18   in the morning on the day that they are going to

19   collect, because that is the order of the day, and

20   then they usually begin labeling early in the

21   morning or relatively early in the morning, on the

22   last day of storage because it takes some time to

282

1   do all the preparations and get the post-infusion

2   sampling done, and the donor and the staff don't

3   want to stay around too late.

4             One could express this difference in time

5   from start to finish, as it were, in hours rather

6   than days, but again the common practice and what

7   we have experience in is doing just exactly that

8   system and knowing that those platelets work all

9   the way until the end of the day.

10             I think that one example of documentation

11   of the fact that these platelets don't all

12   immediately die at the last moment is some work

13   that we have presented previously, I believe at

14   this meeting, with transfusing platelets on day 6

15   and day 7 when we had no other platelets to

16   transfuse to thrombocytopenic patients, and those

17   platelets worked quite well.

18             So, I am not really concerned about the

19   day of the day after the last day of storage.  I

20   think that for comparative sakes, we are better off

21   just leaving it the way that we have been doing it.

22             DR. ALLEN:  Other questions or comments

283

1   for Dr. AuBuchon?

2             [No response.]

3             DR. ALLEN:  Thank you very much.

4             Our third presentation on this topic is

5   Dr. Edward Snyder, Department of Laboratory

6   Medicine, Yale, New Haven Hospital.

7                Presentation - Edward Snyder, M.D.

8             DR. SNYDER:  Thank you very much.  I

9   appreciate the opportunity to address the

10   committee.

11             Just a couple of things before I start.  I

12   think it is important to remember that platelets

13   really are not red cells.  You can freeze red cells

14   quite adequately, but platelets defy freezing under

15   most conditions.

16             There are a fair number of differences

17   between them and trying to treat them the same, I

18   think causes some problems.  The BEST committee and

19   the various experts, including the FDA, have worked

20   hard to come up with a protocol because when you

21   look at the details, which is where the devil

22   resides, if you were to use an apheresis product,

284

1   for example, and infused that on the last day,

2   there are volume considerations.  You can't take a

3   7-day product that you have stored and

4   plateletpherese the recipient, and then that

5   afternoon, reinfuse the pheresed product as a fresh

6   material plus the test product, because the volume

7   shifts would throw all of your calculations off.

8             So, we have thought a lot about exactly

9   how this should be done.  What I am going to

10   present to you is what I euphemistically call the

11   front end considerations, which is the protocol for

12   the storage of fresh, for the treatment of fresh

13   radiolabeled platelets.

14             [Slide.]

15             A conflict of interest statement.  Since

16   we do radiolabeling for so many different

17   companies, I thought it would be useful just to

18   push this up, so it could be on the record.

19             [Slide.]

20             The purpose of this study was really to

21   validate a dual platelet radiolabeling protocol

22   using both indium and chromium, whole blood

285

1   platelets en-tube, a French phrase there, and it

2   really wasn't a problem with indium.  The problem

3   was with chromium.

4             Prior to this time, almost all the

5   chromium had been labeled using a full bag of blood

6   that was collected, a protocol that Slichter and

7   others had developed and had been for years.  If we

8   now had to concern ourselves with volumes, we had

9   to shift the collecting a small amount of volume,

10   which is why we can collect the product, label it,

11   and infuse it on the day of storage, the last day

12   of storage, because the presumption is that 43 ml

13   or 60 ml isn't going to change the blood volume,

14   whereas, doing a full platelet pheresis or drawing

15   500 ml of whole blood would affect the blood

16   volume.  Even if you reinfused everything at the

17   end, there would still be too many shifts, and

18   there was too much riding on the results for that

19   variable.

20             So, the question was really could chromium

21   be labeled in a tube in a small volume.  We also

22   wanted to determine if the protocol was robust,

286

1   because we realized that there are multiple labs

2   doing this, and there needs to be some little

3   wiggle room, and I will go through a lot of the

4   variables, and If we could use this technique to

5   validate Murphy's law.

6             This whole thing was predicated on the

7   ability to label fresh platelets, whole blood

8   platelets, and inject them using chromium or

9   indium, because if you only used indium, and you

10   had never used chromium for fresh, then, you only

11   used the other radionuclide for the test product,

12   and there would be bias potential, so that wasn't

13   acceptable.

14             There were three test sites that validated

15   the front end, our lab, Red Cross at Norfolk with

16   Pam Whitley, and Dartmouth, Dr. AuBuchon.  This was

17   all based on the best protocol from 2004, which was

18   a work-in-progress, so the data I am giving you is

19   relatively small numbers.  I think there is a total

20   of 25, an N of 25, because everything was sort of

21   moving, it was a moving target, we needed to get

22   data for this presentation, as well as to see as we

287

1   went along if we were moving in the right

2   direction.

3             The concept was to adopt universally, this

4   is a generic protocol that would be used by all

5   manufacturers and other groups wanting to do these

6   studies.

7             [Slide.]

8             So, the processing basically was IRB

9   approved and Radiation Safety Committee approved at

10   the various organizations, only volunteer donors

11   were used, normal volunteer donors.

12             The variables that we evaluated, and you

13   will see data on, was the volume of whole blood to

14   be processed, was there a critical concern, and

15   again the concern was the known low labeling

16   efficiency of chromium being very low, in the 15

17   percent range as opposed to indium, much higher

18   levels for indium being in the 80 percent, 80 to 85

19   percent level, was the concern that in the volume

20   of whole blood, you are going to get enough

21   platelets to label with chromium that you can

22   actually detect survival and recovery with the

288

1   chromium.

2             The quantity of platelets, blood, to be

3   collected and injected is inherent, again in number

4   one.  The labeling environment, could we do it in a

5   tube, was it too sensitive, was it too fickle that

6   you really had to do things exactly right or it

7   wasn't going to work, the radioactive dose for

8   labeling and injection, and this has to do with the

9   size of the sodium iodide crystal in the gamma

10   counter.  Some places use a 2-inch crystal, some

11   use a 3-inch crystal.  The efficiencies are much

12   lower for the 2-inch crystal.

13             Couple that with the low labeling

14   efficiency of chromium to begin with, and if you

15   have a donor who you pick that happens to have a

16   platelet count of, let's say, 150,000 per

17   microliter, which is okay but low, as opposed to

18   someone with 350,000, you might get too little

19   radioactivity to actually be able to do an

20   appropriate analysis.

21             Then, also the labeling efficiency in the

22   tube, and what about the radionuclide elution,

289

1   would that be a problem.

2             [Slide.]

3             Other variables included the survival

4   characteristics, as I mentioned, of chromium at the

5   labeling efficiency, which is what LE is.  The

6   techniques for elution, standard preparation, and

7   estimated blood volume varied among labs, and we

8   tried to standardize that.

9             The sample times, some labs tested, for

10   example, we used to test at 5 minutes, 30 minutes,

11   1 hour, 2 hours, 3 hours, other places would just

12   do 2 hours, some 3 hours, some didn't think you

13   should do anything for the first 24 hours, some

14   went out to 10 days, some went out for 7 days of

15   sampling.  These are all variables that needed to

16   be evaluated.

17             The settings of the window on the gamma

18   counter, for chromium and indium, you have to make

19   sure that you didn't have overlap, took that into

20   account.  The size of the crystal I mentioned, the

21   time of counting, and then the calculations and the

22   curve fitting that was modeled with the COST

290

1   program, some of the things that Dr. AuBuchon has

2   already mentioned.

3             [Slide.]

4             So, these are comparisons of variables and

5   let me walk you through this.  The variables are

6   listed in the left column.  This is Yale-indium.  I

7   couldn't put a hyphen in because it kept flipping

8   over when I did that, so Yale-indium,

9   Yale-chromium, Red Cross-indium, Red

10   Cross-chromium, and Dartmouth-Hitchcock-indium, and

11   Dr. AuBuchon's data, as you will see, he did not

12   look at chromium, so he didn't do paired studies,

13   he just did chromium.

14             This is all studies done on fresh blood

15   platelets, what I call the front end labeling.  So,

16   the whole blood, this is to show you that there was

17   variability in the results you are going to see

18   among labs, showing as the testimony of the

19   robustness of the protocol.

20             Again, these were done before everything

21   was completely finalized, so there is some wiggle

22   room, if you will.

291

1             43 ml is what we collected.  We collected

2   86 ml for chromium.  So, on a donor, we collected

3   about 130 ml from that individual.  We used twice

4   the amount of chromium, but it was divided into

5   two, 43 ml aliquots because we weren't sure how

6   much label we were going to get.  So, that is how

7   we chose to do it.

8             The Red Cross used 60 ml, and they also

9   collected two, 60 ml aliquots from the same donor,

10   so they collected 180 at one time.  Hitchcock used

11   43 ml for their indium study.

12             7 ml of ACD was added at Yale for both

13   indium and chromium, as well as for Dartmouth,

14   whereas, the Red Cross used 60 ml with 9 ml of ACD.

15   The ratios are roughly the same.  It's about 6.1

16   here and about 6.6.

17             The concept was if you use more ACD, you

18   acidify the platelets to a lower level, and you

19   therefore don't have to wait for a two-hour hold

20   period, which is of practical importance.

21             Because we used 7 ml, we waited 2 hours

22   before we did processing.  Red Cross waited less

292

1   than 30 minutes because they use 9 ml, and

2   Hitchcock used 2 hours for their 43 ml with the

3   indium.

4             The microcuries injected again.  We used

5   40 microcuries of indium and 25 of chromium, higher

6   than the others because we have a two-inch crystal,

7   which is what we have been using for years.  This

8   obviously went through Radiation Safety, was well

9   within the guidelines of the Nuclear Regulatory

10   Commission, not a problem, but still you would like

11   to use less microcuries of anything if you can.

12             15 microcuries was what was used for the

13   other two labs.  Again, they had three-inch

14   crystals, and they had good efficiencies of

15   counting at that level.

16             The other things that were of

17   consideration were how the standards were measured.

18   You could measure it by volume and pipetting an

19   aliquot into a dilution, or you could weigh it to

20   do it gravimetrically, so Vol. vs Grav. is

21   volumetrically versus gravimetrically.

22             The overwhelming consensus was gravimetric

293

1   should be used.  This was before we did ours, so we

2   used volume, so we used volume, but gravity was

3   what was used, weighing it was what was used by the

4   other two labs.

5             The crystal size, I mentioned, and then

6   the draw times, we sampled at 1, 2, 3 hours, and

7   then daily out to day 7, and then day 10.  You have

8   to make allowances for the weekend, and not come in

9   Saturday and Sunday, so there are some practical

10   considerations, but on a more scientific note,

11   measuring at day 10 with indium will allow you to

12   determine if any other blood elements have been

13   labeled with indium and are confounding the data

14   that you are trying to collect and assuming it is

15   only a platelet you are looking at.

16             You do spin out the red cells, but there

17   could be some other confounding indium in there.

18             The other two groups just measured at 3

19   hours and then daily out to 7 to 10 days, whereas,

20   Dr. AuBuchon only measured out to day 7.  He didn't

21   go out to day 10.  Not true, you went out to day

22   10?  Well, that is what I thought I heard you say.

294

1   I will recheck further.

2             [Slide.]

3             These are the results of our studies.  I

4   am just going to show the first three columns

5   because the other ones I have graphs of, so it is

6   easier to read.  I apologize for the small figures.

7             The average platelet count was 290,000 per

8   microliter.  This was an N of 5 for only the

9   studies at Yale.  An average count of 290,000,

10   which was a reasonably good number.  We started by

11   labeling with 200 microcuries of chromium.  That is

12   what was added into the vessel to do the labeling,

13   which was a very large amount.  It is also

14   expensive.

15             We started with 100 microcuries of indium.

16   We started with more chromium because of the

17   concern over the low labeling efficiency, and then

18   the standard deviations are down here.  The next

19   slide will give you the other information here.

20             [Slide.]

21             Labeling efficiency was determined using a

22   dose calibrator, which is fairly standard, and the

295

1   efficiency was defined as the activity of the

2   platelets divided by the activity of the platelets

3   plus the activity of the supernatant after labeling

4   x 100.

5             [Slide.]

6             What you see here with chromium, the

7   labeling efficiency was about 16 percent or so, but

8   with indium, the efficiency was about 80 percent.

9   The numbers are different from the five because we

10   did a few practice to see what the efficiency was

11   going to be.

12             So, this is what we were looking at.

13   There was clearly a lower efficiency with chromium,

14   but it could be labeled in a tube.  We weren't sure

15   if it was going to be efficient or not.

16             [Slide.]

17             The elution was another concern, the

18   elution, the average counts.  In the supernatant

19   minus the background divided by a constant,

20   multiplied by counts per minute on the standard

21   minus the background.

22             [Slide.]

296

1             What we found was that the elution for

2   chromium was about 5 percent, for indium it was

3   about between 1 and 2 percent, which were within

4   reasonable levels, so the assumption was that

5   labeling in a tube versus a bag was going to show

6   reasonable efficiency of labeling and reasonable

7   lack of elution, if you will.

8             [Slide.]

9             The amount injected, in ml, we injected

10   about 7 ml of platelets with the chromium, which

11   got us about 20 to 25 microcuries.  So, we injected

12   all the platelets we had, because we needed to get

13   a level somewhere, at least 25 to 30, so we could

14   measure.

15             With indium, injected a little less,

16   because with indium, what you do is you just draw

17   up, you have such a good labeling efficiency, you

18   just draw up what you need based on the amount

19   irradiation, not on the size of the platelet

20   injectate.  You don't inject as much as you can,

21   because you would obviously inject too much.

22             So, using 2 or 3 ml of radiolabeled

297

1   platelets, we were able to get in somewhere around

2   35 microcuries of indium.  So, it looked like we

3   had a fighting chance at least to inject enough

4   chromium.  We had hoped to get reasonable labeling.

5             [Slide.]

6             The total platelets injected, for the

7   chromium, was about 1.5 x 10

9 and about 1 x 109 for

8   indium, which meant that you were in the ballpark.

9   If someone had a 289,000 platelet count using the

10   technique we used with the various spin speed and

11   times expressed in the protocol, you would have

12   enough, between 1 to 1.5 billion cells to be able

13   to inject.

14             So, it looked like you didn't have to be

15   overly concerned that you weren't going to have

16   enough platelets assuming that there was a normal

17   donor.

18             [Slide.]

19             This is to show you the variability.  This

20   gets to Susan's question about variability.  So,

21   these are our five subjects, and what we have here

22   is day 7 recovery for chromium and indium, and day

298

1   survival, in hours, for chromium and indium.

2             Now, the day 7 versus day 10 means that we

3   weren't sure, if you sampled the day, your last

4   sample was taken at day 7 and put in the COST

5   program, would you get a different result than if

6   you went out to day 10 and put it in the COST

7   program.

8             It makes more sense if you look at it the

9   other way, that you don't do day 7, you just do out

10   to day 10.  If you don't get to day 7, are the

11   curves different, are the results different.  So,

12   that is what we were testing.

13             These were fresh platelets.  It wasn't a

14   day 7 store.  It was the last sample was taken on

15   day 7 after injection of the fresh material, or day

16   10.

17             So, what you see here is a fair amount of

18   variability.  I will get to the means in a moment.

19   But 51 percent, 68, 60, 52, 45 for chromium if you

20   looked at sampling on day 7.  On day 10, the

21   results were similar.  Here is a 46, here is a 65.

22   Unfortunately, the same variability that applies to

299

1   red cells, applies to platelets, as well.

2             The survivals were similarly variable, 192

3   hours versus 227 hours, and 216 hours, the

4   reference is 9 days, 24 x 9 for the hours, x the

5   number of days.  Basically, the results were quite

6   similar.  Again, between day 10 for chromium and

7   indium, and day 10 versus day 7, chromium and

8   indium, and I will show you those data in a second.

9             [Slide.]

10             So, the recovery for the chromium, the

11   last sample being on day 7, was about 58, 59

12   percent.  For indium, it was a little higher, 58,

13   59 percent, as well.  For day 10, the chromium and

14   the indium were again similar, and they were almost

15   the same.

16             [Slide.]

17             From our experience, this is survival for

18   a N of 5 in hours.  This black line is 9 days just

19   as a reference.  Again, if you sample your last

20   sample on day 7 with chromium versus the last

21   sample on day 10 for chromium, you got almost the

22   same results.

300

1             Similarly, with indium for day 7 to day

2   10, and indium versus chromium were similar, with

3   all the caveats that we use, had lower labeling

4   efficiency and injected lower amounts.

5             [Slide.]

6             So, the raw numbers, the numbers are, the

7   means were 55.8 chromium, 58.9 indium, with a

8   standard deviation of 9 to 13.  The day 10 results

9   meaning the sample, if you looked at a day 10

10   sample, was 56 to roughly 60 with the standard

11   deviation of 7.8 to 13.  It was quite similar,

12   55.8/56.2, 58.9/59.9, and the same thing with

13   survivals, they were quite similar, 221 versus 232,

14   221 versus 214, it was quite close, so there was no

15   major discrepancy, and it seemed reasonable that

16   this would work well.

17             [Slide.]

18             Tidewater's data here.  This is the 3-hour

19   recovery for indium and the indium survival in

20   hours.  This is their chromium.  They did the same

21   things we did only they had a few variations as I

22   mentioned.

301

1             The chromium recovery, with the first and

2   only sample on day 1 being a 3 hour, we used the 1,

3   2, and 3 hour.  They used only a 3 hour, and this

4   is the chromium survival.  You can see the

5   variability.

6             Here is a recovery of 30 percent.  Here is

7   a 99 percent, 178 hours versus 221 hours.  There is

8   a fair amount of variability again as we mentioned.

9   The point is that they got 64.3 for their indium,

10   we had gotten about 60; 60.9 for chromium, we were

11   about 59.  We had about 216 hours, they had 192,

12   and they had 190 for their chromium.

13             So, again, it appeared to be quite robust

14   despite the variability and despite the fact that

15   these are small N's, the numbers are reassuring.

16             [Slide.]

17             Dr. AuBuchon's data for indium using just

18   a 3-hour point of recovery, which I believe I got

19   at least that much right, 61 percent was their

20   recovery again based on only a 3-hour sample on the

21   first day.  We had a lot of discussions about how

22   critical is the first hour point, the second hour,

302

1   et cetera, so it is a meaningful consideration

2   here.

3             Then, his survival was 213 hours with an N

4   of 11, and he had sampled out, I thought it was day

5   7, it goes out to day 10.  So, again, robust among

6   three laboratories.

7             [Slide.]

8             So, the processing variables to review.

9   The volume of whole blood processed, 43 to 60.  It

10   looks like it doesn't much matter, but we are going

11   to settle on one volume.

12             Quantity of platelets to collect and

13   inject appeared to be acceptable for chromium,

14   which was important. We knew it would do well for

15   indium.

16             The labeling environment, it appears can

17   be adequately transferred into a tube, so we don't

18   need to use a full unit of blood, and you can draw

19   small volumes.

20             The radioactivity dose, I mentioned

21   appeared to be acceptable, the labeling efficiency

22   and the elutions appeared to not be of any concern.

303

1             [Slide.]

2             The recovery and survival, low labeling

3   efficiencies appeared to be adequate with chromium.

4   The techniques for standards and estimated blood

5   volume were similar among the labs and appeared to

6   be fine.

7             The sampling days didn't appear to be a

8   big concern, but we are going to perform one

9   protocol, which will probably be a 3-hour sample,

10   and then daily out to day 7, and then skipping, and

11   then going out to day 10 or 11 or 12 actually.

12             The window settings, crystal size, I think

13   we are going to require I believe a 3-inch crystal

14   because it is the least amount of radioactivity, so

15   if you are going to do this, you have got to get

16   into the business, you have got to buy the machine

17   and get a larger size crystal.  I should talk, we

18   now have a 3-inch crystal, so I can say that.

19             The cost calculations were done and showed

20   uniformity, as Dr. AuBuchon showed.

21             [Slide.]

22             In summary, en-tube fresh whole blood

304

1   radiolabeling, using chromium is feasible even with

2   a low labeling efficiency and even with the 2-inch

3   crystal.  The BEST protocol is robust and

4   reproducible, some local flexibility is acceptable

5   and does not appear to substantially affect

6   outcome.

7             We will recommend universal adoption of

8   the BEST protocol when it is finished, hopefully,

9   this month, and it appears therefore that this

10   front end analysis is amenable and acceptable to do

11   the Murphy inferiority creep evaluation, so that

12   the FDA would have a green light based on what we

13   have seen here as pilot data, but fairly

14   reasonable, to go ahead and effect a change in the

15   standards by which we determine if products are

16   appropriate for licensure.

17             [Slide.]

18             Just to acknowledge the various labs, the

19   various people, a special thanks not only to the

20   labs and the personnel who did all the work, but to

21   Dr. Scott Murphy, Larry Dumont, who did a superb

22   job, along with Andy Heaton, Stein Holme, and

305

1   Sheryl Slichter, with the calculations and all the

2   other aspects of this, and also the entire BEST

3   collaborative.

4             Thank you very much.

5             DR. ALLEN:  Thank you, Dr. Snyder.

6             Comments or questions for Dr. Snyder?

7             Dr. Strong.

8             DR. STRONG:  I will ask the same question

9   of you on the 66/50 issue.

10             DR. SNYDER:  I am allowed 3 seconds of

11   reflection?  In my heart of hearts, I would like to

12   see 50 percent, but I think 66 percent, I think you

13   would have to show--I think 66 percent is probably

14   a better way to go, so I probably would choose to

15   the higher standard based on the fact that the

16   purpose of the FDA is to protect the public health,

17   and I think the public health is better served by

18   adding that 16 percent.

19             If there were overriding considerations,

20   we would have to rethink, but I think as an initial

21   standard, I think 66 percent is probably doable.

22             DR. STRONG:  According to these

306

1   preliminary studies, none of these passed.

2             DR. SNYDER:  None of these were compared

3   to a test.  These were only evaluation of the

4   fresh.  I didn't do any of the comparisons that Dr.

5   AuBuchon did.  This is apples and gorillas here.

6   What you were looking at was just the numerator.  I

7   had the denominator, I didn't have the numerator.

8             The numbers were similar to what Jim

9   showed, and that is an interesting question about

10   the fact that fresh is less of a recovery than an

11   apheresis product, probably because of the spinning

12   that occurs when it is being collected and

13   processed, but that is the way things are.

14             DR. ALLEN:  Other comments or questions?

15             [No response.]

16             DR. ALLEN:  Thank you very much.

17             We will now move quickly to the open

18   public hearing.  I understand that we have got two

19   speakers on Topic II, Allene Carr-Greer from AABB

20   and Mike Fitzpatrick from ABC again.

21             Before we get into those, I have to read

22   my statement, and I apologize.

307

1             Open public hearing announcement for

2   general matters meetings.

3             Both the Food and Drug Administration and

4   the public believe in a transparent process for

5   information gathering and decisionmaking.  To

6   ensure such transparency at the open public hearing

7   session of the Advisory Committee meeting, FDA

8   believes that it is important to understand the

9   context of an individual's presentation.

10             For this reason, FDA encourages you, the

11   open public hearing speaker, at the beginning of

12   your written or oral statement to advise the

13   committee of any financial relationship that you

14   may have with any company or any group that is

15   likely to be impacted by the topic of this meeting.

16   For example, the financial information may include

17   a company's or a group's payment of your travel,

18   lodging, or other expenses in connection with your

19   attendance at the meeting.

20             Likewise, FDA encourages you at the

21   beginning of your statement to advise the committee

22   if you do not have any such financial

308

1   relationships.  If you choose not to address this

2   issue of financial relationships at the beginning

3   of your statement, it will not preclude you from

4   speaking.

5             Ms. Carr-Greer.

6                       Open Public Hearing

7             MS. CARR-GREER:  I am sorry I didn't catch

8   you before you did that.  In the interest of time,

9   we are leaving ours as a written statement, and I

10   think ABC did not have one.

11             DR. FITZPATRICK:  We didn't have a written

12   statement, and with the representation of our

13   membership on the BEST panel, we support their

14   recommendation.

15             DR. ALLEN:  Thank you.

16             Any other comments?  Please identify

17   yourself.

18             MR. DUMONT:  I am Larry Dumont and Gambro

19   paid for my way here today.  I wanted to make a

20   comment on the 50 percent question on survival.

21             As I presented at the workshop in May,

22   there is an analytical consideration to the 50

309

1   percent.  The reason for that is because

2   survival--we didn't go into the details--but

3   survival is actually an extrapolated number in the

4   analysis, and there is a lot more uncertainty

5   associated with that estimate than there is with

6   the recovery.

7             Since the criteria are that you meet both

8   the recovery and survival, if you keep with the 66

9   percent for both of those, then, to power the

10   studies appropriately, the sample size nearly

11   doubles because of the uncertainty associated with

12   the survival.

13             The clinicians, all these platelet guru

14   clinicians, tell me that the recovery clinically is

15   more important than the survival, so melding that

16   consideration with the analytical aspect, it seems

17   like there is too much of a burden of proof from a

18   testing standpoint if you match those both at 66

19   percent as oppose to 66 and 50, and that

20   presentation is available on the FDA web site.

21             Thanks.

22             DR. ALLEN:  Thank you.

310

1             Any other public comments?

2             [No response.]

3             DR. ALLEN:  The public comment period is

4   now closed.

5             We will move to the presentation by Dr.

6   Vostal on the FDA Current Thinking and the

7   Questions for the Committee.

8               FDA Current Thinking and Questions

9           for the Committee - Jaro Vostal, M.D., Ph.D.

10             DR. VOSTAL:  Thank you very much.

11             First of all, I would like to express my

12   appreciation to the members of the expert panel we

13   had at the workshop, to the members of the BEST

14   collaborative, and certainly to Drs. AuBuchon and

15   Snyder in terms of being able to help us formulate

16   this approach and do the initial validation.  We

17   really appreciate that.

18             In terms of using this evaluation, we

19   really like it because it is independent of any

20   collection device, it's a manual collection, and it

21   also prevents us or keeps us away from the quality

22   creep that was described earlier.

311

1             So, we are committed to accepting this

2   approach if the committee agrees with that.

3             [Slide.]

4             The first question is:  Do committee

5   members agree that a new approach for platelet

6   evaluation comparing the recovery and survival of

7   stored test platelets to a reference preparation of

8   freshly drawn platelets from the same healthy

9   individual would represent a scientific

10   advancement?

11             DR. ALLEN:  Why don't you read all of them

12   and then we will come back and discuss.

13             DR. VOSTAL:  Okay.

14             [Slide.]

15             The second question is:  If FDA were to

16   adopt a standard as described in Question No. 1,

17   would the following criteria be appropriate:

18             a.  Recovery of greater than 66 percent.

19             b.  Survival of greater than 66 percent.

20             This represents the recovery and survival

21   of radiolabeled platelets.

22             Dr. Smallwood informed me that the

312

1   Question No. 1 is worded is slightly differently

2   than what was originally posted on our web site.

3             DR. ALLEN:  Thank you.

4             Committee Discussion and Recommendations

5             DR. ALLEN:  Can we go back to Question No.

6   1, please.  This question is now open for

7   discussion by the committee.

8             Dr. Goldsmith.

9             DR. GOLDSMITH:  Do we have a good

10   definition of fresh platelets in terms of handling

11   and reproducibility lab to lab, and place by place?

12   I am just concerned about the introduction of

13   variability there.

14             DR. ALLEN:  Was that defined by the BEST

15   collaboration?

16             DR. VOSTAL:  The protocol for fresh

17   platelet preparation was put together by a BEST

18   committee and actually by Dr. Snyder.

19             DR. SNYDER:  Yes.  The platelets are drawn

20   and immediately put into a container, and then ACD

21   is added, so the fresh is minutes old, it is not

22   stored for a period of time and you come back hours

313

1   later.  The processing begins immediately, and that

2   would be standard.

3             DR. ALLEN:  Is that satisfactory?  Okay.

4             Any other questions or discussion?  Are we

5   ready to vote?  Okay.  Dr. Smallwood.

6             DR. SMALLWOOD:  Question No. 1.  Do

7   committee members agree that a new approach for

8   platelet evaluation comparing the recovery and

9   survival of stored test platelets to a reference

10   preparation of freshly drawn platelets from the

11   same healthy individual would represent a

12   scientific advancement?

13             Dr. Allen.

14             DR. ALLEN:  Yes.

15             DR. SMALLWOOD:  Dr. Davis.

16             DR. DAVIS:  Yes.

17             DR. SMALLWOOD:  Dr. DiMichele.

18             DR. DiMICHELE:  Definitely yes.

19             DR. SMALLWOOD:  Dr. Doppelt.

20             DR. DOPPELT:  Yes.

21             DR. SMALLWOOD:  Dr. Goldsmith.

22             DR. GOLDSMITH:  Yes.

314

1             DR. SMALLWOOD:  Dr. Klein.

2             DR. KLEIN:  Yes.

3             DR. SMALLWOOD:  Dr. Laal.

4             DR. LAAL:  Yes.

5             DR. SMALLWOOD:  Dr. Harvath.

6             DR. HARVATH:  Yes.

7             DR. SMALLWOOD:  Dr. Kuehnert.

8             DR. KUEHNERT:  Yes.

9             DR. SMALLWOOD:  Dr. Leitman.

10             DR. LEITMAN:  Yes.

11             DR. SMALLWOOD:  Dr. Quirolo.

12             DR. QUIROLO:  Yes.

13             DR. SMALLWOOD:  Dr. Schreiber.

14             DR. SCHREIBER:  Yes.

15             DR. SMALLWOOD:  Dr. Whittaker.

16             DR. WHITTAKER:  Yes.

17             DR. SMALLWOOD:  Ms. Knowles.

18             MS. KNOWLES:  Yes.

19             DR. SMALLWOOD:  Dr. Strong, if you could

20   vote.

21             DR. STRONG:  That would be great.  Yes.

22             DR. ALLEN:  The results of voting.  It is

315

1   a unanimous Yes for Question 1.  The non-voting

2   industry representative agrees with the Yes vote.

3             DR. ALLEN:  Thank you.

4             We will go on to consideration of Question

5   2.  Dr. Smallwood, in terms of protocol here, since

6   we have got an (a) and a (b), do you want a single

7   vote on both (a) and (b), or do you want us to take

8   a vote separately on those two?

9             DR. SMALLWOOD:  I would defer to the staff

10   in terms of how you want this question interpreted.

11   Two votes.

12             DR. ALLEN:  Why don't we have discussion

13   on the whole question and then we will take a vote

14   separately on (a) and (b), and certainly based on

15   the discussion we heard, if somebody would like to

16   propose a modification for 50 percent, that is

17   obviously something that could be discussed and

18   might be substituted or whatever.

19             Let's wait and see how the discussion goes

20   on that point.  The question is open for

21   discussion.

22             DR. KUEHNERT:  Maybe I will just be the

316

1   first and just go back to the basics.  What are the

2   data again for picking this number?  Let's start

3   with recovery.

4             DR. VOSTAL:  When this was initially

5   proposed, I think there was relatively little data.

6   Dr. Murphy more or less picked that out from his

7   experience with doing platelet survival studies for

8   decades.

9             Looking at the data that was collected by

10   Dr. Snyder and Dr. AuBuchon, I think it was a

11   pretty good estimation of what that number should

12   be, but initially, there was very little data.

13             DR. ALLEN:  I found his review article in

14   Transfusion to be very helpful.  I mean obviously,

15   not being an expert in this area at all, and his

16   discussion of the splenic sequestration, the

17   studies that suggest splenic sequestration of a

18   certain percentage of normal healthy platelets,

19   which explains why your recovery is less than 100

20   percent, I found reasonably convincing.

21             It was on that basis that he then chose

22   the semi-arbitrary point of 66 percent.

317

1             DR. KUEHNERT:  But there is actually two

2   numbers. There is what gets sequestered and

3   recovered is 66 percent of the fresh, but then what

4   this is talking about is the ratio of the test

5   product to fresh, right?  Before, I asked about the

6   first point, but now I am asking about this

7   recovery ratio, and that is where it looks like

8   there are data, but the data that we have are Dr.

9   Snyder and Dr. AuBuchon's, that is the world of

10   data that we are dealing with, is that right?

11             DR. VOSTAL:  In terms of looking at the

12   ratio, yes.

13             DR. SCHREIBER:  In the survival, will it

14   be a set time period that you have to subscribe to?

15   Do you just follow it all the way down or do you

16   extrapolate?

17             DR. VOSTAL:  I think you collect data

18   until there is no more signal, and then you

19   extrapolate to the end of the curve, so you can get

20   your full survival time.  That is done by the

21   software.

22             DR. ALLEN:  Dr. Strong.

318

1             DR. STRONG:  Just a general comment.  I

2   think in the world of science and medicine that to

3   get all of the experts in a given field to agree on

4   anything is somewhat of an accomplishment, so I

5   would hate to ignore the consensus agreement of the

6   experts at the conference that came up with 66/50.

7             Actually, I think the last time we talked

8   about this, I asked you the question, if the

9   consensus was to change it to 66/50, would you

10   agree, and at least at that time, you did.  Just a

11   comment.

12             DR. VOSTAL:  I think that we did reach an

13   agreement at the workshop, but it was a more or

14   less soft agreement.  As you saw here, even Jim

15   admitted that you can argue the other side of the

16   coin.

17             From our point of view, you know, Jim

18   argues that it should be 50 percent because in

19   thrombocytopenic patients, you get a decreased

20   survival, but you have got to keep in mind that you

21   do all these studies in healthy volunteers, healthy

22   donors, and so we will be conceding quality right

319

1   upfront in healthy donors even without going to

2   patients.

3             As Dr. Epstein pointed out, we have all

4   the current platelet products meet the standard,

5   and there is really no reason to lower the standard

6   at the moment.

7             DR. STRONG:  I will voice my prejudice

8   here, because I happened to work with Dr. Slichter,

9   who spent her entire life doing this sort of thing,

10   and I think most of the experts agree that recovery

11   is more important that survival.

12             Also, the comment that Larry made, that it

13   will significantly complicate the studies to equate

14   those since you really are doing extrapolation on a

15   survival curve, therefore, you have to have a much

16   larger number to get your statistics.

17             DR. ALLEN:  Certainly, we could consider

18   adding a footnote, if you will, based on the

19   comment from Dr. Harvath earlier in the afternoon

20   that if you are doing studies that aren't in

21   healthy people, that perhaps you could consider

22   using a lower survival number, something like that.

320

1   Liana, I don't want to put words in your mouth.

2             DR. HARVATH:  Unfortunately, I wasn't at

3   that workshop.  I heard about it and am familiar

4   with some of the data, but my concern is in moving

5   away from what the recommendation was from the

6   experts, as you have pointed out, in the absence of

7   data to argue for deviating from the consensus of

8   the experts, I feel that I personally cannot

9   support the 66 percent for that particular point.

10             I don't have the data in front of me to

11   convince me to do otherwise.  The reason I asked

12   that question is sometimes a sponsor may come to

13   the FDA to perform a study, and it is possible that

14   the main outcome of that study will be performed in

15   people who are thrombocytopenic, and that t the

16   majority of these products will be used in people

17   who are thrombocytopenic or who are massively

18   hemorrhaging because of some other problem.

19             So, I have difficulty supporting the

20   change from the recommendation of the expert panel

21   at 50 percent to 66, but I would like to see the

22   data to support that difference.

321

1             DR. EPSTEIN:  I think that Dr. Haddad

2   outlined a hierarchy of approaches to validating

3   novel platelet products based on risk, and what was

4   being said is that for minimally modified

5   platelets, we generally accept the recovery and

6   survival as a surrogate for thrombocytopenic

7   efficacy, and that we move to trials in bleeding,

8   not just thrombocytopenic, but there are sort of

9   two grades there, too, stable thrombocytopenic and

10   bleeding thrombocytopenic patients.

11             In those studies, the endpoint is actually

12   not recovery or survival, it is hemostatic

13   efficacy.  So, the general paradigm here is that we

14   don't let the companies go into studies in

15   thrombocytopenic patients with an endpoint of

16   hemostatic efficacy if they haven't at least shown

17   recovery and survival in normal patients.

18             So, I think that part of the problem here

19   is that the issue is a little bit moot.  Could a

20   situation arise in which the studies can only be

21   done in the thrombocytopenic patients?  Suppose the

22   answer is yes.  Is that the normal course of

322

1   events?  No.

2             Normal course of events is that we are

3   first looking for recovery and survival in healthy

4   normals.

5             DR. SNYDER:  I think we are mixing things

6   here. The in vivo studies here are for radiolabeled

7   survivals in normal volunteers.  Corrected count

8   increments have really nothing to do with that 66

9   percent.  That was not a number that was put up as

10   you need a 66 percent corrected count increment.

11   We are talking about apples and oranges here.

12             You have to think hard to think of a study

13   where you would actually be able to do radiolabeled

14   survivals in someone who was thrombocytopenic.

15   First of all, you would have to give allogeneic

16   platelet because they don't have enough platelets

17   to label in the first place, plus they would be

18   sick.

19             We actually talked about that concept of

20   whether you could go into a leukemic patient and

21   infuse a radiolabeled platelet, and the ethics of

22   it, I mean it would be nothing for the benefit of

323

1   the patient at all.  It would be hard to get them

2   through an IRB today, I would think.

3             Just to keep the committee clear that the

4   thrombocytopenic patients really have nothing to do

5   with what the committee was evaluating here.  This

6   was only in normal volunteers, looking at a

7   radiolabel, which is Phase II, and then the third

8   step would be to go into patients with transfusions

9   if the product showed that it was adequate to go to

10   that level.

11             DR. ALLEN:  Dr. Leitman.

12             DR. LEITMAN:  I was part of that expert

13   panel, and I think what Jaro said is true, that

14   recommendation for survival of 50 percent of fresh

15   was soft.  We were focused on recovery, because we

16   have always been focused on recovery.  That is what

17   we think of for critical patient care.

18             I am influenced in my thinking today by

19   the data that I think Dr. AuBuchon showed for both

20   7-day storage and 10-day storage, that those

21   components met this criteria of 60 percent of

22   fresh.

324

1             Those are very, in traditional transfusion

2   medicine thinking, those are extreme circumstances.

3   We are all used to storing for 5 days.  So, if a

4   large N study, multi-center study shows that you

5   can accomplish that, I don't see a problem with

6   raising the bar.

7             Again, I am reflecting everything that has

8   been said.  These studies are done in healthy

9   subjects, and the highest bar should be in

10   autologous healthy subjects, and we are not

11   applying that to sick patients, but you wanted

12   higher in the perfect setting because it will go

13   lower in sick, bleeding people with low counts.

14             DR. ALLEN:  Dr. Heaton.

15             DR. HEATON:  Yes, I am Andrew Heaton of

16   Chiron, also the Heaton, the architect of the

17   method that you heard described.  I would make a

18   couple of key points, I think.

19             First, when you are putting stored

20   platelets into a human, the level of variability in

21   the recovery is relatively modest, but if you look

22   at the variability in survival, from time to time,

325

1   it is very much greater than the variability in the

2   recovery.

3             So, my suggestion would be that you go

4   with the 66 percent recovery, but the 50 percent

5   survival, not because you are lowering the

6   standard, but more to allow for the higher degree

7   of variability that you will observe in the studies

8   that will be performed to support licensure

9   approval.

10             DR. ALLEN:  Dr. Klein.

11             DR. KLEIN:  The issues as I see them are

12   that if you would move to a 66 percent survival,

13   you are trying to say that using this surrogate

14   will help you determine that, in fact, you have a

15   better quality platelet.

16             Now, that may or may not be true, we just

17   don't know.  By and large, in a practical sense, it

18   doesn't matter because the patients who are getting

19   these platelets are going to use them up long

20   before you are going to have to worry too much

21   about survival.

22             The other side of that is that in order

326

1   for people to do studies, they are going to have to

2   have a much large N to meet this relatively

3   subjective requirement, the 66 percent, and that

4   may actually be a hurdle that is unnecessary, so I

5   think perhaps what the committee has to determine

6   is because of the variability in survival data, it

7   is reasonable to go with a higher level, thinking

8   that you are giving a better platelet which you may

9   not be, but complicating perhaps the process of a

10   licensure for some of these products.

11             DR. ALLEN:  Other comments or discussion?

12             DR. DiMICHELE:  Based on the discussion so

13   far, I mean I think the question is do we need to

14   consider survival.  I mean why did the discussion

15   group consider survival at all if it's not so

16   important in assessing anything from clinical

17   importance to platelet function?

18             DR. KLEIN:  I think the answer for that is

19   that you would really like to have a bleeding

20   endpoint with a halting of either bleeding or

21   prevention of bleeding.  In the absence of that,

22   you do the best that you can, and probably the best

327

1   that you can is starting with in vitro testing and

2   then using both recovery and survival.

3             DR. DiMICHELE:  But why not recovery

4   alone?

5             DR. KLEIN:  Without taking anything away

6   means you have a little less in the way of your

7   certainty that, in fact, what you are giving is

8   hemostatically effective.

9             DR. DiMICHELE:  So, what you are saying is

10   despite the ambiguous relevance of the survival

11   data, I mean I guess we are being asked to pick one

12   arbitrary number versus another for a parameter

13   that nobody seems to care very much about, I mean

14   in terms of its clinical relevance or what it is

15   telling us about the product.

16             DR. KLEIN:  Well, they care a little bit

17   about it. The previously cited Sheryl Slichter used

18   a viability assay, which was a combination of

19   recovery and survival, and thank goodness, you

20   didn't propose that for us today.  It would be one

21   more variable to put into the mix.

22             DR. DiMICHELE:  I am just thinking about,

328

1   you know, when we used sort of survival data in

2   clotting factor therapy, I mean we use it for

3   various reasons.  I mean we use it to make sure

4   that no one has an occult antibody, we use it to

5   actually determine treatment schedule, but that is

6   not what it is being used for here, so I am just

7   asking the question.

8             DR. LEITMAN:  I think we didn't use this

9   survival data because it was all over the place in

10   older studies.  It was so variable from study to

11   study, and there wasn't a conformity of the way of

12   doing those studies, and there was more conformity

13   in recovery.  Maybe Andy--I think he just said

14   that, but the data seems better now, and if it is

15   there, then, it is useful.

16             DR. HEATON:  I think a key issue here is

17   the level of reproducibility of survival is much

18   better now than it used to be.  Secondly, there is

19   an importance in survival in that you can get

20   platelet products of clinical situations with very

21   adequate recovery, but very, very short survival,

22   which would compromise patient care.

329

1             So, survival has some value, it just has

2   more variability than does recovery.

3             DR. DiMICHELE:  It just makes me wonder,

4   though, whether that parameter, I mean obviously,

5   there is a product variability and then a host

6   variability, and the question is, is whether in

7   that case, the host variability exceeds the product

8   variability, you know, the variability and what

9   that variability is due to.

10             That has been my experience certainly in

11   using platelets clinically, is that it often has

12   more to do with what is going on in the patient

13   than potentially what the product is.

14             DR. HEATON:  The key issue is you need to

15   have the platelets there for them to be able to

16   function.  We did do a series of studies measuring

17   radioisotope functionality, as well as recovery,

18   and you get quite good correlation between

19   post-transfusion recovery and post-transfusion

20   function as documented by isotope retention, and

21   aggregation columns or adhesion columns.

22             DR. DiMICHELE:  I am only addressing my

330

1   comments to survival, not to recovery.

2             DR. ALLEN:  Andrew, one further question

3   to you, as the parent of the technique here.  Given

4   the argument, if you are doing your studies in

5   normal humans as opposed to people with a

6   pathological process or disease of some sort, might

7   you be able to accept 66 percent or a higher level

8   for the normal studies if you then said we will

9   accept a greater variability in people with a

10   disease process?

11             DR. HEATON:  In theory, yes, but if you

12   consider the practice in a normal individual, a

13   certain percentage of your platelets die every day

14   of old age, and a certain percentage are used to

15   support your hemostatic process.

16             In a patient, the ratio is heavily biased

17   towards usage and away from dying of old age.

18   Therefore, in a patient circumstance, recovery is

19   of much greater importance to you than survival

20   because so few of the platelets live long enough to

21   die of old age.

22             DR. QUIROLO:  Is it even conceivable that

331

1   the survival would be below 50 percent?  I mean

2   making this 50 percent, does that sort of make the

3   survival moot, that it really doesn't matter at

4   that point?

5             DR. HEATON:  No, you do occasionally see

6   reduced post-transfusion survivals.  It is not very

7   common, and usually it's associated with reductions

8   in post-transfusion recovery, but you do

9   occasionally see reductions in post-transfusion

10   survival.  In the early days with cold storage

11   platelets, you saw reduced survival, but relatively

12   normal recovery, and it was clearly a significant

13   improvement when we shifted to room temperature

14   storage platelets.

15             DR. ALLEN:  Dr. Haddam.

16             DR. HADDAM:  Just one quick word about

17   survival that you mentioned.  One patient

18   population that survival would be important in is

19   the pediatric population because maybe one

20   transfusion in a kid would allow that kid not to

21   have to come back a second time for a transfusion.

22             DR. LACHENBRUCH:  I just wanted to ask,

332

1   what is known about the correlation between

2   recovery and survival, because if they are

3   measuring almost the same thing or something pretty

4   close, then, perhaps only one of them is needed.

5             DR. KLEIN:  I think the problem is that

6   they may or may not be depending upon what the

7   process is, and I think a very good example is just

8   what Andy Heaton said, was cold stored platelets,

9   where clearly there is a major difference between

10   recovery and survival, and I think you can't

11   necessarily predict, if you are going to be using a

12   new process, that the one will necessarily

13   correlate well with the other.

14             That is why I would argue that they both

15   are necessary although one is probably more

16   important in terms of what we are talking about

17   today than the other.

18             DR. ALLEN:  Dr. Heaton, do you want to add

19   to that?

20             DR. HEATON:  Yes, the key issue, I would

21   agree with what Harvey said, under normal

22   physiologic circumstances, there is very good

333

1   correlation, but in the patient circumstances, you

2   don't always get that correlation, and under

3   certain specific storage conditions, you don't

4   always get that correlation.

5             One of the areas that is increasingly

6   happening is as the platelet pheresis developers

7   produce new machines, they are elutriating the

8   platelets, and therefore, preferentially selecting

9   a young population of platelets, and that will

10   cause differences in the survival to recovery

11   ratio, different to that which you would expect in

12   a normal representative sample of the patients'

13   circulating platelets.

14             DR. ALLEN:  Other comments or questions?

15             DR. LEITMAN:  Again, on this survival, we

16   have been focusing on the bleeding, critically ill

17   thrombocytopenic patient.  There is a large

18   proportion of outpatients with chronic

19   thrombocytopenia, and if you can reduce their need

20   to come to clinic from three times a week to two

21   times a week, or from two times a week to one time

22   a week, that is a great advantage to them and to

334

1   management of the clinic, so survival is important.

2             DR. DiMICHELE:  Just to clarify the aspect

3   of my question, it was really to what extent it

4   reflected the product and to what extent it would

5   reflect what is going on in the patient, so that is

6   what I was trying to gauge and whether it was

7   important or not.

8             DR. ALLEN:  Other comments or questions?

9             [No response.]

10             DR. ALLEN:  We will, first of all, take a

11   vote on 2a, the recovery question.

12             DR. KLEIN:  The way this is worded is

13   appropriate, it is not worded optimal.  One could

14   say that 66 percent is appropriate, but maybe 50

15   percent is better, more practical, however you want

16   to say it.

17             DR. ALLEN:  Do you want to propose a

18   wording change?

19             Any other comments with regard to Dr.

20   Klein's--

21             DR. VOSTAL:  Also, you can consider that

22   50 percent is not a magic number.  We could also

335

1   come up with something in between 50 and 66

2   percent.

3             DR. ALLEN:  A sliding scale, if you will.

4             DR. VOSTAL:  I would also like to add that

5   if variability and survival is an issue, maybe we

6   can deal with that through statistical approach and

7   give it variability, or some way of dealing with

8   that, so we don't have to necessarily lower the

9   mean, but we can extend the parameters around that

10   number.

11             DR. ALLEN:  I suspect we don't have enough

12   hard information today to try to tackle that issue.

13   I think it is the sort of thing that again, with

14   collection of data over time, one might conclude

15   that that is an appropriate decision, but I don't

16   think we are able to do that today based on what we

17   have.

18             Dr. Smallwood, let's call the question for

19   2a.

20             DR. SMALLWOOD:  For the record, Question

21   No. 2a.

22             If FDA were to adopt a standard as

336

1   described in Question 1, would the following

2   criteria be appropriate? Recovery greater than or

3   equal to 66 percent.

4             Dr. Allen.

5             DR. ALLEN:  Yes.

6             DR. SMALLWOOD:  Dr. Davis.

7             DR. DAVIS:  No.

8             DR. SMALLWOOD:  Dr. DiMichele.

9             DR. DiMICHELE:  Yes.

10             DR. SMALLWOOD:  Dr. Doppelt.

11             DR. DOPPELT:  Yes.

12             DR. SMALLWOOD:  Dr. Goldsmith.

13             DR. GOLDSMITH:  Yes.

14             DR. SMALLWOOD:  Dr. Klein.

15             DR. KLEIN:  Yes.

16             DR. SMALLWOOD:  Dr. Laal.

17             DR. LAAL:  Yes.

18             DR. SMALLWOOD:  Dr. Harvath.

19             DR. HARVATH:  Yes.

20             DR. SMALLWOOD:  Dr. Kuehnert.

21             DR. KUEHNERT:  Yes.

22             DR. SMALLWOOD:  Dr. Leitman.

337

1             DR. LEITMAN:  Yes.

2             DR. SMALLWOOD:  Dr. Quirolo.

3             DR. QUIROLO:  Yes.

4             DR. SMALLWOOD:  Dr. Schreiber.

5             DR. SCHREIBER:  Yes.

6             DR. SMALLWOOD:  Dr. Whittaker.

7             DR. WHITTAKER:  Yes.

8             DR. SMALLWOOD:  Ms. Knowles.

9             MS. KNOWLES:  Yes.

10             DR. SMALLWOOD:  Dr. Strong, your opinion.

11             DR. STRONG:  Yes.

12             DR. SMALLWOOD:  The results of voting are

13   13 Yes votes, 1 No vote, and the non-voting

14   industry representative agrees with the Yes vote.

15             DR. ALLEN:  Thank you.

16             Any further discussion on Question 2b?

17   What I would propose is that we vote on 2b as

18   presented by the FDA, and if that is not accepted,

19   we will entertain discussion about (a), we don't

20   want to vote on it at all, or (b) whether we want

21   to vote on it on a different level.

22             So, we will take the question as submitted

338

1   at 66 percent, first of all.

2             Dr. Smallwood.

3             DR. SMALLWOOD:  Question No. 2b.  If FDA

4   were to adopt a standard as described in Question

5   1, would the following criteria be appropriate?

6   Survival greater than or equal to 66 percent.

7             Dr. Allen.

8             DR. ALLEN:  No.

9             DR. SMALLWOOD:  Dr. Davis.

10             DR. DAVIS:  Yes.

11             DR. SMALLWOOD:  Dr. DiMichele.

12             DR. DiMICHELE:  Abstain on the basis of

13   undecidedness.

14             DR. SMALLWOOD:  Dr. Doppelt.

15             DR. DOPPELT:  Yes, but I am also a little

16   confused.

17             DR. KUEHNERT:  Could we hold on a second,

18   because I think you said if there is a majority no

19   vote, then, we are going to rediscuss it, is that

20   what you said?  I just wanted to clarify.

21             DR. ALLEN:  Yes, if the--

22             DR. KUEHNERT:  Or if we all abstain?

339

1             DR. ALLEN:  The question is would a

2   criterion of survival of 66 percent or greater be

3   appropriate, and if the answer is no, that is not

4   an appropriate level, we will continue with

5   discussion as to whether another level should be

6   suggested or whether we will just leave it as we

7   are not making any recommendation for this one.

8             So, a no would be to reject the 66 percent

9   level.

10             DR. SMALLWOOD:  Dr. Doppelt, your vote

11   remains the same?

12             DR. DOPPELT:  Yes, I am still with yes,

13   but we are using the word appropriate, so however

14   you want to define it.

15             DR. SMALLWOOD:  Dr. Goldsmith.

16             DR. GOLDSMITH:  Yes.

17             DR. SMALLWOOD:  Dr. Klein.

18             DR. KLEIN:  I think it is appropriate, I

19   don't think it's ideal.  I guess that is a yes.

20             DR. SMALLWOOD:  Dr. Laal.

21             DR. LAAL:  Abstain.

22             DR. SMALLWOOD:  Dr. Harvath.

340

1             DR. HARVATH:  Abstain.

2             DR. SMALLWOOD:  Dr. Kuehnert.

3             DR. KUEHNERT:  What if it's a majority

4   abstain?  I feel like I was decided, if one can be

5   decided on abstaining, I was going to decide on

6   that, because I think it needs to be worked out

7   further.

8             DR. ALLEN:  That is perfectly okay.  If

9   the majority of the committee votes abstain, I

10   think it is very clear that the committee is saying

11   that there is not enough information that has been

12   presented to make a decision on this one.

13             We are strictly advisory to the FDA.  The

14   FDA will take this information under consideration

15   and make their own determination as to whether they

16   want to go one way or another or to seek further

17   information on it.

18             DR. KUEHNERT:  Then, I will decidedly vote

19   to abstain if that is possible.

20             DR. SMALLWOOD:  Dr. Leitman.

21             DR. LEITMAN:  I vote yes, but I don't

22   really care that much.

341

1             DR. SMALLWOOD:  Dr. Quirolo.

2             DR. QUIROLO:  Yes.

3             DR. SMALLWOOD:  Dr. Schreiber.

4             DR. SCHREIBER:  Yes.

5             DR. SMALLWOOD:  Dr. Whittaker.

6             DR. WHITTAKER:  Abstain.

7             DR. SMALLWOOD:  Ms. Knowles.

8             MS. KNOWLES:  Abstain.

9             DR. SMALLWOOD:  Dr. Strong, your opinion.

10             DR. STRONG:  How about no, but I don't

11   really care that much.

12             DR. SMALLWOOD:  The results of voting are

13   as follows:  7 Yes votes, 1 No vote, 6 abstentions,

14   and the non-voting industry rep agrees with the No

15   vote.

16             DR. ALLEN:  Of all the votes cast, there

17   was not a majority, am I correct on that, or it

18   reached the 50 percent level?  As I interpret this,

19   the vote was split, the committee was split.  It

20   did not carry, but I think the FDA has a sense of

21   where the committee is coming from on this issue.

22             Thank you, all.  I think we are at our

342

1   break time. We have got a 15-minute break.  I would

2   like to ask that everybody be back here at--I am

3   going to say that it's 3:50, I got different

4   timepieces--I am going to say it's 3:50.  Please be

5   wandering back in at 4 o'clock and ready to start

6   at 4:04.

7             [Break.]

8             DR. ALLEN:  Topic III, Experience with

9   Monitoring of Bacterial Contamination of Platelets.

10             Introduction and background by Dr. Vostal,

11   FDA.

12           III. Experience with Monitoring of Bacterial

13                    Contamination of Platelets

14                   Introduction and Background

15                        Jaro Vostal, M.D.

16             DR. VOSTAL:  Thank you very much for the

17   opportunity to share with you where we are in terms

18   of bacterial detection and transfusion products,

19   particularly of platelet transfusion products.

20             [Slide.]

21             A lot of things have happened in terms of

22   bacterial contamination detection in the last three

343

1   or four years, and I would just like to point out

2   some of the major events that have happened.

3             In 2001, the FDA approved two devices, the

4   BacT/Alert and the Pall BDS device, and they were

5   cleared by the FDA for quality control of platelet

6   products.

7             In 2002, there was an open letter to the

8   blood collection community urging bacterial testing

9   of platelet products, and this was a letter from

10   Drs. Brecher, AuBuchon, Yomtovian, Ness, and

11   Blajchman.

12             In this letter, they pointed out the

13   urgency in terms of being able to detect bacteria

14   in contaminated platelet units, and they urged

15   everyone to start doing bacterial testing.

16             In the same year, in December, we

17   presented at a BPAC meeting, the outline of studies

18   that would be necessary to clear bacterial

19   detection devices for a release indication of

20   platelet products.

21             Then, if we fast forward to 2004, the AABB

22   came out with a new standard requiring bacterial

344

1   testing prior to release of platelet products.

2   This standard didn't actually specify the testing

3   method, and it was up to the individual or

4   individual centers to decide if they wanted to do

5   culture with a clear device or if they wanted to do

6   a dipstick method, looking for pH change or glucose

7   change.

8             Later that year, PHS, including FDA and

9   CDC, expressed concerns about a lack of testing

10   standardization and product shortages due to

11   potential excessive false positive rates.  It turns

12   out that a lot of these methods used by different

13   centers are not standardized, and the lack of the

14   standardization could lead to a number of false

15   positives and false positive rates and the discard

16   of those units.

17             There was a discussion between PHS and

18   AABB, and the outcome of those discussions was the

19   formation of a task force for bacterial

20   contamination to facilitate the discussion between

21   all these parties involved.

22             [Slide.]

345

1             This slide summarizes the concern that the

2   FDA has regarding validation of bacterial detection

3   systems.

4             The standard conditions for culturing of

5   platelets 24 hours post collection have not been

6   defined, and they have not been defined in terms of

7   the volumes taken, whether you require anaerobic

8   cultures, the time of sampling, and the duration of

9   the hold prior to release of the units.

10             In addition, sensitivity, specificity, and

11   the predictive value of bacterial detection

12   systems, including cultures and dipsticks, have not

13   been determined.

14             [Slide.]

15             There are additional PHS concerns that

16   have been voiced mainly by the CDC, and there are

17   standard approaches needed for confirmation of

18   positive test results, identification of the

19   organisms, speciation of the organisms, and then

20   aggregated reporting of these results, so that we

21   can have a clear picture of what the contamination

22   rates are across the country.

346

1             In addition, there has been discussion

2   about notification of donors when the results may

3   suggest endogenous bacteremia.  Also, there is a

4   need for estimation of residual risk of bacterial

5   contamination is desirable as a benchmark.

6             [Slide.]

7             So, these are the unmet needs for the

8   blood system.  We need better clarification of

9   labeling for units released after bacterial

10   detection.  We need validation of bacterial culture

11   on pre-storage pooled units.  We need data to

12   support extension of platelet storage out to 7

13   days.

14             [Slide.]

15             The FDA clearance of bacterial detection

16   devices is based on their intended use.  There are

17   two uses that we recognize.  One is for quality

18   control, quality control indication.  This is the

19   sampling of a small number of collected products to

20   assure collection process is in control, and this

21   could be as few as 4 per month.

22             The key point here is that the centers can

347

1   transfuse the product without waiting for the

2   results.  The quality control is really to look at

3   the collection process whether the collection is in

4   control, not necessarily whether those units are

5   fit to be released for transfusion.

6             A product release indication is a

7   screening of all products prior to release for

8   transfusion.  Here, the decision to transfuse

9   depends on the results of that test.

10             The products that have been tested by a

11   devices cleared for this can be labeled as "tested

12   negative for bacteria by" whatever method was used.

13             [Slide.]

14             So, in terms of QC labeling, quality

15   control labeling for BacT/Alert culture bottles,

16   now, this is what was approved in 2001, and the

17   labeling states, The BacT/Alert System and culture

18   bottles may be used for quality control testing of

19   platelets.  The laboratory should follow its own

20   quality control procedures for this use.

21             The BacT/Alert System, including the

22   culture bottles, should not be used in determining

348

1   suitability for release of platelets for

2   transfusion.  Users considering such release

3   testing should first consult CBER for the

4   appropriate clinical studies.

5             So, that is the labeling that is currently

6   on this device.

7             [Slide.]

8             Here is how this device was approved or

9   cleared for this indication.  For quality control,

10   we look at in vitro testing, and usually, the

11   device is tested on platelet products intentionally

12   contaminated with variable levels of bacteria, what

13   is common referred to as a spiking study.

14             The in vitro testing identifies the

15   analytical sensitivity for particular bacterial

16   species and the adequate sampling times, sampling

17   volumes, and hold periods to improve sensitivity of

18   that device.

19             Of course, there are limitations to the

20   information you can get from in vitro studies.

21   That is, the clinical sensitivity, specificity, and

22   predictive value are unknown.  Conditions of use

349

1   may very considerably.  There could be on-line

2   versus off-line testing, percentage of units

3   sampled, testing during shelf life versus at

4   outdate, use or non-use of anaerobic cultures.

5             [Slide.]

6             So, our current thinking on evaluation of

7   bacterial detection devices for release of platelet

8   products for transfusion is as follows.  This

9   criteria is more stringent because more stringent

10   validation criteria are needed than for a QC

11   testing claim because the intended use is to

12   assure, with defined confidence, that released

13   products are not contaminated.

14             That is, the bacteria, if present, will be

15   below a certain level, and the frequency of

16   contamination is predictably low.  Now, in addition

17   to the in vitro studies, clinical studies are

18   needed to establish the false negative rate and the

19   false positive rate for the device under actual use

20   conditions.

21             [Slide.]

22             So, the evaluation process goes initially

350

1   the same as for QC testing.  We do the in vitro

2   testing, but then we requested a field trial to

3   demonstrate the performance of the devices under

4   actual use conditions.

5             This field trial consists of sampling of

6   transfusion products from routine collections, and

7   for culture-based devices, the trial should

8   demonstrate that culture results of a sample taken

9   early in storage period are predictive of results

10   of a sample taken at the time of release of product

11   or at the end of storage period.

12             [Slide.]

13             Previously, we have presented this

14   approach to the BPAC committee back in December of

15   2002, and this was the initial design of these

16   studies.  We proposed that during a storage period

17   of up to 7 days, their initial sample should be

18   taken early in the storage period, somewhere around

19   day 1, and then a second sample should be taken at

20   day 5, followed by another sample at day 7.

21             The results of these samples will be

22   compared to see if they can confirm the results of

351

1   the initial culture. If this a contaminated unit,

2   the bacteria will proliferate during storage, so

3   these two samplings have a lot easier time of

4   detecting bacteria, so they serve as the standard

5   for the original culture.

6             Initially, the BPAC committee agreed with

7   this approach, however, none of the companies

8   involved in the production of these devices have

9   stepped forward and done these types of studies,

10   and the problem has been that because of the

11   bacterial contamination rate being low, the size of

12   the study has to be relatively large, somewhere

13   around 50,000 units tested.

14             Therefore, doing a study like this on that

15   many platelets has been a substantial hurdle to

16   getting these studies done.  So, we re-analyzed in

17   information we would require for clearing of these

18   devices, and we have sort of adapted the model or

19   the study model to fit those needs.

20             [Slide.]

21             We have a new proposed design for a field

22   trial of bacterial cultures.  We propose to use a

352

1   validated 7-day collection container, so the cells

2   will be stored out under conditions that would be

3   acceptable.  We would then perform a culture on day

4   1, and a second culture and a Gram stain on day 7.

5             So, we have eliminated the day 5 culture,

6   which would bring down the cost significantly.  The

7   day 1 culture will approximate the prevalent risk

8   of unscreened products, and currently, we think

9   that risk is about 1 to 2,000 to 4,000 of units

10   contaminated.

11             The day 7 culture and Gram stain will

12   define the performance of day 1 culture and measure

13   the residual risk of platelet products pre-screened

14   on day 1.

15             [Slide.]

16             The primary objective will be to

17   demonstrate that the 95 percent upper confidence

18   limit is less than 1 per 5,000 based on a point

19   estimate of the 7-day contamination rate of 1 to

20   10,000.

21             If risk of unscreened platelets

22   significantly exceeds risk of pre-screened 7-day

353

1   platelets, then, we were willing to approve the

2   extended dating for pre-screened platelets, and

3   label the culture device based on its performance.

4             So, what we are saying is we are going to

5   try to get an idea of what the residual risk is of

6   7-day platelets, and if that residual risk comes in

7   somewhere around this contamination rate, 1 in

8   10,000, that would be sufficient enough, compared

9   to unscreened platelet products, would be a

10   sufficient improvement over unscreened platelet

11   products for us to be able to allow approval of the

12   device for bacterial testing, as well as approval

13   of 7-day platelets.

14             [Slide.]

15             We have suggested some cost-saving options

16   for the field trial.  As I pointed out, we proposed

17   a single culture on day 7.  This obviates the need

18   for a previously suggested culture on day 5.

19             The day 1 culture is already being done

20   according to the AABB standard, so practically,

21   half the study is already paid for right there.

22   Expired products on day 7 of storage could be

354

1   shipped to a central testing location.  In this

2   way, you could save on the amount of personnel

3   needed to do the testing and also standardize the

4   procedures that are used to test these products.

5             The day 7 cultures could be tested using

6   pooled samples.  Previously, we proposed that each

7   sample should be sampled with a single bottle, but

8   as I pointed out, there is a significant growth of

9   bacteria during storage, and those platelet loads

10   by day 7 could be very high, up to 10

6 CFU/ml.

11   Therefore, since these devices are sensitive at 10

12   CFU/ml, you can have significant dilution and still

13   detect contaminated units.

14             So, you can dilute these 5-fold to

15   10-fold, and still have sufficient sensitivity to

16   confirm the initial culture.

17             This last point also talks about

18   collecting a small volume and pooling it into the

19   culture bottle to get sample pooling.

20             [Slide.]

21             As I pointed out earlier, this study will

22   still require a large number of platelet products,

355

1   and probably somewhere around 50,000.  This will

2   give us reasonable assurance that the residual risk

3   is less than 1 to 10,000 for pre-screened units.

4             If we agree on pooling of 10 samples for

5   the final day 7 culture, the number of actual

6   cultures could be reduced down to 5,000.  Of

7   course, the individual units from positive pools

8   would need to be cultured independently, and Gram

9   stain on individual units is needed to rule out

10   senescent cultures.

11             [Slide.]

12             In addition, there has been a lot of

13   discussion about pre-storage pooling of random

14   donor platelets.  This is an issue that is closely

15   tied to validation of bacterial detection devices

16   because a validated detection device is needed for

17   approval of these types of products.

18             So, to approve pre-storage pooled

19   platelets, what we need is a container, a container

20   that has been validated to store these pooled

21   products out to 7 days.  The validation of these

22   types of containers will require a comparison of

356

1   pre- and post-storage pooled platelet performance

2   in a thrombocytopenic patient population, and these

3   would be corrected count increment type studies

4   that we touched on earlier today.

5             Besides having a validated container, we

6   also need a validated bacterial detection device.

7   Again, we need to define the analytical sensitivity

8   of the detection device for bacteria in a pool of

9   random donor platelets, and these would be spiking

10   studies.  They are designed to account for the fact

11   that you have a dilution of the inoculum when you

12   pool those products together.

13             So, it is a more stringent test on those

14   bacterial detection devices.

15             There will be need for a validation in a

16   field trial for detection of bacteria in the pooled

17   products, however, this could be combined or even a

18   study that is done on single units could be

19   substituted for a field trial of pooled products.

20   We may be able to extrapolate data from a field

21   trial on single unit testing if analytical

22   sensitivity is sufficient to detect bacteria in

357

1   pools where one unit was contaminated.

2             So, this pretty much summarizes where we

3   are today with the bacterial detection and products

4   that are associated with bacterial detection

5   devices.

6             Thank you.

7             DR. ALLEN:  Thank you.

8             Comments or questions for Dr. Vostal?

9             DR. LEITMAN:  I have a question, Jaro.

10   So, you will have two cultures per component, a

11   drawn on the day of collection, one drawn on day 7.

12   I assume the culture sample taken on day 1 is held

13   for 7 days, so you might get discrepant results.

14             So, the culture drawn on day 1 doesn't

15   grow anything, and the culture drawn on day 7 does.

16   You don't really have a gold standard.  How do you

17   know which one to accept?

18             DR. VOSTAL:  I think the cultures would be

19   held for the same amount of time.  Both cultures

20   can be held for 7 days.

21             DR. LEITMAN:  What if they are still

22   discrepant, how do you know that the day 7 reflects

358

1   something that was undetectable on day 1 versus the

2   introduction of an organism during processing in

3   sample acquisition, which we see all the time?

4             DR. VOSTAL:  Right.  I mean that could

5   happen, and there could be a situation where

6   bacteria are present at day 1 at such a low

7   concentration that they are not picked up, and then

8   they peak by day 7.  Now, that would indicate that

9   that initial culture sensitivity or the device was

10   not sensitive enough to pick it up early on in

11   storage at day 1.

12             If the sensitivity was low, you would miss

13   it until the bacteria grew out to a very high

14   level.

15             DR. LEITMAN:  I understand that, but I am

16   not sure that is the correct conclusion.  If you

17   detect it on day 7, and it wasn't detectible on day

18   1, one conclusion is that it was present at such a

19   low concentration, that it didn't become evident in

20   a bacterial storage medium designed to let it grow

21   to large numbers, BacT/Alert or BacTech, or

22   something like that.

359

1             But you know for sure that if you detect

2   it on day 7, not day 1, that the reason that is, is

3   because it was there, but not at high enough

4   concentrations.  The alternative reason is that it

5   is introduced through processing and sampling

6   acquisition.  I don't know how you tell which is

7   which.

8             DR. ALLEN:  I think you make a very good

9   point.  What it basically says is you never can

10   tell for sure whether the sensitivity of the day 1

11   culture, you know, what the exact sensitivity of

12   the day 1 culture is.

13             What you are checking for is the

14   sensitivity of day 1 culture plus the risk of

15   introduction during the manipulation, and probably

16   the risk is so low, it is not going to make a lot

17   of difference, you just lump it all together and

18   call it the sensitivity of the system.

19             I think your point is a very valid one,

20   but I am not sure given the level that we

21   anticipate day 7 being positive versus day 1, that

22   it makes a lot of difference as to why.  We want to

360

1   certainly make certain that whatever we are doing

2   in terms of procedures, that we are absolutely

3   minimizing any risk of introducing organisms,

4   however.  I mean that is a concern I think we all

5   have with any manipulation.

6             DR. KUEHNERT:  I have been involved with

7   some of the discussions, and Jay might want to

8   follow up, but I mean this is a concern.  Blood

9   centers are performing this now in quality control.

10   There are false positive rates that are actually

11   highest at the point that screening is introduced,

12   and then becomes lower as we feel more comfortable

13   with the technique, but definitely, this is

14   something that would need to be addressed in the

15   protocol, is confirming whether indeed this was a

16   true positive versus a false positive.

17             You could go back and reculture the

18   sample, which at that point will not be day 1

19   anymore, but you would expect that culture to

20   reconfirm if it was truly positive at the outset.

21   I think that the group would be including that in

22   the protocol, if that helps.

361

1             DR. VOSTAL:  I think, to answer Dr.

2   Leitman's question, I think what she is getting at

3   is the contamination is not coming from the donor

4   or at the time of collection, it is coming through

5   handling of the product, and I guess we have to

6   make an assumption that this is going to be a

7   closed system and that there is going to be minimal

8   introduction of bacteria during processing.

9             DR. KUEHNERT:  Was that the question?

10             DR. LEITMAN:  Yes.

11             DR. ALLEN:  That is an assumption.

12             DR. KUEHNERT:  That would still be an

13   issue which, if contamination of the unit occurred

14   during processing, that would be a problem you

15   would still want to detect, versus, I thought you

16   were getting at a problem where contamination

17   occurred during the process of actually culturing

18   the unit, which would be a false positive.

19             DR. LEITMAN:  Anything not related to the

20   donor having bacteremia, anything introduced

21   through the process of--I suppose you could

22   introduce something into the component unrelated to

362

1   the donor, but would be a real bacterial

2   contamination, so something related to sampling of

3   the product, to get the sample for culture.

4             If it was discrepant in the other

5   direction, so it's present on day 1, and not

6   present on day 7, you would say oh, that's probably

7   a sampling problem, but you don't know.

8             DR. ALLEN:  Again, I am not up on all of

9   the current devices.  My understanding is that the

10   current manufacturer of platelets, especially, the

11   apheresis platelets, the system is as closed as it

12   is possible to make it to try to avoid anything

13   during any introduction of bacterial contaminants

14   during the processing of the platelets.  Am I

15   correct on that?

16             DR. LEITMAN:  It is just a precautionary,

17   you just want to be careful in how you look at the

18   data that comes back.

19             DR. ALLEN:  The other question, I suspect

20   this may have already been done, is to look very

21   carefully in spiked units at the rate of growth of

22   the bacteria.  I would not necessarily agree that

363

1   they can reach a level of 10

6/ml by day 5 or 7.  My

2   guess would be that by day 2 or 3, that they would

3   have reached perhaps 10

8 to 109/ml, and you may have

4   die-off by the time you reach out to day 5, which

5   is again the point about doing the Gram stain, you

6   know, you could detect organisms there, maybe even

7   better than you could by culture at the far end.

8             I don't know what the growth patterns are,

9   but I think that is something that ought to be

10   looked at.  It wouldn't take very many units and

11   could be quickly done.

12             DR. KLEIN:  Actually, that has been done

13   and most of the organisms that you pick up late are

14   very slow growers, and I was going to ask the other

15   question.  I mean with the low sensitivity of the

16   Gram stain, what is the utility of that test?  It

17   seems to me that that is useless in terms of really

18   finding out anything.

19             DR. VOSTAL:  The Gram stain really is a

20   backup to if you have a senescent culture, one that

21   peaked early and the conditions in the bag get so

22   bad, that the bacteria died off, and they wouldn't

364

1   proliferate in a culture type device whereas, you

2   might be able to pick them up in a Gram stain.

3             DR. KLEIN:  That would be an awful lot of

4   work for a very small return.

5             DR. ALLEN:  But I think that is a question

6   that could be easily answered in a very small, very

7   focused pre-design study to determine whether that

8   would be helpful or not.  I think that is a very

9   good question.

10             I want to just clarify and lay to rest, I

11   know that the AABB made their recommendations that

12   included dipsticks for pH and glucose, and I assume

13   that you are excluding that from your definition of

14   bacterial detection units.

15             DR. VOSTAL:  From our point of view, that

16   is not a validated system.

17             DR. ALLEN:  Fine, that's fine.  That's all

18   I want.

19             DR. VOSTAL:  The study design will be for

20   devices that would validated for these products.

21             DR. HARVATH:  Jaro, could you tell us what

22   the thinking is in terms of how long these cultures

365

1   would need to be held?  I know in the CFR, for the

2   standard thioglycolate testing of, let's say, drug

3   products, I think the requirement used to be 14

4   days, that you had to culture out.

5             Is that going to be true for platelet

6   products for validating, clearing this device?

7             DR. VOSTAL:  The question is whether once

8   the device is validated and approved, how long do

9   you have to hold the culture?  Most of these

10   bacteria grow out quite quickly, so you are looking

11   for, it's a release test, so you want to get an

12   answer as soon as possible, so probably you will be

13   able to release them within 24 to 48 hours.

14             It wouldn't make sense to hold them out to

15   7 days if your platelet shelf life is 7 days.

16             DR. HARVATH:  Could I follow up with that

17   on another question, because the cell and tissue

18   people who are working in somatic cell therapies

19   are struggling with this issue, so let's say people

20   were able to validate this technology with platelet

21   products, would they have to go through this entire

22   process again, for example, to validate it for

366

1   somatic cell therapies, where they are now being

2   told by FDA they have to hold the cultures for 7 to

3   14 days?

4             DR. VOSTAL:  I think that is a different

5   system totally, because you have a lot longer shelf

6   life for those products as opposed to platelet

7   products.  You know, you can afford to hold tissue

8   products for an extended period of time.

9             DR. HARVATH:  Actually, some of those

10   products are collected and used in a relatively

11   short period of time, or in an example, for cord

12   blood, they could be collected, tested, and only

13   stored if the cultures are negative, but people are

14   encountering the same problem for cord blood, and

15   they are using these very same systems that you are

16   talking about for platelets, and that is really

17   what I was thinking about, because those of us who

18   have INDs for cord blood banking have been told

19   that the processes have not been validated for

20   using these bacterial culture systems, these

21   automated systems.

22             So, I think that if this process got

367

1   worked out for platelet products, it would be

2   really advantageous to other types of blood-derived

3   cell products to be able to extend the use in those

4   areas, as well.

5             DR. ALLEN:  Dr. Kuehnert.

6             DR. KUEHNERT:  I just had two questions.

7   One, you talked about study power, and I have these

8   numbers on residual risk.  Are these based on 80

9   percent power to detect, or how are these

10   statistics derived to get the 50,000 number?

11             DR. VOSTAL:  I believe it was 80 percent

12   power.

13             DR. KUEHNERT:  The other question I had

14   concerned, you know, there is a lot of discussion

15   about cost, and I think Dr. Kleinman is going to

16   also talk about it, but has it been worked out

17   exactly what--I mean some of the methodology here

18   is actually being determined by perception of what

19   the costs are and what the most cost intensive

20   methods are.

21             Has that been worked out?  I mean is

22   adding a culture, one culture, say, at day 5, you

368

1   know, is that a large addition to the cost of it,

2   or actually, is it just resources and personnel to

3   do any cultures, has that been worked out at all?

4             DR. VOSTAL:  I think we have had some

5   discussion about where and how to cut the costs of

6   these studies.  I think there is a fixed component,

7   which involves personnel and dedicated time to do

8   these studies, but certainly, you know, if you have

9   to pay for the culture bottles, you know, getting

10   rid of at least one culture, the 5-day culture,

11   would produce some savings.

12             DR. KUEHNERT:  So, if, say, there was a

13   contribution to contribute blood culture bottles,

14   the cost analysis might change significantly, is

15   that right?

16             DR. VOSTAL:  Yes, I would expect so.

17             DR. ALLEN:  Other specific questions or

18   comments for Dr. Vostal?  Yes, Dr. Laal.

19             DR. LAAL:  Is identification of the

20   bacteria any part of this entire process that you

21   are suggesting?  So, what you grow in day 1 and

22   what you grow from the day 7, would you compare if

369

1   it's the same bacteria, or have you introduced

2   something into the system?

3             DR. VOSTAL:  Yes, I think there is a great

4   deal of interest in being able to speciate the type

5   of bacteria that are detected on day 1, as well as

6   on day 7.

7             DR. ALLEN:  Dr. Epstein.

8             DR. EPSTEIN:  I just want to come back to

9   Dr. Harvath's point about applying these automated

10   bacterial detection systems to tissue products,

11   cellular and tissue products.

12             The main point is that it requires a

13   case-by-case validation because there is so much

14   disparity in the properties of all the different

15   tissues, that people might be interested in

16   monitoring these methods.

17             One obvious difference is the solid

18   tissues versus things that are liquid or in a gel

19   state, and then there is the whole issue of

20   different profiles of organisms and the issue of

21   potential contaminations that could affect the

22   culture system, like Wharton's jelly, what effect

370

1   would that have on a cord blood sample, I don't

2   know.

3             So, I think that FDA's point of view is

4   that we are not in a position to generalize.  The

5   difference with respect to blood products is that

6   they are made in a highly stereotypic way, that

7   they are made by thousands of different blood

8   centers, but that they follow very common paradigms

9   for the characterized products.

10             DR. ALLEN:  Dr. Strong.

11             DR. STRONG:  Just a quick comment on

12   tissue.  Certainly, these systems are being used by

13   tissue banks, but obviously off label, and they are

14   said to be validated, but not under the same

15   circumstances that we are discussing here.

16             DR. ALLEN:  We will have a chance to

17   discuss this further.  One final question.  My

18   understanding is that you do not have a formal

19   question for us.  You are asking for committee

20   discussion and recommendations.  Is that on the

21   modification of the study design primarily, or how

22   would you like us to focus our discussion when we

371

1   get to that point?

2             DR. VOSTAL:  It is really an informational

3   session for the committee, but we would appreciate

4   and comments and discussion on the design and how

5   we are proceeding in terms of validating these

6   devices.

7             DR. ALLEN:  Fine.  Thank you.

8             Our second presentation is by Dr. Jerry

9   Holmberg, who is the Senior Advisor for Policy and

10   Executive Secretary of the DHHS Advisory Committee

11   on Blood Safety and Availability.

12        Summary of ACBSA Meeting: Bacterial Contamination

13                      Jerry Holmberg, Ph.D.

14             DR. HOLMBERG:  Thank you.  What I am doing

15   today is going to give you an overview of what we

16   discussed at our April meeting this year and a

17   little bit of background.  The main topic at that

18   meeting was the bacterial contamination issue.

19             [Slide.]

20             Some of the things that the Advisory

21   Committee for Blood Safety and Availability has

22   been involved with has been with the reduction of

372

1   bacterial contamination of blood products, the

2   reduction of errors, TRALI, of course, availability

3   of blood products, and also the constant monitoring

4   of emerging diseases.

5             [Slide.]

6             Just to give you a comprehensive overview

7   and move very quickly, though, some of the things

8   that we did discuss at the last meeting was the

9   National Blood Reserve, which has been a

10   recommendation to the committee and from the

11   committee to the Secretary.

12             We have also discussed the HHS blood

13   monitoring system, which we call the Blood

14   Availability and Safety Inventory System, which

15   when it is completed, it will be an interface with

16   the Secretary's Command Center.

17             Also, we did review the Medicare

18   Modernization Act and specifically some of the

19   HOPPS rules that were published in January 6 of

20   this year.  Some of those changes took place in

21   April for the HOPPS and the reimbursement was

22   changed to reflect back to the January 1 of that

373

1   year.

2             I understand from our colleagues up in CMS

3   that the reimbursement issues will be once again

4   the ruling for 2005, will be published in August,

5   and there will be time for comments on that.

6             [Slide.]

7             We also discussed very briefly a new FDA

8   requirement.  This was an initiative by Secretary

9   Thompson to have bar coding requirements for human

10   drug products. This was effective April 26th.

11   There is a two-year phase-in period, and I think

12   that Elizabeth Callahan has presented this before

13   at the BPAC.

14             Just to bring this up again, I think the

15   primary emphasis was to make sure that blood banks

16   and hospitals were moving in the right direction

17   for incorporating bar codes into their technology

18   in hopes of reducing errors.

19             [Slide.]

20             The main topic of the meeting, though, was

21   the impact and the assessment of methods to reduce

22   the risk of bacterial contamination of platelets.

374

1             [Slide.]

2             A little bit of history on this.  The

3   College of American Pathologists actually had done

4   a Phase I requirement.  This was basically an

5   endpoint testing including swirling methodology.

6   This went into effect in December of 2002, and

7   then, of course, the American Association of Blood

8   Banks introduced their standard, Standard 5.1.5.1,

9   which states that the blood bank or transfusion

10   service shall have methods to limit and detect

11   bacterial contamination in all platelet components.

12   The implementation was March 1st.

13             [Slide.]

14             What really brought this discussion to the

15   table at our April meeting was that my boss, Dr.

16   Beato, who is the Acting Assistant Secretary for

17   Health, did write the American Association of Blood

18   Banks a letter in which he requested that AABB

19   carefully consider delaying the implementation

20   until a clear plan is developed, and also a comment

21   was that it may cause potentially serious and

22   possible unintended effects on the availability of

375

1   platelet products for patient care.

2             Some of the areas of discussion that were

3   mentioned in that letter that really needed to be

4   addressed were quality control methods applicable

5   to pre-release testing, potential extension of

6   platelet dating, pool of whole blood derived

7   platelets, and surveillance and reporting protocols

8   for positive test results.

9             I do appreciate some of the questions that

10   were asked of Jaro concerning the surveillance and

11   the reporting, because that has been an issue of

12   making sure that we have the identification and the

13   surveillance there.

14             One of the problems that we have is that

15   we did not have sufficient data to really look at

16   how this would affect this implementation of the

17   standard, both the CAP Phase I requirement, and

18   also the AABB standard, how that would affect the

19   availability.

20             [Slide.]

21             I think we can skip this slide, but hold

22   on just for a second.  The estimated risk is 1 to

376

1   1,000, to 1 to 3,000 units of platelets, and the

2   mortality rate is thought to be approximately 1 to

3   60,000 transfusions.

4             As you can see, Jaro has already mentioned

5   quite a bit of the activity that has taken place,

6   not only at BPAC, but also CDC, the CDC study, the

7   BaCon study, and also workshops that the FDA has

8   had.

9             [Slide.]

10             We did rely on the American Association of

11   Blood Banks and also America's Blood Centers for

12   some of their survey data to see what was the

13   impact of availability of apheresis and whole blood

14   derived platelets.

15             This is the survey from the American Blood

16   Centers, days to expiration, as you can see, most

17   of the apheresis products had 4 to 3 days

18   availability on them, and the whole blood is

19   relatively consistent all the way through.

20             [Slide.]

21             The American Association of Blood Banks,

22   and when I use the hospital blood banks, I am

377

1   really referring to blood banks being any facility

2   that collects and distributes blood products as

3   opposed a transfusion service.

4             With the AABB survey, increase in platelet

5   shortage at the blood banks.  Seventy percent of

6   their participants in their survey said that there

7   was no increase in platelet shortage, yet, 12

8   percent said yes, and 18 percent said yes, but they

9   didn't know why.

10             [Slide.]

11             If we look at those people that said that

12   there were shortages, an N of 10, 6 of them said

13   that there was a shortage, less than 10 percent

14   shortage in platelets.

15             [Slide.]

16             Also looking at those blood banks, again

17   referring to collection and distribution, in

18   outdating was there an increase in outdating, 21

19   said no, and you can see that the remainder said

20   yes with different percentages of increase in

21   outdating.

22             Now these are the facilities that are

378

1   collecting blood products and distributing.

2             [Slide.]

3             Now, let's turn to the hospital

4   transfusion service.  Has QC testing impacted the

5   availability of apheresis and whole blood derived

6   platelets?

7             Again, one of the things I mentioned at

8   the bottom of each slide, is that it does not

9   include the shortage of the platelet product type

10   in some of these questions that were asked.

11             Increase in platelet shortage at

12   transfusion service--this was an N of 232--58

13   percent said no, that there was not an increase in

14   shortage, 16 percent said yes, and 26 once again

15   said yes, but they didn't know why.

16             [Slide.]

17             Again, looking at the percentage of

18   transfusion services affected by platelet

19   shortages, the majority said that there was a less

20   than 10 percent with 27 reporting 10 to 25 percent.

21             [Slide.]

22             If we also look at the transfusion

379

1   services only in their outdating, 147 of 233 said

2   that there were no increase in outdating, and 51

3   said that they really didn't know whether there was

4   or not.  It really didn't tell us much.

5             [Slide.]

6             We looked at the ABC's survey respondents,

7   54 of the ABC's 77 centers responded.  They collect

8   80 percent of ABC's blood supply.  Thirty-nine of

9   the centers produce whole blood derived platelets

10   and 85 percent produce apheresis platelets.

11             [Slide.]

12             What Dr. Fitzpatrick did in this survey

13   was to look at the centers that had performed

14   bacterial testing before February 2004.  There were

15   12 centers, 7 of the 12 said there were no changes

16   in outdates, and 5 said that there were increased

17   changes in outdates.

18             [Slide.]

19             If we look again at the ABC survey of

20   blood centers, these are the facilities that were

21   testing apheresis products, an N of 54.  Ninety-two

22   percent had implemented bacterial testing.  Two

380

1   percent said they did not make apheresis products,

2   and 2 percent planned to implement, and 4 percent

3   said that the hospitals would be responsible for

4   doing the apheresis platelet testing.  This is

5   apheresis only.

6             [Slide.]

7             If we look at the methodologies that were

8   used by these blood centers to test the apheresis

9   platelets, 78 percent used the BacT/Alert, 20

10   percent use the Pall BDS, and 2 percent used the

11   dipstick.  That 2 percent is one blood center.

12             [Slide.]

13             Now, if we look at the blood centers,

14   again this is ABC survey, testing whole blood

15   derived platelets, 39 percent said that they had

16   implemented testing for whole blood derived

17   platelets, 31 percent said that the hospital would

18   do it, and 24 percent said they didn't make the

19   product at all, and 6 percent said they planned to

20   implement testing of whole blood derived platelets.

21             [Slide.]

22             This slide is the slide that concerns me

381

1   most.  This was the method used to test whole blood

2   derived platelets for bacterial contamination, with

3   62 percent of the whole blood derived platelets

4   being tested by the dipstick methodology endpoint

5   testing.

6             Somebody commented already about the

7   practicality of the Gram stain, and zero percent

8   for the Gram stain, 24 percent were going to be

9   using the Pall BDS, and 14 percent BacT/Alert.

10             [Slide.]

11             When we look at the preliminary data from

12   the blood centers, again, this represents data

13   presented at the April meeting.  Since then, the

14   numbers have increased, the amount of donors

15   tested, donor units to have increased.  The New

16   York Blood Center had 5 true positives.  That is 1

17   in 4,100, false positives 5, again in the same

18   number.  Florida Blood Service had a lower rate, 1

19   in 1,790, and it looks like my typing got a little

20   creative there, 5 in 2,000.  I think that is 2,147.

21   Puget Sound, 1 in 1,800, and false positives, 1 in

22   600.  The totals there being true positives were

382

1   16, 1 in 2,500.

2             So, the estimate that has been presented

3   has been about 1 to 2 to 3,000 rate.

4             [Slide.]

5             The conclusion that was presented or the

6   conclusions that were reached concerning all of

7   this that was presented to the Advisory Committee

8   on Blood Safety and Availability was that many, but

9   not all facilities reported a reduction in

10   production and use of whole blood derived

11   platelets.

12             The outdating has increased especially

13   towards midweek, but seems to be manageable, had a

14   lot of comments from hospitals where their

15   inventory was very tight on a Wednesday, many

16   variables not controlled in both culture techniques

17   and off-label end stage testing.  I will come back

18   to that in just a minute.

19             [Slide.]

20             Nonstandardization, both of endpoint

21   testing and cultures, has led to many variables:

22   time to culture the product, aerobic and/or

383

1   anaerobic testing.  You all referred to that

2   already.  The sampling volume that is used.

3   Detection is based on metabolic activity and may

4   not detect the endotoxins.

5             Determining sensitivity, specificity, and

6   predictive value of early product cultures are

7   currently believed to require large and expensive

8   field trials.

9             [Slide.]

10             The recommendation of the Advisory

11   Committee on Blood Safety and Availability to the

12   Secretary of Health and Human Services is, and I

13   will read this for you:

14             Whereas, the DHHS ACBSA recognizes the

15   importance of methods to reduce the risk of

16   bacterial contamination in both apheresis and whole

17   blood derived platelets; and whereas the community

18   also recognizes the potential for limited

19   availability of platelets, particularly whole blood

20   derived platelets, and whereas the current five-day

21   shelf life of apheresis and whole blood derived

22   platelets and restrictions on whole blood derived

384

1   pre-storage pooling has been identified as barriers

2   to the optimal implementation of bacterial

3   detection in platelets, the committee encourages

4   dialog among DHHS agencies, blood programs, and

5   manufacturers to ensure for:

6             Facilitation of prompt development of

7   technologies, the design and completion of feasible

8   studies, and the satisfaction of licensing

9   requirements to permit both the pre-storage pooling

10   of whole blood derived platelets and extension of

11   platelet dating.

12             [Slide.]

13             If anybody would like to review the

14   complete transcripts, you can go to our web site:

15   hhs.gov/bloodsafety, the transcripts, the

16   presentations from all of the speakers, and the

17   recommendations to the Secretary are listed there,

18   along with a summary of the minutes.

19             Our next meeting will be August 27th and

20   28th, and we will address some of the same issues.

21   We are going to have a follow-up on the bacterial

22   detection studies and some of the progress that

385

1   Steve Kleinman will be mentioning in just a few

2   minutes.

3             Also, I want to mention that the venue has

4   changed.  We will not be at the place we have been

5   the last two times.  We will be back at the Capitol

6   Hyatt, which is up on Capitol Hill on New Jersey

7   Avenue.

8             Any questions?

9             DR. ALLEN:  Thank you, Dr. Holmberg.

10             Yes, Dr. Klein.

11             DR. KLEIN:  The most impressive statistic

12   that I saw in your slides was the fact that more

13   than 40 percent of the 232 hospitals had problems

14   with platelet availability, if I heard your data

15   correctly.  I find that absolutely astonishing

16   whether the 26 percent note was due to testing or

17   not.

18             Did I hear that correctly?  The majority

19   did not, but more than 40 percent having a problem,

20   I find absolutely astonishing.

21             DR. HOLMBERG:  Right, yes.

22             DR. KLEIN:  That was correct?

386

1             DR. HOLMBERG:  This is hospitals.

2             DR. KLEIN:  Hospitals.  The other comment

3   I would make, since I commented on the Gram stain,

4   is that the dipstick is great for swimming pools

5   and for hot tubs.

6             DR. ALLEN:  Other questions?  Dr. Strong.

7             DR. STRONG:  Just a precautionary note,

8   the issue of false positives, I think we haven't

9   standardized the collection of that kind of data

10   either.  For example, in the slide you show, since

11   our number was so high, is because we have also

12   seen problems with the culture system itself.

13             We have had culture bottles that come to

14   us positive, which we count as false positives, and

15   we have had cells in the incubator fail, which

16   weren't immediately identified, which caused

17   positive results to be reported on a number of

18   units, so there are other reasons for false

19   positives other than just contamination at the time

20   of sampling.

21             I would also confirm the use of the

22   dipstick since we are doing that for whole blood

387

1   platelets, and we have done probably 15- or 20,000

2   now, and we have a pretty high false positive rate,

3   and not a single one that has picked up a true

4   positive.

5             DR. ALLEN:  Are you doing that at, what,

6   24 hours after collection?

7             DR. STRONG:  No, we do it at the time of

8   order, so before pooling, each unit is sampled for

9   dipstick.

10             DR. ALLEN:  That is very interesting

11   information.

12             DR. STRONG:  We also have had to lower the

13   bar on that one because in the attempts to validate

14   that sampling system, using criteria of 6.5 for pH,

15   for example, 250 for glucose, we had about a 15

16   percent failure rate, none of which confirmed for

17   positivity, and, in addition, the dipstick was

18   unable to pick up staph cultures, for example.

19             It picks up E. coli nicely at day 4 or day

20   5, but that is about it.  So, we have had to drop

21   the standard down to a pH of 6 and a glucose of

22   zero, and we still get lots of false positives and

388

1   no pick-ups.

2             DR. ALLEN:  Thank you.

3             Other questions or comments for Dr.

4   Holmberg?

5             [No response.]

6             DR. ALLEN:  Thank you.

7             In our third presentation on this topic,

8   Dr. Steve Kleinman is presenting on behalf of the

9   American Association of Blood Banks Task Force.

10               Presentation - Steven Kleinman, M.D.

11             DR. KLEINMAN:  Thanks and good afternoon.

12             Just an introductory remark before the

13   slides come on.  I think we shouldn't lose sight of

14   the fact that although we are not doing something

15   that is perfect, we are removing bacterially

16   contaminated units that otherwise would go to

17   patients, and in some cases, would cause serious

18   morbidity and kill a few people, so while progress

19   is not perfect, I do think we are making progress

20   in the clinical safety area.

21             At any rate, I am going to talk today

22   about the Interorganizational Task Force that has

389

1   been established.

2             [Slide.]

3             You have heard this before, the standard

4   was introduced by AABB in March of 2004, but the

5   standard was talked about for over a year and

6   published, and had plenty of time for comment, and

7   it is something that was well known to be on the

8   horizon before it actually was introduced.

9             [Slide.]

10             So, the purpose of the task force, in

11   general terms, since the AABB put this standard in

12   place, it seemed logical that the AABB should have

13   a mechanism by which to evaluate and move forward

14   on all of the implementation issues that have

15   arisen.

16             So, this task force is an opportunity to

17   centralize all those issues within one body in the

18   American Association of Blood Banks, as well as

19   involve people, many interested parties outside of

20   the AABB.

21             So, in a general sense, the task force is

22   monitoring all the activity or trying to monitor

390

1   the activity that is going forward with testing,

2   and you saw the results of the first AABB survey

3   that Jerry presented.

4             Thirdly, the task force provides a forum

5   for discussions amongst many interested parties, as

6   I will go into, and that involves obtaining the

7   input from agencies, such as FDA and other parts of

8   HHS.

9             [Slide.]

10             Representation on the task force, American

11   Association of Blood Banks, America's Blood

12   Centers, American Red Cross, Department of Defense

13   as blood collectors.

14             Within that membership, we have

15   representation from transfusion services, blood

16   collection facilities, transfusion services that

17   use a lot of apheresis platelets, transfusion

18   services that use a lot of whole blood derived

19   platelets, chairs of multiple committees within

20   AABB, and then we have representatives from FDA,

21   CDC, the Availability Committee within HHS, and

22   NHLBI, about I think 15 to 20 people all together.

391

1   We are also working to get liaisons from the

2   American Society of Microbiology and the Infectious

3   Disease Society.

4             [Slide.]

5             More specifically, what has the task force

6   done?  It has only been in operation since the

7   beginning of May, and I think we have made some

8   progress, but we have run into the summer, which

9   means it's a little bit more difficult to convene

10   large groups of people and move forward rapidly.

11             At any rate, we did have a public meeting

12   that took place in the beginning of June that

13   included all of the task force, as well as many

14   representatives from many of the companies that

15   current make bacterial detection systems or are

16   contemplating bringing tests to market.

17             One of the aims of the task force is to

18   continue with data collection efforts, and

19   currently, we are planning another survey of member

20   institutions to ask questions both about platelet

21   availability, outdating, methods of culture use,

22   results, impacts that have come from the system.

392

1             We feel that another survey is needed

2   because the first survey went out right at the

3   time, within a month or two, of people actually

4   implementing or beginning to implement the

5   standard, and we think that another look at what is

6   happening will give us a little bit more of what is

7   happening in a more stable sense now.

8             So, we are hopeful that that will get

9   underway next month of September.

10             Additionally, a number of issues have been

11   discussed that need to be I think standardized, as

12   many of the speakers have mentioned.  So, our task

13   force is preparing guidance that will be written

14   and sent out to membership.

15             Some of the items that we are going to

16   discuss that haven't really been explicitly

17   discussed up until now is what to do if you get a

18   positive culture.  I can tell you one of the things

19   that we are strongly recommending, I guess we can't

20   require anything, but we are strongly recommending,

21   is that all organisms be speciated, identified that

22   is.

393

1             We are also recommending that any

2   co-components or plan to recommend that

3   co-components from positive platelet units be

4   retrieved and not transfused, so red cells and FFP

5   be retrieved.

6             With regard to what we are calling

7   surrogate tests, not direct bacterial detection

8   systems, something like pH or glucose monitoring,

9   we will have some guidance on what to do to follow

10   up those test results.

11             We want to also address the issue, we

12   think it only occurs rarely, but clearly it can

13   occur, that you transfuse a platelet unit and

14   because at least in the most commonly used system,

15   the BacT/Alert, you keep the culture going for 7

16   days, you may come up with a positive test result

17   after the unit has already been transfused to the

18   patient.  So, we are coming up with some guidance

19   to how to manage that situation.

20             Fourthly, we are going to address the

21   issue of depending on your findings of cultures,

22   when should you actually notify donors of the fact

394

1   that they have had a positive culture and what kind

2   of recommendation should you make to them.

3             Additional items that will be included in

4   the guidance that we are working on, are

5   standardized definitions for what represents a

6   confirmed positive versus a presumptive positive

7   versus a false positive that people can then use to

8   guide their data collection, so that when we have

9   various organizations presenting their data, they

10   are all using the same definitions.

11             That is a first step to compiling some

12   kind of national experience.  The next step then is

13   to be able to find a way to have everybody report

14   their data into a central mechanism.  Well, that is

15   the problem with all data on blood collection in

16   the U.S., that we don't have centralized

17   mechanisms, each organization has its own

18   mechanism, but hopefully, we will have standardized

19   definitions by which we can capture information.

20             We are also working to decide what

21   organisms might be reportable and whether there are

22   any public health issues involved in certain

395

1   isolates.

2             So, all of these things, we hope to be

3   addressed in a guidance document, one or more

4   guidance documents with the goal of getting the

5   first one out, I would guess by the beginning,

6   well, sometime next month or September again.

7             [Slide.]

8             I reference additional actions.  I

9   referenced this point before.  We want to monitor

10   the effects of bacterial detection testing on

11   platelet availability.  We hope to do that through

12   the survey mechanism or other mechanisms that we

13   may be able to define.

14             The two crucial issues that were

15   identified by HHS are also subjects for this task

16   force, and they are helping to trying to talk

17   through the issues in a clinical protocol, as you

18   have heard Jaro discuss, that would permit the

19   extension of the dating period from 5 days to 7

20   days.  You have only heard a little bit about these

21   studies.  I will go into them a little bit more.

22             They are very complex actually and

396

1   logistically very difficult to bring off.

2             Secondly, we have the other issue of

3   pooling whole blood derived platelets, and just a

4   couple comments about that.

5             As you heard, because of the logistics of

6   testing individual platelet units, and presumably

7   also because of the expense, most people have not

8   decided to perform a culture on individual platelet

9   units, because that would be 5 to 6 cultures for

10   one platelet transfusion dose versus one culture on

11   an apheresis product.

12             That is why people have resorted to these

13   unsatisfactory surrogate measures like pH and

14   glucose.  We agree that that is not a very

15   satisfactory way to decrease bacterial risk to

16   patients, and it seems like the best way to solve

17   that problem is to do what they do in Europe, and

18   that is to pool platelets and be able to store

19   them, and then take a culture of the pool.

20             Now, they use a different technology in

21   Europe to make their platelets, whole blood derived

22   platelets, as I am sure you know, but I think many

397

1   people on the task force consider this an important

2   priority because that is going to permit the

3   application of the more sensitive culture

4   techniques, the bacterial detection.

5             However, there is one other potential

6   approach to those units, and that is, if

7   manufacturers who are working on point of release

8   testing can get their products to market and

9   validated and licensed by FDA, then, we may have

10   alternatives, so we see two approaches to the whole

11   blood derived platelet problem, and we are hopeful

12   that we can move one or both of those forward.

13             Finally, I mention the potential public

14   health reporting issues that we are also

15   considering.

16             [Slide.]

17             I want to just go a little bit into the

18   extended platelet storage protocol that we spent a

19   good deal of our meeting talking about at our first

20   task force meeting.  I thank Dr. Mark Brecher for

21   putting some of this information together.  That is

22   why you see the UNC logo on the slides.

398

1             [Slide.]

2             So, the specific aim of the study is to

3   assess the predictive value--and you have heard

4   Jaro talk about this, so hopefully, it is sort of

5   an evolving aim, I think what I have on these

6   slides may be slightly different from what Jaro

7   said, but it is in the same ballpark--to assess the

8   predictive value of an early bacterial culture of

9   platelets (that means a sample taken on day 1

10   and/or day 2 of storage) for detection of bacterial

11   contamination.

12             This would then allow for a product claim

13   because sensitivity, specificity, and predictive

14   value in a clinical trial is something that the FDA

15   is requiring, and for release testing, so

16   hopefully, this study would allow manufacturers who

17   participate to derive the data to make that claim,

18   and then we would have an actual release test.

19             The endpoints, as we understood them at

20   least a month ago, were detection of greater than

21   80 percent of bacterially contaminated platelets by

22   early culture, and then with a 7-day residual risk

399

1   of less than 1 in 10,000.  These are, as I said,

2   kind of evolving and we are going to have some more

3   discussions on what they should be.

4             [Slide.]

5             This summarizes a study done up in Canada

6   with the Pall System, and there are other small

7   studies out there where people have looked at early

8   versus late cultures.  This slide was presented

9   previously at the meeting.

10             It is not quite indicative of the study

11   because the early cultures were not all done on day

12   1.  They were actually done on the day that the

13   hospital received the product, and that often was

14   day 3 or day 4, and then the later cultures were

15   done after pooling, at the time of release, and

16   they did find one additional positive further one

17   that had been negative at the time that they tested

18   it.

19             I only show this slide to illustrate the

20   kind of data that we might be able to generate by

21   such a study.

22             [Slide.]

400

1             What about the size of the study?  Well,

2   we have heard this number 50,000, and I have

3   reproduced it here, and that is what we think will

4   provide a reasonable assurance of a residual risk

5   less than 1 in 10,000 for the early culture.

6             [Slide.]

7             These are the BacT/Alert bottles, both the

8   aerobic and the anaerobic, and in one formulation

9   of the study, where we do a culture at day 1,

10   perhaps another culture at day 2, perhaps a culture

11   at day 7 or later, where you can see there is a lot

12   of culturing to be done if we go this route.

13             Why did I put day 7 or later?  We have a

14   real problem, I think right now, and that is if we

15   do this study with an endpoint at day 7, and our

16   platelet storage containers get better, and we are

17   able to store out platelets for 10 days, will we

18   have to repeat this same clinical trial with

19   10-day-old platelets.

20             I think that would be an unfortunate

21   occurrence if we can't find a way to get around

22   that with an initial study protocol.  So, maybe one

401

1   solution is to, instead of testing at day 7, even

2   though the containers are only validated for 7-day

3   storage, can we test later in anticipation of the

4   future rather than having to repeat what is going

5   to turn out to be a logistically complicated and

6   expensive study.

7             [Slide.]

8             Well, here is what we might expect.  If we

9   do culture 50,000 platelets, and our rate of

10   positivity is 1 in 1,000, and we actually detected

11   50 positive cultures early, and we wanted to be

12   sure that we were detecting at least 80

13   percent--and again, these criteria appear to have

14   changed, maybe the criteria is going to be just the

15   residual risk of 1 in 10,000--you can see that

16   depending on our initial culture positivity rate,

17   we would have a certain number of allowable

18   positive late cultures to meet the criteria.

19             In a 1 in 4,000 rate, which seems to be

20   about the true positive rate that we think we are

21   seeing now, 1 in 2,500 from the early data, but it

22   seems like it might be a little bit less than that,

402

1   you can see that if we only had 3 contaminations on

2   day 7, the whole study would fail.

3             I think Dr. Leitman's statement about the

4   fact that you can have false positive cultures

5   introduced at the time you take the sample, even

6   when you take the day 7 sample, and from the

7   initial data, finding out that that happens just as

8   frequently as true positive cultures, I think it

9   really warrants serious consideration about whether

10   we are setting ourselves up to fail because of

11   contamination problems that presumably we don't

12   know how to avoid at this point.

13             [Slide.]

14             Well, let's say we could do this and we

15   were going to do 3 sets of cultures per unit, and

16   we weren't doing the pooling that was mentioned,

17   this is a lot of culturing, because you have to

18   keep your cultures going for 7 days, and in order

19   to do this, the actual testing in one year, we

20   would probably need at least four laboratories and

21   lots of equipment.

22             The initial estimate, and I know I have

403

1   seen it, suppose that we have to pay for

2   everything, and that manufacturers aren't

3   supporting this at all, but it eliminates, it

4   doesn't completely account for all of the costs, it

5   just really accounts for the testing costs, and not

6   other costs of the study, is over $5 million.

7             Another point that is important is this is

8   not like doing a study on any other infectious

9   disease marker because here you can't just retain a

10   tube, remember you need 50,000 outdated platelet

11   products.  We don't want to outdate 50,000 platelet

12   products.

13             We hope we are not outdating apheresis

14   products at that rate, or we are really not doing

15   out job, so one of the proposals is that we would

16   actually do this study with units that were with

17   whole blood derived platelets that are actually

18   specially manufactured for this study, so that we

19   don't have to wait until if a platelet is made and

20   sent out and then outdates, we would have all of

21   the logistics of getting that platelet back from

22   the hospital into the center.

404

1             I don't think that is a cost issue, I

2   think that is a logistics issue that would make the

3   study unworkable, so we have arrived at the

4   possibility of specially manufacturing platelets

5   just for the purpose of the study.

6             [Slide.]

7             Well, I want to go back to the more basic

8   question, because the study is complicated, it is

9   probably doable, but do we really need a clinical

10   study to increase our storage time to 7 days?

11             A couple of points that are worth making.

12   There was a time, as you know, when storage was 7

13   days, and that was at a time when there was no

14   bacterial testing.  Many European countries have

15   already extended their storage to 7 days using the

16   same bacterial detection systems that we are still

17   testing in this country.

18             Now, they use different production methods

19   for their platelets, but they haven't done the

20   clinical studies in order to do this.  Now,

21   obviously, we regulate the way we choose to

22   regulate in this country, and they regulate the way

405

1   they choose to, but can't we use that precedent,

2   shouldn't that be telling us something?

3             [Slide.]

4             We know that the major clinically

5   significant bacteria are detectable prior to 5

6   days.  I mean there is ample evidence from spiking

7   studies and from even clinical isolate studies

8   around the world, that there are very few bacteria

9   that actually come up after 5 days.

10             Most of them, maybe they are not all there

11   by 1 day, I understand that, so we are sort of all

12   expecting the study to succeed, so how strong is

13   the data already to tell us that we would not

14   impair safety and can we do it with spiking

15   studies.  Obviously, that has not been the decision

16   to this point.

17             Now, if we are really concerned about

18   getting 7-day platelet storage, if we have to do

19   this study, we are clearly several years away.  We

20   haven't gotten the study underway, it is going to

21   be a minimum 1-year study, it is going to have to

22   go through regulatory review, so clearly, we are

406

1   not going to get there very quickly, and then,

2   finally, does this study set a precedent for

3   requirements to further extend platelet storage

4   beyond 7 days.  I have already covered that point.

5             So, I think we should think carefully.  If

6   we have to do the study, we need to find out how to

7   make it logistically possible.  Clearly, if the

8   sample size was smaller, it would be easier to do.

9             I will conclude there.

10             DR. ALLEN:  Thank you.

11             Dr. Klein.

12             DR. KLEIN:  Steve, does your $5.4 million

13   estimate, is that for testing one manufacturer's

14   system?

15             DR. KLEINMAN:  Yes, that is for the

16   BacT/Alert assuming that the equipment and the

17   culture bottles would need to be paid for out of

18   the study funds, and it includes aerobic and

19   anaerobic cultures.  I am sure there is a way to

20   bring that cost down.

21             DR. KLEIN:  But all those people who are

22   using other systems, then, would have to have their

407

1   own $5.4 million studies, is that correct?

2             DR. KLEINMAN:  I don't know.  I think it

3   may be possible to involve multiple manufacturers

4   in the same study, but again, that might increase

5   the cost.  That hasn't been budgeted in at this

6   point.

7             DR. ALLEN:  But that certainly makes

8   sense, doesn't it?  I mean the question is whether

9   you are validating the system or whether you are

10   validating the process of doing the testing, and I

11   haven't thought that one through.

12             DR. KLEIN:  I think you are clearly

13   validating the system.

14             DR. ALLEN:  It is an interesting question

15   of whether you would need to have an N of 50,000

16   for each method proposed.

17             DR. KUEHNERT:  I mean some systems are

18   very, very similar, like BacT/Alert versus BacTech,

19   for instance, automated culture systems versus, you

20   know, other systems have different mechanisms like

21   the Pall BDS, so I guess it would depend on the

22   differences between the systems.

408

1             DR. KLEINMAN:  I think another question

2   would be whether other manufacturers who are

3   working on other tests that aren't currently

4   licensed, also would have a product available and

5   want to take advantage of this study to run their

6   test, as well.  That all needs to be worked out

7   still.

8             DR. ALLEN:  Steve, you have quoted the

9   presentation from Jim AuBuchon about the one late

10   culture that was found.  Do you know what the

11   organism there, was it a skin contaminant?

12             DR. KLEINMAN:  It was a Staph, I can't

13   remember if it was an epidermidis.

14   Coagulase-negative staph, I think, Staph

15   epidermidis probably.

16             DR. ALLEN:  Can you say just a little bit

17   more about--you mentioned the public health

18   reporting issue?

19             DR. KLEINMAN:  The concern here is, will

20   we detect--I mean the primary concern is we need to

21   know the list of notifiable diseases in various

22   states, I mean this gets into the question that if

409

1   we find an infection in an asymptomatic person, is

2   it reportable.  I don't know, but if we do find a

3   bug that we think is coming from bacteremia, and

4   it's a notifiable bug, then, I assume, at least our

5   colleagues at CDC have been advocating that we

6   should have  a system in place to get those

7   reported, and that is one of the things that we are

8   working through.

9             Now, whether we need to do that or not is

10   still something that I think the task force is

11   debating.  It is not really, you know, you wouldn't

12   choose to do blood cultures in asymptomatic donors

13   to do bacterial surveillance in your community.

14   You would probably go to hospitalized patients to

15   find out what kind of really bad organisms might be

16   floating around in the community, so maybe Matt

17   would like to address that.

18             DR. ALLEN:  I will address it, having

19   worked in both state and local health departments,

20   as well as the Federal Government, I think that is

21   a non-issue.  Certainly, for something that clearly

22   is an important public health problem, the West

410

1   Nile virus, the blood collection centers have found

2   a way to deal with it.  I think this is an

3   non-issue.

4             With very rare exceptions, that being the

5   decades old instance of the donor at the NIH

6   Clinical Center who had, what was it, Salmonella

7   osteomyelitis.  You know, with that one exception

8   that I am aware of, I think this is a non-issue.

9             DR. KUEHNERT:  Could I say something?  I

10   think Steve said it well, I don't think this is

11   something, a surveillance system someone would want

12   to use to detect infections on a prospective basis

13   as a way to definitively detect infections that are

14   reportable, but I think if they occur, you want to

15   know about them, and that is just what this is,

16   just making sure that there is a mechanism for the

17   blood banking system to interface with the state

18   and territorial epidemiologists.

19             So, for instance, something like you never

20   thought might happen, like a Listeria bacteremia,

21   or Salmonella, it is not that we think that is

22   going to happen all the time, but that it is good

411

1   to have a mechanism, so that it gets reported and

2   followed up.

3             DR. LEITMAN:  Steve, when we asked the

4   head of our hospital micro department whether we

5   should use anaerobic cultures, and we started this

6   back in September of 03, he commented that if you

7   hold the culture for 5 to 7 days, anaerobic

8   organisms of significance will grow out in an

9   aerobic culture, only more slowly, and advised

10   against doing that.  So, we took his advice.

11             As everyone here knows, the predominant

12   organisms are skin organisms, staph and strep, and

13   then there are others, but anaerobes, significant

14   anaerobes occur, but they are very rare, and what

15   is the evidence that you don't get them with a

16   7-day aerobic culture?

17             DR. KLEINMAN:  I don't really know the

18   data well enough to give you a good answer.  I

19   think the ones that you get in the anaerobic model

20   tend to be the propion and bactam acnes, and I

21   guess you can get them in the aerobic model, as

22   well, so I don't really know what the advantage of

412

1   doing the anaerobic culture.

2             I think there are a couple, there is at

3   least one report of clostridial sepsis that killed

4   a patient, and that organism would not have been

5   picked up by an aerobic bottle, as I understand it,

6   and that is where the advocates for doing anaerobic

7   cultures I think get--advocate that policy, but I

8   think, as you will find, I don't think it has been

9   systematically surveyed across the people who are

10   doing cultures, but I know the majority of large

11   blood collectors who are using the BacT/Alert are

12   not doing anaerobic cultures at this point in time.

13             I couldn't give you a percentage, but I

14   think it is the large majority.  Most people, when

15   they have independently evaluated this, have

16   decided not to do anaerobic culture.  Of course,

17   the Pall BDS system depends on oxygen levels, it

18   won't pick up any anaerobes.

19             So, I think the large majority of culture

20   tested platelets in the U.S. are not being tested

21   by the anaerobic system.  Whether or not we need to

22   do anaerobic cultures as part of this study is

413

1   still something that is under discussion.

2             DR. ALLEN:  Dr. Epstein.

3             DR. KLEINMAN:  Thank you, Jim.

4             I just wanted to comment on a number of

5   the issues that have come up in the discussion.

6   First of all, about the false positives, there is

7   no way to eliminate false positive cultures, but

8   there has been a lot of discussion about a plan to

9   save samples, so they are available for reculture.

10   One could also use a titer, since titers should

11   rise over time in real positives.

12             So, that is a recognized issue, but not a

13   totally solved problem, but retention samples and

14   reculturing is intended to be part of the protocol,

15   and the analytic data that have seen in terms of

16   the model is talking about confirmed positives,

17   recognizing that there may be still some small

18   error in what you call a confirmed positive.

19             The second point is on Steve's slide where

20   you talked about three positives at 7 days could

21   sink the ship if you hadn't expected upfront a rate

22   of 1 in 4,000, but that was predicated on the

414

1   notion that the screen has to be 80 percent

2   sensitive, and FDA's current thinking is that we

3   shouldn't predetermine that, that the study will

4   tell us the sensitivity upfront test, and that can

5   then be dealt with by truth in labeling, that we

6   are actually more focused on whether there is net

7   risk reduction and what is the residual risk.

8             Let me just comment that if you do a study

9   of 50,000, you could have five positives at 7 days,

10   missed at day 1/2.  That is still a residual risk

11   point estimate of 1 in 10,000, and the 95 percent

12   upper confidence bound is 1 in 5,000, which we

13   would find acceptable because that is still less

14   than the upfront rate found even in the best

15   systems day of 1 in 4,000.

16             So, we are not focused on the 80 percent

17   sensitivity issue.  We would let the study tell us

18   how good the test is.

19             Then, with respect to the issue of, well,

20   what about tests that come forward down the road,

21   other technologies, including endpoint tests.

22   Again, that has had discussion, and the general

415

1   approach is that, well, we ought to really save the

2   samples from the study, so that the samples can be

3   reanalyzed without all the logistics of a new

4   study.

5             The other approach that could be discussed

6   is if newcomer tests are analytically equivalent.

7   Then, we might have a much better basis to accept

8   them than we do now, because right now we have

9   tests with certain analytical characterizations

10   based on in vitro studies, but we don't know

11   meaningful they are clinically, whereas, if we know

12   how meaningful they are clinically, and then we

13   have analytic equivalence, then, we might be better

14   positioned to deal with newcomer technologies with

15   much, much smaller, quote, unquote "bridging"

16   studies.

17             That begs the question of who pays the

18   upfront cost, and I think it is one of the

19   arguments in favor of public funding for the study

20   that has been discussed.

21             Then, other dimensions of the study

22   haven't been sort of fully developed here, but FDA

416

1   also sees a need to validate a standardized upfront

2   procedure, because one of the problems is that

3   there isn't presently standardization of the

4   volume, the sampling time, the number of bottles,

5   should it contain an anaerobic culture, whereas,

6   there has been reasonable speculation about the

7   ability to detect facultative anaerobes, the bottom

8   line is that only a study will really tell us what

9   we are missing.

10             So, the question is if you can fund the

11   study, isn't it desirable to find out with the

12   anaerobic culture.  I don't disagree with the

13   speculation, but we have had sepsis and fatality

14   from anaerobes.  What we don't know is the

15   frequency.

16             Lastly, the issue of the 7-day culture.

17   We have sort of been a little loose.  Sometimes we

18   mean holding the culture for 7 days when that is

19   the day 1 culture, and sometimes we mean the

20   culture performed on day 7.

21             I just think everybody needs to remember

22   that in the BioMerieux system, the culture can

417

1   continue to be incubated essentially as long as you

2   want, but in the Pall BDS system, it is a one-shot

3   deal, and you basically get your readout at the

4   chosen time of 24 or 38 hours, and that's it, and

5   that there is also this issue of sampling error,

6   that the colony count could be below the threshold

7   for a culture in either of those two systems.

8             If it is low enough, you can simply miss

9   it in your sample, and then it is not going to grow

10   out, it doesn't matter how long you hold it.  That

11   is again one of the issues with measuring the

12   sensitivity upfront culture because you are also

13   looking at its threshold, and the systems have been

14   validated for thresholds of between 10 and 100

15   CFU/ml in the sample as taken on day 1 or day 2.

16             The whole question is, well, how good is

17   that really as a predictor of culture negativity on

18   day 5, day 7, or day 10, and the answer is nobody

19   knows.

20             So, these are just some of the

21   methodological issues that have been under

22   discussion, and as Steve says, they are continuing

418

1   to evolve, but I think that the big breakthrough

2   has been that if you put aside for the moment of

3   where will we find the money, there has been a lot

4   of movement toward consensus about what should the

5   study accomplish, and I think that that has been a

6   tremendous step forward with the cooperation of a

7   lot of groups and individuals.

8             DR. KLEINMAN:  I think the removal of that

9   80 percent criteria clearly makes it a much more

10   reasonable endpoint in order to have a potentially

11   successful out, and to therefore compare to the

12   current product, which is an untested product, and

13   to show that safety is increased even with the

14   7-day stored product.  I mean that is the aim.

15             So, I think there is a lot of movement

16   there.

17             DR. ALLEN:  Any other questions for Dr.

18   Kleinman at this point?  You will be around for the

19   rest of this afternoon?  Okay.

20             We now move to the open public session.  I

21   have on my agenda here, one person who wishes to

22   speak, Dr. Boris Rotman, BCR Diagnostics.

419

1             I have to read my little statement, first

2   of all.

3             Open public hearing announcement for

4   general matters meetings.

5             Both the Food and Drug Administration and

6   the public believe in a transparent process for

7   information gathering and decisionmaking.  To

8   ensure such transparency at the open public hearing

9   session of the Advisory Committee meeting, FDA

10   believes that it is important to understand the

11   context of an individual's presentation.

12             For this reason, FDA encourages you, the

13   open public hearing speaker, at the beginning of

14   your written or oral statement to advise the

15   committee of any financial relationship that you

16   may have with any company or any group that is

17   likely to be impacted by the topic of this meeting.

18   For example, the financial information may include

19   a company's or a group's payment of your travel,

20   lodging, or other expenses in connection with your

21   attendance at the meeting.

22             Likewise, FDA encourages you at the

420

1   beginning of your statement to advise the committee

2   if you do not have any such financial

3   relationships.  If you choose not to address this

4   issue of financial relationships at the beginning

5   of your statement, it will not preclude you from

6   speaking.

7             Thank you.

8                       Open Public Hearing

9             DR. ROTMAN:  First of all, I am a newcomer

10   according to the last comment.  I am a Professor of

11   Brown University Medical School, but I have a

12   conflict of interest as it is clear here.  I own a

13   small company that is funded by NIH.

14             [Slide.]

15             This is my first meeting here, and I am

16   sort of interested in all the comments, because I

17   think that the system capable of detecting bacteria

18   in real time will avoid many of the pitfalls that

19   we have been hearing all afternoon.

20             This slide indicates what has been

21   commented the prevalence with the current

22   methodology.

421

1             [Slide.]

2             These are the advantages of testing at the

3   point of care.  First of all, the storage can be

4   extended, and there are no false negatives because

5   all the bacteria are tested directly from the

6   platelets.  There is no problem with laboratory

7   medium or the kinetics of measuring bacteria on day

8   1 or 2.

9             Third, we all know the number of bacteria

10   is much larger at the end of the storage period, so

11   the chance of false negatives is much lower.

12             Fourth, distribution will be very, very

13   simple. This is a system that also is inexpensive

14   and can be used in the hospitals.

15             [Slide.]

16             I going to go briefly here.  I don't want

17   to overwhelm the audience with scientific

18   information.

19             The system is based on three technologies.

20   One is new, it is using spores as nanodetectors.

21   The other is well known and is increasing assay

22   sensitivity that is obtained using extremely small

422

1   volumes.  In this case, we are using 5 picoliters.

2             Then, we use a state of the art image

3   analysis and data acquisition.  I ought to mention

4   that the device is actually composed of 80,000

5   independent biosensors as I am going to show you in

6   a minute.

7             [Slide.]

8             These are the characteristics of the

9   instrument. It can detect at the level of

10   individual cells.

11             [Slide.]

12             It gives results in about 15 minutes.

13             [Slide.]

14             Low cost as I mentioned, and then it has a

15   very large dynamic range.

16             [Slide.]

17             The instrument is relatively inexpensive.

18             [Slide.]

19             This is the principle.  Bacteria emits

20   biochemical signals that are traditional.  We use

21   the enzymes that have been used for hundreds of

22   years to identify bacteria.  Our favorite is

423

1   aminopeptidase.

2             Then, we use as a detector, a spore, that

3   in the presence of the bacteria, emit fluorescent

4   light, which is captured by a standard methodology.

5             [Slide.]

6             This is the cross section of the

7   instrument.  Actually, we call it a biochip, and

8   each biochip has, as I said, 80,000 of these wells

9   that we call microcolanders, because they act as

10   collecting and filtrating agents or compartments.

11             Each microcolander has about 200 spores,

12   and a single bacterium that happens to fall into

13   one of the colanders will trigger a chain reaction,

14   which makes the spore emit fluorescent light.

15             [Slide.]

16             This is an actual view of the biochip

17   under the electron microscope.  Each of the

18   microchips is 20 microns in diameter.

19             [Slide.]

20             This is an operation.  The test sample is

21   mixed with the spore.  It filters through the

22   biochip.  It is incubated and then, as I mentioned,

424

1   the chain reaction occurs and fluorescence appears,

2   and an image is captured and analyzed by standard

3   computing systems.

4             [Slide.]

5             This is an actual picture of the assay and

6   illustrates how simple it is.  In this particular

7   section of the biochip, there are 4 bacteria as is

8   shown on the right.

9             On the left, we have the biochip

10   photograph under normal light.

11             [Slide.]

12             This is about our company.  It is a small

13   company in Rhode Island, and we have been awarded a

14   Phase II grant specifically for detection of

15   bacteria in platelets.

16             [Slide.]

17             This is our list of publications.

18             Thank you.

19             DR. ALLEN:  Thank you, Dr. Rotman.

20             Any questions or comments from the

21   committee members?  Yes.

22             DR. STRONG:  Can you put this in the

425

1   marketplace in the next two months?

2             [Laughter.]

3             DR. ROTMAN:  Well, as a matter of fact,

4   no.  The purpose of this talk, I am glad you asked

5   this question, is that we would like to get from

6   FDA some guidelines because obviously, this method

7   cannot be tested by the conventional spiking.  We

8   need to have an entirely different guidelines.

9             We estimate that we will be in the market,

10   we are hoping to send FDA data in about 8 or 9

11   months.

12             DR. ALLEN:  So, future technology.

13             DR. DiMICHELE:  Do you think there is the

14   potential that this methodology might be too

15   sensitive in terms of, you know, I mean in terms of

16   it would be very interesting in terms of what you

17   would set the threshold for, you know, clinical or

18   potentially pathological sensitivity versus--maybe

19   not.

20             DR. ROTMAN:  Yes, I am afraid that it will

21   open a Pandora box for FDA, because what is

22   acceptable, 100 bacteria?  1,000?  1?  It is up to

426

1   FDA.

2             To give you an idea, right now the

3   methodology is probably detecting a million per ml.

4             DR. ALLEN:  Thank you very much.

5             This, I think is a promising avenue of

6   discussion for the future.

7             One more comment.  Would you identify

8   yourself, please.

9             DR. HOLME:  Stein Holme, Pall Corporation.

10             I have a few comments regarding the safety

11   risk of 7-day stored platelet as compared to 5-day

12   pre-storage pooled platelet product.

13             In my view, the risk of sepsis with a

14   7-day stored platelet product is much larger than

15   what you will find with a 5-day pre-storage pooled

16   product.  The reason for that is that based on the

17   studies we have done, the levels of bacteria, the

18   high levels of bacteria occur between 3 to 4 days

19   of storage.

20             By shifting the shelf life to 7 days, the

21   average life or shelf life of the platelets that

22   are going to be transfused is going to be increased

427

1   from 2 to 3 days, to 3 to 4 days.  That means that

2   the dose that potentially will get in a

3   contaminated product with shelf life of 7 days is

4   going to be several logs higher, is going to be

5   10,000, 100,000 higher.

6             However, with regard to pre-storage pooled

7   product, we talk about a much lower level of

8   platelet since we are pooling up to 4 to 6 platelet

9   product together, the increased level potentially

10   could be only 5 times, not several log times

11   higher.

12             So, the safety you see in terms of risk of

13   sepsis to a patient is going to be quite different

14   in terms of 7 days versus 5 days pre-storage pool.

15   I think that needs to be considered.

16             Also, in Pall Corporation, we have

17   developed, we basically know what the sensitivity

18   of our system is.  It is about 1 CFU/ml.  So,

19   basically, if you get a bug into the culture pouch,

20   it will be detected.

21             Currently, we heard today that random

22   donor platelet post-stored using dipstick or pH or

428

1   swirling is not satisfactory sensitive, so instead

2   what we have is to propose, is to have for

3   pre-storage 5-day pooled storage product, we can

4   have a device which is much, much more sensitive

5   where you can take a sample on day 1 or day 2 of

6   storage and have a much safer system in terms of

7   the risk of sepsis than you will have today with

8   the 5-day stored platelet that has been transfused

9   post-storage.

10             Thank you.

11             MR. WAGNER:  Just one comment.  Steve

12   Wagner, American Red Cross.

13             The ability to detect one organism in 5

14   picoliters is I believe equivalent to having a

15   limit of sensitivity of 2 x 10

8 organisms per ml.

16   Please correct me if I am wrong.

17             DR. ROTMAN:  You are correct that you are

18   wrong.

19             This is one organism in the biochip, and

20   we are using half an ml of platelets, and biochip

21   filters the sample.  That is the crucial thing

22   about the biochip, that it can use 1, 2, or 3 ml

429

1   samples, and in the 80,000 wells, 1 bacterium

2   landed out of the half ml of the sample, so that

3   means 2 bacteria per ml.

4             DR. ALLEN:  Thank you for the question and

5   the clarification.  I am going to ask that we move

6   on.  As this system progresses to the point that it

7   might be considered for application, I would trust

8   that a lot of these issues will be worked out and

9   that information, if it is brought before the

10   committee, information will be provided.

11             I would like to now declare the public

12   comment session closed, and we will turn the

13   committee's discussion as our final item for the

14   day to providing recommendations or advice to the

15   FDA with regard to this issue of bacterial

16   contamination detection and type of clinical trials

17   and clinical study design that might be necessary

18   to move it from the QA or QC stage to the product

19   release stage.

20             Is that correctly stated, Jay?

21             DR. EPSTEIN:  Well, this was an

22   informational topic, and we would appreciate a

430

1   general discussion by the committee at the

2   committee's discretion, but we aren't posing

3   specific questions.

4             Committee Discussion and Recommendations

5             DR. ALLEN:  I guess a question I have got,

6   and I will direct this to Dr. Epstein, Dr. Vostal,

7   whoever else at the FDA might want to answer it, I

8   think this is a very important study to do.

9   Obviously, with detection products that have very

10   different designs, it makes it difficult because

11   perhaps you need to replicate the study using each

12   one of the different methods.

13             Nonetheless, we also need to look at the

14   broader system and how well it is working in

15   general, and it would seem to me that this is a

16   high priority.  I would like to suggest that we

17   need to bring the stakeholders to the table and try

18   to get them to come to some degree of agreement in

19   terms of study design and funding.

20             If we could get, whether it's 4 million or

21   5 million or 6 million, whatever the cost is going

22   to be, we know that a large number of our blood

431

1   centers are already doing the early testing as

2   urged by the AABB.  That is one group of

3   stakeholders, the manufacturers are another, and

4   certainly the Public Health Service generally, and

5   I am using that to indicate all of the agencies of

6   the Public Health Service including the FDA, the

7   NIH, and the CDC, certainly have an interest in

8   this also, and it would seem to me that we ought to

9   try to get perhaps a consortium of funding from

10   different sources that meet the need and that would

11   enable us, as quickly and efficiently as possible,

12   to answer the question using all of the licensed,

13   the QC license systems that are out there.

14             Other comments or recommendations?  Yes,

15   Dr. Goldsmith.

16             DR. GOLDSMITH:  Maybe just one.  Do I

17   understand this correctly, is there a dichotomy in

18   clinical use of platelets, that some actually have

19   testing in the field, those that follow the AABB

20   principles, and there are other blood centers or

21   hospitals or clinics that do not follow these

22   principles current?  Just a question.

432

1             DR. ALLEN:  Dr. Strong, you are nodding

2   your head yes.

3             DR. STRONG:  I think that is correct.

4   AABB is a voluntary organization, so if you choose

5   to belong, then, you would have to meet that

6   standard, but you can choose not to belong, which

7   many hospitals do, and some have chosen not to

8   belong for these kinds of reasons.

9             DR. KUEHNERT:  I just want to add I think

10   there is a further dichotomy developing, and that

11   is between apheresis and whole blood drive

12   platelets, and that is of just as great a concern,

13   I think, as we are seeing, how these surrogate

14   tests are performing, and it is creating a very

15   concerning gap in potential patient safety.

16             DR. ALLEN:  Dr. Doppelt.

17             DR. DOPPELT:  I am still just a little bit

18   confused about the role or indication to do the

19   Gram stain in the study.  I mean we heard sort of

20   two scenarios.  One is that if you have a low level

21   of contamination early on, by waiting longer, you

22   would get a positive culture that otherwise would

433

1   have been negative.

2             The other flip side is I thought I heard

3   it said that if the contamination is high, that it

4   might somehow overgrow the material as a median,

5   the cultures would be negative.

6             Is that a theoretical thought or is that

7   real?

8             DR. ALLEN:  It is theoretically possible

9   that you could get a positive culture early that

10   essentially would die out by the time that you do

11   the testing.  Whether that is practically possible

12   or not, I don't know.  My guess would be that it is

13   not highly likely to occur and that you would still

14   get positive cultures even at a very late date, if

15   have got a highly pathogenic organism that very

16   happily grows in platelets at room temperature or

17   the ambient temperature at which they are stored,

18   it is not room temperature.

19             DR. DOPPELT:  I mean it just seems that if

20   it is more theoretical, then, you are adding a lot

21   of time and expense to the study, that might not

22   have much of a return.

434

1             DR. ALLEN:  If there is a question about

2   it, it's a study that could be done very easily.  I

3   know if you go back to some of the earlier

4   materials that were provided to the committee in

5   past deliberations on bacterial contaminations, Dr.

6   Brecher from the University of North Carolina has

7   done a number of inoculation studies looking at

8   growth patterns, and so on.

9             I haven't read these papers in detail, at

10   least to the extent that I was on the committee for

11   one of the discussions, not for one of the ones,

12   but it is very easy to answer these with small,

13   mini studies, and you can make a determination of

14   whether or not Gram stains are necessary.

15             I certainly would agree with Dr. Klein

16   that doing 50,000 Gram stains on materials is not

17   going to be highly efficient or productive

18   mechanisms.

19             Dr. Strong.

20             DR. STRONG:  Just to comment on that.  It

21   depends on the detection system one is using.  The

22   BacT/Alert is a continuous monitoring system, and

435

1   certainly a high degree of detection would be

2   picked up throughout the course of the incubation

3   period.

4             It is an issue probably more for the

5   single point detection systems where you get only

6   one measurement, in which case you might get an

7   intermediate growth after that assay was performed

8   that would be missed.

9             I have a question for FDA concerning the

10   pooling proposal, that if we used pools of 10, we

11   could reduce the number of samples that we had to

12   test at the 7-day point of detection.

13             Would that mean we would also need to do

14   pools at the beginning to compare day 1 with day 7?

15             DR. VOSTAL:  I think that pooling upfront

16   would not be advisable because you lose sensitivity

17   when you pool, you know, by the number of units

18   that you combine.  I think since we are already

19   stressing the sensitivity of the devices by trying

20   to get a sample as early as possible, pooling would

21   even decrease that.

22             DR. KUEHNERT:  I just wanted to bring up a

436

1   point, because it was mentioned before, how this

2   study would be useful to apply to other methods,

3   and there was some discussion about storing samples

4   for the study.  I am not sure how they would be

5   stored and I doubt they would preserve the platelet

6   function, and may result in samples that would not

7   be optimal.

8             I am just wondering if that is something

9   that FDA and others would consider to be usable

10   samples for later analysis should there be a method

11   that has been approved and could be applied.

12             DR. EPSTEIN:  I think platelet function

13   would no longer be relevant because we are looking

14   at these samples for bacterial detection, and this

15   has not been talked through, but one concept would

16   be you save things prospectively, but then you only

17   retain the positives, because then you could ask

18   the question whether some other new technology

19   could pick up all the positives that were detected

20   by this technology.

21             I think it is feasible to save the

22   samples, they are just frozen aliquot--

437

1             DR. KUEHNERT:  If the platelets then broke

2   down, I mean obviously, the function is not an

3   issue.

4             DR. EPSTEIN:  They just don't matter

5   anymore.

6             DR. KUEHNERT:  The function of it is not

7   an issue, but it is not thought that there would be

8   any inhibitory effect of whatever the platelet

9   breakdown products would be that would change  the

10   sensitivity of the culture.

11             DR. EPSTEIN:  I think that is a fair

12   point, and you would have to find that out.  Before

13   you used the repository, you would want to do a

14   model experiment to make sure that the freeze/thaw,

15   or whatever the storage method, didn't create

16   sensitivity artifact.

17             DR. ALLEN:  Dr. Strong.

18             DR. STRONG:  We routinely have a retention

19   sample on all of our units that we sample, and in

20   order to go back and confirm whether this is a true

21   positive or false positive, but one of the issues

22   that has come up there, at least with the company,

438

1   is that the bags that we are using the retention

2   sample for have not been validated to maintain the

3   viability of the bacteria, whatever that may be,

4   and also they have argued that we are just missing

5   the anaerobes.

6             So, we have indeed gone back and done a

7   study to culture those that turn up positive for

8   both aerobe and anaerobic organisms, but none of

9   those have turned out to be positive, but I think

10   that is going to bring up the issue of validating

11   whatever storage medium bag method one uses in

12   order to use retention samples to sort out true

13   from false positives.

14             DR. ALLEN:  Certainly, in general, and I

15   am not talking specifically about platelets because

16   I don't know, but as a general principle, if your

17   bacterium, especially your more fastidious

18   organisms in a low inoculum are not in a friendly

19   or growth conducive environment, they will tend  to

20   die off over time.

21             So, it is possible that you could have

22   some probably skin contaminants present on day 1

439

1   that would not be present at day 7, may not be

2   picked up by the limited culture systems and the

3   limited inoculum size that is used for the day 1

4   culture.

5             I am not sure that those are really going

6   to be clinically significant anyhow.  Again, it's a

7   theoretical argument that probably can be best

8   answered by studies of the type that can be done

9   quickly and easily in the laboratory using a small

10   number of specimens to answer the question.

11             DR. STRONG:  Well, that is one of the

12   challenges of these studies is what is a clinically

13   significant infection.  In some respects, we have

14   increased the risk by having to delay release of

15   platelets 24 hours.

16             We know that most of our infections

17   anecdotally over the past have occurred with day 4

18   or day 5 platelets rather than day 1 or day 2

19   platelets, and now we have extended the storage

20   prior to release, so what is the relevant

21   contamination number that will be clinically

22   important.

440

1             Our experience had been the true positives

2   tend to come up early, they are positive within 24

3   to 48 hours, and it is the highly questionable ones

4   that are coming up like day 5, which we can't

5   confirm and are in question.

6             DR. ALLEN:  In terms of what is clinically

7   relevant, depends on the host also.  Each one of

8   us, you know, I have got to go back and brush my

9   teeth before I go to bed tonight, I am going to

10   shower my blood system with bacteria in that

11   process, you know, I am going to handle it without

12   any problem.

13             The fact is that does happen to normal

14   humans all the time.  It is when you get your

15   high-risk patients, those who are immunosuppressed,

16   those who are critically ill in  the hospital,

17   exactly the kind of patient often that is going to

18   be getting platelets and other blood products that

19   even relatively normally, you know, the kinds of

20   bacteria to which we are normally exposed and can

21   handle without any problem is going to cause

22   potential problems.

441

1             On the other hand, it is surprising what

2   people can handle on a routine basis if it's not

3   presented in too large a number, you know, we have

4   got a relatively efficient humoral and cellular

5   immunity system to handle that sort of thing, and

6   if it functions at all, we can handle a certain

7   amount of stuff.

8             DR. STRONG:  We actually have an anecdotal

9   example of just that.  We had a contaminated unit,

10   this was prior to culture, in which it just

11   happened that the same patient got both splits from

12   this unit, and the day 3 unit caused no problems,

13   and the day 4 unit generated a septic reaction, and

14   this is a patient in the ICU.

15             DR. DiMICHELE:  I apologize if this was

16   discussed while I was out of the room for a while,

17   but I just wanted to comment on the design, on the

18   revision of the design.

19             To save money, which is a really good

20   reason to revise your design, but to go to just day

21   1 and day 7, I am assuming at this point that there

22   are many different factors still that are out with

442

1   respect to ascertaining whether we are going to go

2   to 7 days of platelet storage, bacterial

3   contamination just being one of them.

4             I am just wondering, you know, you are

5   spending a lot of time and effort looking at

6   upfront testing on platelets to see whether it

7   makes a difference at time of release.  There is

8   still a possibility that for other reasons, you may

9   not be able to go to 7-day platelets, is that

10   possible, because at that point, I am just

11   wondering if maybe you may still need to do day 5

12   testing just to anticipate that possibility.

13             DR. VOSTAL:  I think the two main

14   questions are bacterial detection and then platelet

15   efficacy, and we have bags approved already for

16   7-day platelets, have been validated by the

17   radiolabeling studies to be able to store platelets

18   out to 7 days.

19             So, from that point of view, we are ready

20   to go, it is really a question of whether we can

21   get a bacterial detection device.

22             DR. DiMICHELE:  But in terms of redefining

443

1   the standards for platelet efficacy, like what we

2   have been talking about today, will those be

3   applied now to the 7-day bags, and is there a

4   potential that there might be different validation

5   that is needed?

6             DR. VOSTAL:  The ones that the bags have

7   been approved, they are set.  We are not going to

8   reapply the new criteria to them, but any new bags

9   from a manufacturer that comes in now will go

10   through the type of testing we were talking about

11   today.

12             DR. DiMICHELE:  So, really, it is just

13   this testing that stands in the way at this point.

14   Okay.

15             DR. ALLEN:  Let me be the devil's

16   advocate.  Do a second phase testing and not worry

17   about anywhere between 5 and 10 days, if there are

18   other reasons why it is safe and effective to allow

19   storage of the platelets to that point.

20             DR. KLEIN:  Before the session ends, I

21   just want to make one point that Dr. Kuehnert made,

22   and I just hope that it has been appreciate both by

444

1   the panel and I am sure by the FDA.

2             That is, that right now most of the

3   apheresis platelets in the country are being tested

4   by a culture method, so despite the fact that that

5   is not a pre-release test, they are being cultured

6   and infectious units are being interdicted.

7             It is the apheresis, it is the single

8   donor pooled platelets that are being tested by

9   Gram stain, by swirling, and by dipstick

10   technology, and that is where we are still having a

11   public health issue.  We really do need to move on

12   that before we think about doing lots of other

13   kinds of things where safety actually has been

14   dramatically improved.

15             In this area, safety hasn't been improved,

16   and it is a real issue.

17             DR. ALLEN:  I think that is a very

18   important point that both of you have made, and

19   obviously, integral to that is the issue of whether

20   you are going to require culture or some other

21   methodology, bacterial detection methodology on

22   every single random donor unit that is collected or

445

1   whether you are going to allow early pooling with

2   an appropriate detection method there.

3             I think that perhaps is where a lot of the

4   focus ought to be, I agree, unless you are going to

5   say we aren't going to continue with random donor

6   platelets anymore, and that is a very significant

7   question for the industry to have to wrestle with

8   also.

9             Other comments or questions?  Yes.

10             DR. QUIROLO:  That decision may be made

11   economically because I know in San Francisco, that

12   we don't make random donor platelets because they

13   are too expensive. It is cheaper to go with

14   pheresed platelets, so if you wait long enough,

15   they will just go away.

16             DR. KLEIN:  You may have to wait a long

17   time for that.

18             DR. ALLEN:  Everything is expensive in San

19   Francisco.

20             Dr. Epstein, other directions you would

21   like to have the committee extend its discussion?

22             DR. EPSTEIN:  No.  I think this has been a

446

1   good discussion about the approach to bacterial

2   contamination monitoring.  We understand that the

3   discussion may evolve as the technologies evolve,

4   but we are dealing with the here and now, and we

5   have these culture systems, and we have these

6   unvalidated or poorly validated endpoint tests, and

7   we do need to make progress on both fronts.

8             I would also comment that part of the

9   issue with whole blood derived platelets is that we

10   have not moved toward what are the more

11   economically efficient solutions that are presently

12   available in Europe, namely, pre-storage pooled

13   buffy coat derived, leukoreduced platelet, which is

14   cultured once and used out to 7 days.

15             I mean clearly from the practical point of

16   view, that is a better product.  Whether analogous

17   systems can be put in place for the platelet-rich

18   plasma-derived platelet is an open question, and

19   whether the system would be willing to do the

20   necessary validation to move toward a buffy coat

21   platelet and pre-storage pooled platelet is an open

22   question.

447

1             The Agency stands ready to meet the

2   industry halfway, if there is an incentive and a

3   drive to try to validate either the buffy coat

4   platelet or the pre-storage pooled platelet-rich

5   plasma derived platelet, we would be very pleased

6   to see that development, you know, and help

7   shepherd it along, because that is part of the

8   solution.

9             DR. STRONG:  I can confirm that there is a

10   great deal of interest at least at some centers in

11   doing that.  I think it has been the mechanism, how

12   do we get to that point, you know, what are the

13   requirements.

14             In Canada, they have chosen to accept the

15   European data in order to accomplish that, so I

16   mean I would certainly encourage, as a conflicted

17   member, to consider the European data as a way of

18   validating that.

19             DR. ALLEN:  I think this is a very

20   important area for discussion, not necessarily

21   tonight, but I will take the prerogative of Acting

22   Chair for a few minutes more to urge the FDA to

448

1   seriously consider that, as well as the industry,

2   and perhaps this is an area where the AABB and

3   America's Blood Centers would like to make a

4   recommendation or a proposal to the FDA for

5   consideration in terms of how to move forward on

6   this issue.

7             Other comments or questions or anything

8   else?

9             [No response.]

10             DR. ALLEN:  My agenda says that we are to

11   recess at 6:15, and Dr. Smallwood's clock over

12   there says 6:15:40.

13             So, if there is nothing further, I will

14   adjourn the meeting for the day.  See you all at

15   8:00 a.m. tomorrow.

16             [Whereupon, at 6:15 p.m., the proceedings

17   were recessed, to reconvene at 8:00 a.m., Friday,

18   July 23, 2004.]