- MUMPS virus standard RT-PCR overview

Use of trade names or commercial sources is for identification only and does not imply endorsement by CDC.

The protocol provided is for use to develop and validate a mumps detection RT-PCR test as a research tool. CDC cannot guarantee the performance of any reagents or methods. Results should not be described as having been obtained using a CDC test and no indication should be given that CDC has validated the test.

- Primers for Mumps (PCR product will be 675 bp) :

SH1, AGT AGT GTC GAT GAT CTC AT

SH2R, GCT CAA GCC TTG ATC ATT GA

| Primer sequences are from: Jin L, Beard S, Brown DWG. Genetic heterogeneity of mumps virus in the United Kingdom: identification of two new genotypes. J Infect Dis 1999; 180:82933 |  |
| --- | --- |

CDC has prepared aliquots of mumps virus for distribution as a mumps virus control (0.5 ml) consisting of a cell culture lysate of a wild type mumps virus grown in Vero cells.

# Cycling Parameters for Superscript

40 Cycles

30  min

55oC

2 min

94oC

94oC

15 sec

55oC

30  sec

72oC

30  sec

7 min

72oC

4oC

SUPERSCRIPT RT-PCR PROTOCOL

This is a standard RT-PCR reaction. This is a one-tube reaction, so there is a minimum of specimen handling. Some points to consider:

1. Controls.

Positive control should be RNA extracted from cells infected with mumps virus. It is best to use cells infected with a well-characterized strain of mumps, belonging to a genotype not associated with the outbreak. In addition, the use of vaccine virus as a control is not recommended since in an outbreak setting some cases may be referred for testing who have been recently vaccinated. If there is a question of contamination, this will allow sequence studies to determine the possible source of the virus.

Negative controls should include RNA extracted from uninfected cells as well as a control with water substituted for RNA. When doing large numbers of RNA extractions, it is a good idea to include an extraction control. This mock extraction will show that reagents for RNA extraction have not been contaminated by template.

2. The 30-second extension time shown below is for templates less than 1000 bp in length. For longer templates, it is advisable to increase the extension time.

3. Avoid template contamination. Use dedicated equipment, rooms and hoods for all pre-PCR procedures. Post amplification analysis and processing should be performed in separate rooms using dedicated equipment. Do not share equipment (including lab coats) between pre-PCR and post-PCR procedures.  Use filter tips for all pipetting of pre-PCR procedures and for setting up RT-PCR reactions.

4. We usually use 5 ul of RNA per reaction. Most of the extraction protocols yield 40-50 ul of RNA. We usually extract RNA from one 25 cm2 flask of infected cells or from 100-200 ul of original viral sample. The volume of RNA can be increased, but this will not improve the sensitivity.

MATERIALS:

- Superscript kit (Invitrogen CAT# 10928-042)

- RNasin  (pancreatic RNase inhibitor)

- Ice bucket

- Autoclaved PCR tubes (0.2ul, thin-walled)

- Forward Primer (20uM working dilution prepared in RNase-free water)

- Reverse Primer (20uM working dilution prepared in RNase-free water)

INSTRUCTIONS:

- Thaw all kit reagents except enzymes, vortex and place on ice. Keep enzymes (Superscript and RNasin) on ice at all times. Allow RNA samples to thaw on ice and keep on ice while you are setting up the reactions.

- Label appropriate number of 0.2 ml thin-walled, reaction tubes and place in pre-chilled metal cooling rack. Keep cooling rack on ice for entire protocol.

- Add appropriate volumes (see worksheet in appendix) of reagents 1 through 5 to a pre-chilled 1.5 ml Eppendorf tube. Vortex and keep tube on ice.

- Allow time for pre-mix contents to chill, then add reagents 6 and 7 to pre-mix tube. Vortex and chill briefly on ice.

- Dispense pre-mix (see worksheet) to each reaction tube.

- Using a new, clean pipette tip for each transfer, add RNA to each tube and close the cap.

- Spin the tubes briefly (10,000 rpm for 1 minute) in a chilled microcentrifuge and immediately return the tube to the metal cooling rack.

- While tubes are spinning start the appropriate file on a 9700 thermocycler. Use the “view” command to make sure that you are running the proper conditions. When the block temperature reaches 55 C, hit the pause button. With the instrument paused, transfer the reaction tubes from the metal cooling rack to the heat block of the thermocycler. Close the cover and press the resume button.

Appendix

SUPERSCRIPT RT-PCR REACTION WORKSHEET:

Date:______________                                   Operator:_____________________

| Component | [Final] Conc. | Vol/rxn | # rxns plus 1 | Total vol. | Added Y/N |
| --- | --- | --- | --- | --- | --- |
| 1. 2X Superscript Reaction mix | 1X | 25 ul |  |  |  |
| 2. 5mM MgS04 | 0.8mM | 8 ul |  |  |  |
| 3. Forward Primer (20 uM) | 0.2 uM | 0.5 ul |  |  |  |
| 4. Reverse Primer (20 uM) | 0.2 uM | 0.5 ul |  |  |  |
| 5. ETF water (14.5 – vol RNA) |  | ___ul |  |  |  |
| Vortex and place on ice before adding enzymes |  |  |  |  |  |
| 6. Superscript Enzyme |  | 1 ul |  |  |  |
| 7. RNase Inhibitor (optional) | 20 U. | 0.5 ul |  |  |  |
| Vortex and dispense pre-mix (50ul – RNA vol)_______ into each chilled reaction tube, then add RNA samples |  |  |  |  |  |
| 8. RNA |  | ___ ul |  |  |  |

Forward Primer:_______________                      Reverse Primer:_______________

## Total reaction volume: 50 ul

Pre-mix dispense volume (50ul – vol RNA):________