GMP 12

Good Measurement Practice

on

Standard Operating Procedure Selection

Good laboratory practices, use of proper standards and equipment, and selection of standard operating procedures are essential for providing calibration results with accurate and traceable values with appropriate and suitable uncertainties. The following matrix recommends SOPs based on the parameter, type of test items, and level of uncertainty needed.

| Parameter | Test Item | Recommended SOP |
| --- | --- | --- |
| Mass |  |  |
|  | Railroad test cars | SOP 27, Railroad Test Cars using a Master Track Scale |
|  | Weight carts | SOP 33, Calibration of Weight Carts (References SOP 4 and SOP 7) |
| HB 105-1, Class F ASTM, OIML | Class F Class 5, 6, 7 Class M1, M2, M3 e.g., 10 kg to 250 kg (500 lb) cast iron 1 mg to 5 kg (1 lb to 10 lb) stainless steel | SOP 8, Modified Substitution - may be used if expanded uncertainty is less than 1/3 of the tolerance SOP 7, Single Substitution - to be used, as a minimum, if conditions given for SOP 8 cannot be met NOTE: Balances and standards must be selected properly for these conditions to be met. |
| ASTM, OIML | Class 3, 4 (P) Class F1, F2 e.g., 1 kg kit, 100 g kit | SOP 7, Single Substitution - may be used if expanded uncertainty is less than 1/3 of the tolerance. If uncertainty is greater than 1/3 of the tolerance, then use SOP 4. SOP 4, Double Substitution - to be used for buoyancy corrections and if expanded uncertainty is less than 1/3 of the tolerance Modified SOP 4/5, to be used to incorporate measurement control into SOP 4 NOTE: Balances and standards must be selected properly for these conditions to be met. |
| ASTM, OIML | Class 1, 2 (S, S-1) Class E2 for use in balance calibration | SOP 5, 3-1 Weighing Design (preferred) Modified SOP 4/5, to be used to incorporate measurement control with SOP 4 |
| ASTM, OIML | Class 0, 1 (S) Class E1 for use as laboratory standards | Weighing Designs per TN 952, TN 844, NISTIR 5672, SOP 28, with appropriate check standards |
| Volume |  |  |
| HB 105-2 | Glass flasks | SOP 16, Volume Transfer (acceptable, SOP 14 preferred) |
| HB 105-3 | 20 L test measures (5 gal or 10 gal) | SOP 18, Volume Transfer (single delivery from slicker plate type standard; if glass standards SOP 19). When temperature instability is observed during the calibration process, use SOP 19. |
| HB 105-3 | Large graduated neck type provers - used for meter verification | SOP 19, Volume Transfer for Graduated Neck Type Provers |
| HB 105-4 | LPG provers | SOP 21, Volume Transfer for LPG Provers |
| HB 105-7 | Dynamic Small Volume Provers | SOP 26, Gravimetric Calibration of Dynamic Small Volume Provers |
|  | Laboratory standards Glassware: burets, pipetes, flasks | SOP 13, Gravimetric Calibration with a Single Pan Mechanical Balance OR SOP 14, Gravimetric Calibration with an Electronic Balance |
|  | Laboratory standards Laboratory slicker plate standards | SOP 14, Gravimetric Calibration with an Electronic Balance OR SOP 15, Gravimetric Calibration with an Equal Arm Balance |
|  | Micropipetes | SOP 13, Gravimetric Calibration with a Single Pan Mechanical Balance OR SOP 14, Gravimetric Calibration with an Electronic Balance |
| HB 105-3 | Large graduated neck type provers - used as laboratory standards | SOP 19, Volume Transfer for Graduated Neck Type Provers OR SOP 14, Gravimetric Calibration with an Electronic Balance OR SOP 15, Gravimetric Calibration with an Equal Arm Balance |
| Length |  |  |
|  | Tapes | SOP 11, Tape to Tape OR SOP 12, Bench Method (lower uncertainties) |
|  | Rigid Rules | SOP 10, Rigid Rule |
|  | Pi Tapes | SOP 23, Pi Tape Calibration |
| # Liquid-in-Glass Thermometers |  |  |
| HB 105-6 | Field standards for weights and measures | SOP 25 |
| Timing Devices |  |  |
| HB 105-5 | Field standards for weights and measures | SOP 24 |
| Traffic Speed Gun Tuning Forks |  |  |
|  | For highway official use | SOP 22 |

| Parameter | Process | Recommended SOP |
| --- | --- | --- |
| Measurement Assurance |  |  |
| All Process Measurement Assurance | SOP 30 | Use of process measurement assurance programs |
| Mass | SOP 5, 28 | Use of check standards in procedure TN 952, TN 844, NISTIR 5672 |
| Mass | SOP 4, 6, 7, 8 | SOP 9 |
| Length | SOP 10, 11, 12, 23 | Redundancy built into procedures |
| Volume | SOP 13, 14, 15, 16, 18, 19, 21, 26 | SOP 17, laboratory check standards OR SOP 20, range charts and standard deviation charts |
| Temperature | SOP 25 | Use of check standards in procedure |
| Uncertainty |  |  |
| All parameters | All SOPs | SOP 29, Calculation of Uncertainty |